## Alarms
PRS ID | HD SRS ID | DG SRS ID | Device | Displayed Title | Displayed Instructions | Description | Trigger Conditions | AlarmID | Alarm Name | Priority | Rank | Fault | DGFault | Stop | Clear Immediate | Non Recoverable | No Resume | No Rinseback | No End Treatment | No Blood Recirc | No Dialysate Recirc | Clear Top Only | TreatmentLog | Escalate In | Escalate To | Doxygen Title
N/A | N/A | N/A | N/A | N/A | N/A | Indication for when no alarms is a possible situation | N/A | 0 | ALARM_ID_NO_ALARM | ALARM_PRIORITY_NONE | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Indication for when no alarms is a possible situation
|  |  | HD | Self Test Error | Call service to schedule a repair. | Stuck button POST failure. Stop or Off button detected to be pressed for at least 1 second during test shortly after power up | Stop or Off button are pressed for more than 1 sec during POST. | 1 | ALARM_ID_STUCK_BUTTON_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Stuck button POST failure. Stop or Off button detected to be pressed for at least 1 second during test shortly after power up
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD FPGA POST failure | Minimum HD FPGA major revision is not compatible with this version of HD firmware or the FPGA Id is not correct. | 2 | ALARM_ID_HD_FPGA_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD FPGA POST failure
|  |  | DG | Self Test Error | Call service to schedule a repair. | DG FPGA POST failure | Minimum DG FPGA major and minor revision are not equal to the FPGA major and minor revision in FPGA header | 3 | ALARM_ID_DG_FPGA_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG FPGA POST failure
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD Watchdog POST failure | If voltage reading is less than 22.6Volts when the Watchdog recovered If PIN_SIGNAL is not LOW If backup alarm audio is on when Watchdog expired | 4 | ALARM_ID_HD_WATCHDOG_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD Watchdog POST failure
|  |  | DG | Self Test Error | Call service to schedule a repair. | DG Watchdog POST failure | If voltage reading > 5.0 Volts when the Watchdog recovered If PIN_SIGNAL is not LOW If voltage reading is less than 22.6Volts when the Watchdog recovered | 5 | ALARM_ID_DG_WATCHDOG_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG Watchdog POST failure
|  |  | HD | Self Test Error | Call service to schedule a repair. | UI communication POST failure. UI failed to communicate within a reasonable time after power up | If the UI failed to checkin via the CAN bus within 40 seconds | 6 | ALARM_ID_UI_COMM_POST_FAILED | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | UI communication POST failure. UI failed to communicate within a reasonable time after power up
|  |  | HD | Machine Error | Call service to schedule a repair. | Real-time clock configuration error | If the RTC is not functional as indicated by the bits in the first 3 control registers. | 7 | ALARM_ID_HD_RTC_BATTERY_LOW | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | HD RTC battery low
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD accelerometer failed POST | If the calculated vector length of the accelerometer x, y, z is not the nominal length. | 8 | ALARM_ID_HD_ACCELEROMETER_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD accelerometer failed POST
|  |  | DG | Self Test Error | Call service to schedule a repair. | DG accelerometer failed POST | If the calculated vector length of the accelerometer x, y, z is not the nominal length. | 9 | ALARM_ID_DG_ACCELEROMETER_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG accelerometer failed POST
|  |  | HD | Machine Error | Call service to schedule a repair. | HD blood leak sensor embedded mode communication failed | If the blood leak sensor embedded mode is unresponsive. | 10 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_EMBEDDED_MODE_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood leak sensor setting embedded mode failure.
|  |  | HD | Dialysate Temperature Error | Call service to schedule a repair. | Dialysate temperature is below the target temperature | If the dialysate temperature is below the target temperature for a certain period of time. | 11 | ALARM_ID_HD_DIALYSATE_TEMP_BELOW_TARGET_TEMP | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | ALM_ESC_10_MIN | ALARM_ID_TREATMENT_STOPPED_NO_RINSEBACK | HD dialysate temperature below target temperature
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD firmware image integrity POST test failed | If the calculated CRC of the HD image does not match the expected CRC. | 12 | ALARM_ID_HD_INTEGRITY_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD firmware image integrity POST test failed
|  |  | DG | Self Test Error | Call service to schedule a repair. | DG firmware image integrity POST test failed | If the calculated CRC of the DG firmware image does not matched the expected CRC | 13 | ALARM_ID_DG_INTEGRITY_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG firmware image integrity POST test failed
|  |  | DG | Self Test Error | Call service to schedule a repair. | DG non-volatile usage information group invalid record CRC | If the DG usage information record's CRC does not match the calculated CRC in POST. | 14 | ALARM_ID_DG_INVALID_USAGE_RECORD_CRC | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG invalid usage record CRC
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD non-volatile usage information group invalid record CRC | If the HD usage information record's CRC does not match the calculated CRC in POST. | 15 | ALARM_ID_HD_INVALID_USAGE_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD invalid usage record CRC
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD alarm audio failed POST | If alarm audio current is NOT sufficiently high for at more than 1 second indicating alarm tone is being output | 16 | ALARM_ID_HD_ALARM_AUDIO_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD alarm audio failed POST
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD UI POST failed | If User Interface has not sent POST results within 2 seconds. | 17 | ALARM_ID_HD_UI_POST_FAILED | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD UI POST failed
|  |  | DG | Calibration Error | Call service to schedule a repair. | DG pressure sensors non-volatile calibration record invalid CRC | If the calibration record's CRC does not match the calculated CRC in POST. | 18 | ALARM_ID_DG_PRESSURE_SENSORS_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG load cells invalid calibration record
|  |  | HD | Dialysate Temperature Error | Call service to schedule a repair. | Dialysate temperature is above the target temperature | If the dialysate temperature is above the target temperature for a certain period of time. | 19 | ALARM_ID_HD_DIALYSATE_TEMP_ABOVE_TARGET_TEMP | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | ALM_ESC_10_MIN | ALARM_ID_TREATMENT_STOPPED_NO_RINSEBACK | HD dialysate temperature above target temperature.
|  |  | DG | Calibration Error | Call service to schedule a repair. | DG conductivity sensors invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 20 | ALARM_ID_DG_COND_SENSORS_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG conductivity sensors invalid calibration record
|  |  | DG | Calibration Error | Call service to schedule a repair. | DG drain line volume invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 21 | ALARM_ID_DG_DRAIN_LINE_VOLUME_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG drain line volume invalid calibration record
|  |  | DG | Calibration Error | Call service to schedule a repair. | DG reservoirs invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 22 | ALARM_ID_DG_RESERVOIRS_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG reservoirs invalid calibration record
|  |  | DG | Calibration Error | Call service to schedule a repair. | DG acid concentrate calibration record | If the non-volatile data is valid and if not raise the alarm | 23 | ALARM_ID_DG_ACID_CONCENTRATE_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG acid concentrate calibration record
|  |  | DG | Calibration Error | Call service to schedule a repair. | DG bicarb concentrate calibration record | If the non-volatile data is valid and if not raise the alarm | 24 | ALARM_ID_DG_BICARB_CONCENTRATE_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG bicarb concentrate calibration record
|  |  | DG | Calibration Error | Call service to schedule a repair. | DG accelerometer invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 25 | ALARM_ID_DG_ACCELEROMETERS_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG accelerometer invalid calibration record
|  |  | HD | Calibration Error | Call service to schedule a repair. | HD accelerometer invalid calibration record | If the accel calibration data that was received from NVDataMgmt is NOT legitimate. | 26 | ALARM_ID_HD_ACCELEROMETERS_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD accelerometer invalid calibration record
|  |  | HD | BLD Sensor Error | Call service to schedule a repair. | HD BLD sensor failed zero sequence | BLD fails to complete zero sequence. | 27 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_ZERO_SEQUENCE_FAILED | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood leak sensor zero and self test sequence failed
|  |  | DG | FPGA Error | Call service to schedule a repair. | DG two wire senosors FPGA fault | If the FPGA does not report fresh data or indicates an error for a certain period of time | 28 | ALARM_ID_DG_TWO_WIRE_SENSORS_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG two wire sensors FPGA fault
|  |  | HD | Calibration Error | Call service to schedule a repair. | HD heparin force sensor invalid calibration record | If the heparin force sensor calibration data that was received from NVDataMgmt is NOT legitimate. | 29 | ALARM_ID_HD_HEPARIN_FORCE_SENSOR_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD heparin force sensor invalid calibration record
|  |  | HD | Machine Error | Call service to schedule a repair. | HD Software fault. Software found itself in an unexpected state | If the specific software faults listed in alarm management SW faults has been detected | 30 | ALARM_ID_HD_SOFTWARE_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD Software fault. Software found itself in an unexpected state
|  |  | HD | Blood Pump Error | Call service to schedule a repair. | Blood pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | If the measured MC current is not greater than BP_MAX_CURR_WHEN_STOPPED_MA or is not greater than BP_MAX_CURR_WHEN_RUNNING_MA. | 31 | ALARM_ID_BLOOD_PUMP_MC_CURRENT_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running
|  |  | HD | Blood Pump Error | Call service to schedule a repair. | Blood pump failed motor off check. Measured speed while commanded off | If the measured motor speed is greater than BP_MAX_MOTOR_SPEED_WHILE_OFF_RPM | 32 | ALARM_ID_BLOOD_PUMP_OFF_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood pump failed motor off check. Measured speed while commanded off
|  |  | HD | Blood Pump Error | Call service to schedule a repair. | Blood pump failed motor direction check. Measured vs. commanded | If the set direction versus hall sensors direction versus sign of motor speed DOES NOT agree. | 33 | ALARM_ID_BLOOD_PUMP_MC_DIRECTION_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood pump failed motor direction check. Measured vs. commanded
|  |  | HD | Blood Pump Error | Call service to schedule a repair. | Blood pump failed rotor speed check. Mismatch with rotor and motor speeds | If the delta of the measured motor speed and the commanded motor speed is not within tolerance. | 34 | ALARM_ID_BLOOD_PUMP_ROTOR_SPEED_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood pump failed rotor speed check. Mismatch with rotor and motor speeds
|  |  | HD | Dialysate Pump Error | Call service to schedule a repair. | Dialysis inlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | If the Motor Controller measured DPi pump current > 150.0 mA for more than 250.0 msec. | 35 | ALARM_ID_DIAL_IN_PUMP_MC_CURRENT_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis inlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running
|  |  | HD | Dialysate Pump Error | Call service to schedule a repair. | Dialysis inlet pump failed motor off check. Measured speed while commanded off | If during command off state, the DPi pump motor speed is measured > 100.0 RMP. | 36 | ALARM_ID_DIAL_IN_PUMP_OFF_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis inlet pump failed motor off check. Measured speed while commanded off
|  |  | HD | Dialysate Pump Error | Call service to schedule a repair. | Dialysis inlet pump failed motor direction check. Measured vs. commanded | If the set direction versus hall sensors direction DO NOT agree for more than 250.0 msec. | 37 | ALARM_ID_DIAL_IN_PUMP_MC_DIRECTION_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis inlet pump failed motor direction check. Measured vs. commanded
|  |  | HD | Dialysate Pump Error | Call service to schedule a repair. | Dialysis inlet pump failed rotor speed check. Mismatch with rotor and motor speeds | If the measured motor speed versus the measured motor speed while controlling the target > 5.0 for more than 12 seconds. | 38 | ALARM_ID_DIAL_IN_PUMP_ROTOR_SPEED_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis inlet pump failed rotor speed check. Mismatch with rotor and motor speeds
|  |  | HD | Dialysate Pump Error | Call service to schedule a repair. | Dialysis outlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | If the DPo pump current > 150.0 mA while stopped or Dpo current > 2000.0 mA for more than 2 seconds while running. | 39 | ALARM_ID_DIAL_OUT_PUMP_MC_CURRENT_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis outlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running
|  |  | HD | Dialysate Pump Error | Call service to schedule a repair. | Dialysis outlet pump failed motor off check. Measured speed while commanded off | If the Dpo pump current > 100.0 RPM while commanded off for more than 5000 msec. | 40 | ALARM_ID_DIAL_OUT_PUMP_OFF_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis outlet pump failed motor off check. Measured speed while commanded off
|  |  | HD | Dialysate Pump Error | Call service to schedule a repair. | Dialysis outlet pump failed motor direction check. Measured vs. commanded | If the DPo pump set direction does not agree with hall sensors or set direction does not agree with the sign of the motor speed for more than 25 msec. | 41 | ALARM_ID_DIAL_OUT_PUMP_MC_DIRECTION_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis outlet pump failed motor direction check. Measured vs. commanded
|  |  | HD | Dialysate Pump Error | Call service to schedule a repair. | Dialysis outlet pump failed rotor speed check. Mismatch with rotor and motor speeds | If DPo pump measured rotor speed does not agreed with the measured/commanded motor speed | 42 | ALARM_ID_DIAL_OUT_PUMP_ROTOR_SPEED_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis outlet pump failed rotor speed check. Mismatch with rotor and motor speeds
|  |  | DG | FPGA Error | Call service to schedule a repair. | Main primary heater FPGA fault | If the main primary heater's voltage FPGA does not report fresh data or it reports and error for a certain period of time. | 43 | ALARM_ID_DG_MAIN_PRIMARY_HEATER_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG main primary heater FPGA fault
|  |  | HD | Machine Error | Call service to schedule a repair. | User interface communication timeout | If the sub-system communication timeout has occurred. UI is 5 seconds. DG is 2 seconds. | 44 | ALARM_ID_UI_COMM_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | User interface communication timeout
|  |  | HD | Machine Error | Call service to schedule a repair. | Too many bad CRCs detected on received system messages | If too many bad messages CRCs are received within a set period of time | 45 | ALARM_ID_HD_COMM_TOO_MANY_BAD_CRCS | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD too many bad communications CRC
|  |  | HD | Machine Error | Call service to schedule a repair. | System message that required acknowledgment was not acknowledged | If no retries left in the pending acknowlegement list. | 46 | ALARM_ID_HD_CAN_MESSAGE_NOT_ACKED | ALARM_PRIORITY_HIGH | 3 | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD System message that required acknowledgment was not acknowledged
|  |  | HD | Ultrafiltration Error | Call service to schedule a repair. | Ultrafiltration rate is too high error during treatment | If the UF rate > 2750.0 mL per hour. | 47 | ALARM_ID_UF_RATE_TOO_HIGH_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Ultrafiltration rate is too high error during treatment
|  |  | HD | Ultrafiltration Error | Call service to schedule a repair. | Ultrafiltration volume accuracy error during treatment | If the UF volume (refUFvolume - measureVolume) > 250.0 mL. | 48 | ALARM_ID_UF_VOLUME_ACCURACY_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Ultrafiltration volume accuracy error during treatment
|  |  | HD | Machine Error | Call service to schedule a repair. | HD FPGA communication down for too long | If retries for commands > 5 counts. | 49 | ALARM_ID_HD_FPGA_COMM_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD FPGA communication down for too long
|  |  | DG | Machine Error | Call service to schedule a repair. | DG FPGA not accepting commanded valve states | If the read back FPGA valve states does not match the commanded valve states and the mismatch counter > 3 | 50 | ALARM_ID_DG_VALVE_CONTROL_FAILURE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG FPGA not accepting commanded valve states
|  |  | HD | Blood Pump Error | Call service to schedule a repair. | Blood pump failed motor speed check. Measured vs. commanded | If measured motor speed versus commanded rotor speed exceeded limits. | 51 | ALARM_ID_BLOOD_PUMP_MOTOR_SPEED_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood pump failed motor speed check. Measured vs. commanded
|  |  | HD | Dialysate Pump Error | Call service to schedule a repair. | Dialysate inlet pump failed motor speed check. Measured vs. commanded | If measured motor speed versus commanded motor speed exceeded limits for more than 5 secs. | 52 | ALARM_ID_DIAL_IN_PUMP_MOTOR_SPEED_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysate inlet pump failed motor speed check. Measured vs. commanded
|  |  | HD | Dialysate Pump Error | Call service to schedule a repair. | Dialysate outlet pump failed motor speed check. Measured vs. commanded | If measured motor speed versus commanded motor speed exceeded limits for more than 5 secs. | 53 | ALARM_ID_DIAL_OUT_PUMP_MOTOR_SPEED_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysate outlet pump failed motor speed check. Measured vs. commanded
|  |  | HD | Machine Error | Call service to schedule a repair. | HD critical data integrity check failed | If the critical data failed range check. | 54 | ALARM_ID_HD_CRITICAL_DATA_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD critical data integrity check failed
|  |  | DG | Machine Error | Call service to schedule a repair. | DG critical data integrity check failed | If the critical data failed range check. | 55 | ALARM_ID_DG_CRITICAL_DATA_ERROR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG critical data integrity check failed
|  |  | HD | Machine Error | Call service to schedule a repair. | HD accelerometer error (no readings or FPGA reports error) | If no fresh sample of accelerometer data has been detected for more than 10 msec. | 56 | ALARM_ID_HD_ACCELEROMETER_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD accelerometer error (no readings or FPGA reports error)
|  |  | DG | Machine Error | Call service to schedule a repair. | DG accelerometer error (no readings or FPGA reports error) | If the accel sensor sampling counter as reported by the FPGA indicated 0 count | 57 | ALARM_ID_DG_ACCELEROMETER_FAILURE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG accelerometer error (no readings or FPGA reports error)
|  |  | HD | Valve Error | Call service to schedule a repair. | HD valve homing failed | If the number of failed homing > = 3 counts. | 58 | ALARM_ID_HD_VALVE_HOMING_FAILED | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD valve homing failed
|  |  | HD | Valve Error | Call service to schedule a repair. | HD valve transition time out | if the valve's transition time has timed out (3000msec) | 59 | ALARM_ID_HD_VALVE_TRANSITION_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD valve transition time out
|  |  | HD | Valve Error | Call service to schedule a repair. | HD valve not functional | If the valve bit masks are not as expected in different valve control mode. | 60 | ALARM_ID_HD_VALVE_NOT_FUNCTIONAL | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD valve not functional
|  |  | HD | Valve Error | Call service to schedule a repair. | HD valve current out of range | If the valve current is over the threshold of 1.0 A for 1000 msec. | 61 | ALARM_ID_HD_VALVE_CURRENT_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD valve current out of range
|  |  | HD | Valve Error | Call service to schedule a repair. | HD valve position out of target | If the current position has deviated from the expected position for more than 1 second. | 62 | ALARM_ID_HD_VALVE_POSITION_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD valve position out of target
|  |  | HD | FPGA Error | Call service to schedule a repair. | FPGA reports error | If the FPGA error counter increments for a certain period of time. | 63 | ALARM_ID_HD_BLOOD_LEAK_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood leak FPGA fault
|  |  | DG | Barometric Pressure Error | Call service to schedule a repair. | Barometric pressure is out of range | If the barometric pressure is > 10.1 psi or < 15.4 psi for a certain period of time. | 64 | ALARM_ID_DG_BARO_PRESSURE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG barometric pressure sensor out of range
|  |  | HD | Machine Error | Call service to schedule a repair. | HD requests DG command with invalid parameter fault | If the latest DG command response indicated invalid parameter | 65 | ALARM_ID_DG_COMMAND_INVALID_PARAMETER_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD requests DG command with invalid parameter fault
|  |  | HD | Machine Error | Call service to schedule a repair. | HD blood leak sensor set point set failure | If the set point of the blood leak sensor does not match the set point in the calibration record. | 66 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_SET_POINT_SET_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood leak sensor set point set failure
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD blood pump occlusion self-test failure alarm | If the blood pump pressure > 5000 | 67 | ALARM_ID_HD_BP_OCCLUSION_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood pump occlusion self-test failure alarm
|  |  | HD | Dialysate Recirculation Error | Dialysate has recirculated for too long. | HD active reservoir recirculation out of range | If the recirculation level has exceeded the max allowed of 1.1 | 68 | ALARM_ID_HD_ACTIVE_RESERVOIR_RECIRCULATION_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD active reservoir recirculation out of range
|  |  | HD | Self-Test Error | Call service to schedule a repair. | HD blood leak sensor invalid calibration record | If the CRC of the blood leak sensor's calibration record does not matcht he calculated CRC in POST. | 69 | ALARM_ID_HD_BLOOD_LEAK_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood leak sensor invalid calibration record
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD arterial pressure self-test failure alarm | If arterial pressure < -300 or > 100 mmHg | 70 | ALARM_ID_HD_ARTERIAL_PRESSURE_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD arterial pressure self-test failure alarm
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD venous pressure self-test failure alarm | If the venous pressure < -100 or > 600 mmHg | 71 | ALARM_ID_HD_VENOUS_PRESSURE_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD venous pressure self-test failure alarm
|  |  | HD | No load cell data received | Call service to schedule a repair. | HD No load cell data message receive at least once every 2 seconds | No new load cell data for 2 seconds. | 72 | ALARM_ID_HD_NEW_LOAD_CELL_DATA_MESSAGE_NOT_RECEIVE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD No load cell data message receive at least once every 2 seconds
|  |  | HD | No dialysate temperature data received | Call service to schedule a repair. | HD No dialysate temperature data message receive at least once every 2 seconds | No new dialysate temperature data for 2 seconds. | 73 | ALARM_ID_HD_NEW_DIALYSATE_TEMPERATURE_DATA_MESSAGE_NOT_RECEIVE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD No dialysate temperature data message receive at least once every 2 seconds
|  |  | DG | UV Reactor Error | Call service to schedule a repair. | DG inlet UV reactor is on while there is no flow | If the inlet UV reactor is on while VPi is closed for a certain period of time. | 74 | ALARM_ID_DG_INLET_UV_REACTOR_ON_WITH_NO_FLOW | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG inlet UV reactor is on with no flow
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD syringe pump self-test failure alarm | If the calibration data that was received from NVDataMgmt is NOT legitimate. | 75 | ALARM_ID_HD_SYRINGE_PUMP_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump self-test failure alarm
|  |  | HD | Machine Error | Call service to schedule a repair. | HD monitored voltage is out of range | If the HD monitored voltages are out of range for a certain period of time. | 76 | ALARM_ID_HD_VOLTAGE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD monitored voltage is out of range
|  |  | DG | Machine Error | Call service to schedule a repair. | DG monitored voltage is out of range | If the DG monitored voltages are out of range for a certain period of time. | 77 | ALARM_ID_DG_VOLTAGE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG monitored voltage is out of range
|  |  | HD | Syringe Pump Error | Call service to schedule a repair. | HD syringe pump direction (from encoder) error | If the measured encoder direction of the syringe pump does not agree with the expected direction | 78 | ALARM_ID_HD_SYRINGE_PUMP_ENCODER_DIRECTION_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump direction (from encoder) error
|  |  | HD | Syringe Pump Error | Call service to schedule a repair. | HD syringe pump direction (from controller) error | If the measured direction of the syringe pump motor controller does not agree with the expected direction | 79 | ALARM_ID_HD_SYRINGE_PUMP_CONTROLLER_DIRECTION_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump direction (from controller) error
|  |  | HD | Syringe Pump Error | Call service to schedule a repair. | HD syringe pump fault reported by FPGA | If the FPGA reported the syringe pump fault | 80 | ALARM_ID_HD_SYRINGE_PUMP_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump fault reported by FPGA
|  |  | HD | Syringe Pump Error | Call service to schedule a repair. | HD syringe pump over travel error | If the syringe pump position < SYRINGE_PUMP_RETRACT_POS_MIN | 81 | ALARM_ID_HD_SYRINGE_PUMP_OVER_TRAVEL_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump over travel error
|  |  | HD | Syringe Pump Error | Call service to schedule a repair. | HD syringe pump DAC write failure | If the syringe pump DAC write error is reported by the FPGA | 82 | ALARM_ID_HD_SYRINGE_PUMP_DAC_WRITE_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump DAC write failure
|  |  | HD | Syringe Pump Error | Call service to schedule a repair. | HD syringe pump is running while the BP is off | If the syringe pump is on while the blood pump is off | 83 | ALARM_ID_HD_SYRINGE_PUMP_RUNNING_WHILE_BP_OFF_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump is running while the BP is off
|  |  | DG | RTC Error | Call service to schedule a repair. | DG RTC set year is invalid | If DG RTC set year is < 2021. | 84 | ALARM_ID_DG_SET_RTC_YEAR_INVALID | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG set RTC year is invalid
|  |  | HD | Pump Track Lever Not Fully Latched | Open front door and pull pump track lever all the way down. | HD pump track lever is not fully latched | If the pump track lever is detected as unlatched when it should be. | 85 | ALARM_ID_PUMP_TRACK_LATCH_OPENED | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD pump track latch opened alarm
|  |  | HD | RTC Error | Call service to schedule a repair. | HD RTC set year is invalid | If HD RTC set year is < 2021. | 86 | ALARM_ID_HD_SET_RTC_YEAR_INVALID | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD set RTC year is invalid
|  |  | DG | Calibration Error | Call service to schedule a repair. | DG heating parameters invalid calibration record | If the heating paramters' non-volatile record's CRC does not match the calculated CRC. | 87 | ALARM_ID_DG_HEATING_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG heating invalid calibration record
|  |  | DG | Concentrate Pumps Error | Call service to schedule a repair. | DG concentrate pump hall sensor out of range. | If the concentrate pumps' pulse width is more than 1000 counts for a certain period of time. | 88 | ALARM_ID_DG_CONC_PUMP_HALL_SENSOR_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG concentrate pumps hall sensor out of range
|  |  | DG | UV Reactor Error | Call service to schedule a repair. | DG outlet UV reactor is on while there is no flow | If the outlet UV reactor is on while the RO flow is less than 200 mL/min. | 89 | ALARM_ID_DG_OUTLET_UV_REACTOR_ON_WITH_NO_FLOW | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG outlet UV reactor on with no flow
|  |  | DG | FPGA Error | Call service to schedule a repair. | DG load cells A1/B1 FPGA fault | If the FPGA does not report fresh data or indicates an error for a certain period of time. | 90 | ALARM_ID_DG_LOAD_CELL_A1_B1_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG load cells A1/B1 FPGA fault
|  |  | HD | Dialysate Measurement Error | Call service to schedule a repair. | HD No dialysate flow data receive in the last 3 seconds | If not receiving new dialysate flow readings in expected interval (3 seconds) | 91 | ALARM_ID_HD_DIALYSATE_FLOW_DATA_NOT_RECEIVE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD No dialysate flow data receive in the last 3 seconds
|  |  | DG | Conductivity Sensors Error | Call service to schedule a repair. | DG conductivity sensor bad status reported | If the conductivity sensors's status reports 0x11. | 92 | ALARM_ID_DG_CONDUCTIVITY_SENSOR_BAD_STATUS | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG conductivity sensor bad status.
|  |  | DG | Conductivity Sensors Error | Call service to schedule a repair. | DG outlet primary conductivity sensor out of range | If the outlet primary conductivity sensor, CPo is greater than 30 uS/cm for a certain period of time. | 93 | ALARM_ID_DG_OUTLET_PRIMARY_CONDUCTIVITY_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG outlet primary conductivity out of range
|  |  | DG | Pressure Sensors Error | Call service to schedule a repair. | DG pressure sensor out of range | If either of the pressure sensors is < 0 psia or > 200 psia or the barometric pressure sensor is > 10.1 psi or < 15.4 psi. | 94 | ALARM_ID_DG_PRESSURE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG pressure out of range
|  |  | DG | Watchdog Error | Call service to schedule a repair. | DG watchdog expired | If the watchdog signal is low while not in POST. | 95 | ALARM_ID_DG_WATCHDOG_EXPIRED | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG watchdog expired
|  |  | DG | Inlet Water Temperature High | Inlet water temperature is high.\n- Check the incoming water for issues with cold water supply. | DG inlet water temeprature out of range | If the inlet water temperature is > 39 C. | 96 | ALARM_ID_INLET_WATER_TEMPERATURE_IN_HIGH_RANGE | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Inlet water temperature in high range
|  |  | DG | Dialysate Fill Error | Call service to schedule a repair. | DG fill conductivity out of range | If the average acid or bicard conductivity is > 5% of the normal acid and bicarb conductivity values. | 97 | ALARM_ID_FILL_CONDUCTIVITY_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | Fill conductivity out of range
|  |  | HD | Battery Error | Call service to schedule a repair. | HD battery communication fault | If battery I2C communication fault persistence count > 5. | 98 | ALARM_ID_HD_BATTERY_COMM_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD battery communication fault
|  |  | HD | Syringe Pump Error | Call service to schedule a repair. | HD syringe pump stall alarm | If syringe pump measured rate < stall speed threshold (0.05) and the ramp up has been restarted 3 times after stall has been detected. | 99 | ALARM_ID_HD_SYRINGE_PUMP_STALL | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump stall alarm
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD no cartridge self-test timeout | If during the no cartridge self-tests state, the elapsed self-test time > maximum no cartridge self test time (30secs) | 100 | ALARM_ID_HD_NO_CART_SELF_TEST_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD no cartridge self-test timeout
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD dry self-test timeout | If the dry self test exceeded 180 seconds. | 101 | ALARM_ID_HD_DRY_SELF_TEST_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD dry self-test timeout
|  |  | DG | Self-Test Error | Call service to schedule a repair. | DG heat disinfection temperature gradient out of range | If temperature gradient between Tpo and Thd exceeds maximum 15°C. | 102 | ALARM_ID_DG_HEAT_DISINFECT_TEMP_GRAD_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG heat disinfect temperature gradient out of range
|  |  | HD | Machine Error | Call service to schedule a repair. | HD non-volatile calibration group invalid record CRC | If the computed CRC of the calibration group record is not equal to the record CRC. | 103 | ALARM_ID_HD_NVDATAMGMT_CAL_GROUP_RECORD_CRC_INVALID | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD invalid calibration CRC
|  |  | HD | Air Trap Error | Call service to schedule a repair. | Air trap level sensors reporting illegal combination of air/fluid | If the air trap level readings are not valid for more than 2 seconds. | 104 | ALARM_ID_AIR_TRAP_ILLEGAL_LEVELS | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Air trap level sensors reporting illegal combination of air/fluid
|  |  | DG | Machine Error | Call service to schedule a repair. | DG non-volatile system group invalid record CRC | If the DG NVDATAMGMT CAL group calculated CRC does not mached the expected CRC. | 105 | ALARM_ID_DG_NVDATAMGMT_CAL_GROUP_RECORD_CRC_INVALID | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG invalid calibration CRC
|  |  | DG | Calibration Error | Call service to schedule a repair. | DG flow sensors calibration error | If the calculated CRC of the flow sensors calibration record does not match the calibration records' CRC. | 106 | ALARM_ID_DG_FLOW_SENSORS_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG dialysate flow sensor invalid calibration record
|  |  | HD | Machine Error | Call service to schedule a repair. | HD reports DG restarted fault | If DG restart command has been sent and the DG is not in the expected mode (Standby and Standby Idle). | 107 | ALARM_ID_DG_RESTARTED_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD reports DG restarted fault
|  |  | HD | Syringe Pump Error | Call service to schedule a repair. | HD syringe pump ADC error | If the ADC read is not fresh after POST init and the ADC read counter has not changed. | 108 | ALARM_ID_HD_SYRINGE_PUMP_FPGA_ADC_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump ADC FPGA fault
|  |  | HD | Syringe Pump Error | Call service to schedule a repair. | HD syringe pump volume check error | If the volume delivered is off by more than 5% or 0.1mL. | 109 | ALARM_ID_HD_SYRINGE_PUMP_VOLUME_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump volume check error
|  |  | HD | Syringe Pump Error | Call service to schedule a repair. | HD syringe pump speed check error | If the syringe measured rate is off by more than 5% or 0.1mL. | 110 | ALARM_ID_HD_SYRINGE_PUMP_SPEED_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump speed check error
|  |  | HD | Syringe Pump Error | Call service to schedule a repair. | HD syringe pump not stopped in off state error | If the syringe pump position is not equal to its last position. | 111 | ALARM_ID_HD_SYRINGE_PUMP_NOT_STOPPED_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump not stopped in off state error
|  |  | DG | FPGA Error | Call service to schedule a repair. | DG load cells A2/B2 FPGA fault | If the FPGA does not report fresh data or indicates an error for a certain period of time. | 112 | ALARM_ID_DG_LOAD_CELL_A2_B2_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG load cells A1/B1 FPGA fault
|  |  | DG | Drain Pump Error | Call service to schedule a repair. | DG drain pump current out of range | If the DG drain pump current > 2.2 A when it is in open loop or if the current > 0.1 A when it is off. | 113 | ALARM_ID_DG_DRAIN_PUMP_CURRENT_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG Drain pump current out of range
|  |  | HD | Air Bubble Detector Error | Call service to schedule a repair. | HD venous air bubble detector self-test failure | If the venous air bubble self test did not detect bubble for more than 500 msec. | 114 | ALARM_ID_HD_VENOUS_BUBBLE_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD venous air bubble detector self-test failure
|  |  | DG | Dialysate Temperature Error | Call service to schedule a repair. | DG temperature sensor out of range | If the measured temperature < 0 C or > 80 C for the non-fluid inline sensors and > 120 C for the fluid sensors. | 115 | ALARM_ID_DG_TEMPERATURE_SENSOR_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG temperature sensor out of range
|  |  | DG | Dialysate Temperature Error | Call service to schedule a repair. | DG temperature sensor ADC out of range | If the ADC value of the sensor is not out of range. | 116 | ALARM_ID_DG_TEMPERATURE_SENSORS_ADC_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG temperature sensor ADC fault
|  |  | DG | Machine Error | Call service to schedule a repair. | System message that required acknowledgment was not acknowledged | If no retries left in the pending acknowlegement list. | 117 | ALARM_ID_DG_CAN_MESSAGE_NOT_ACKED | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG CAN message not acked
|  |  | DG | RTC Error | Call service to schedule a repair. | DG RTC configuration error | If the RTC configuration registers report an error. | 118 | ALARM_ID_DG_RTC_CONFIG_ERROR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | DG RTC config error
|  |  | DG | RTC Error | Call service to schedule a repair. | DG RTC battery low | If the RTC battery bit report a low alarm. | 119 | ALARM_ID_DG_RTC_BATTERY_LOW | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | DG RTC battery low error
|  |  | HD | Pre Treament Error | Call service to schedule a repair. | HD pre-treatment wet prime test failure | If the HD pre-treatment wet self test failed. | 120 | ALARM_ID_PRE_TREATMENT_WET_PRIME_TEST_FAILURE | ALARM_PRIORITY_LOW | 910 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD pre-treatment mode wet self-test prime check failure
|  |  | DG | Heaters Error | Call service to schedule a repair. | DG main primary heater voltage out of range | If the main primary heater's voltage is out of range. | 121 | ALARM_ID_DG_MAIN_PRIMARY_HEATER_VOLTAGE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG main primary heater voltage out of range
|  |  | DG | Heaters Error | Call service to schedule a repair. | DG small primary heater voltage out of range | If the small primary heater's voltage is out of range. | 122 | ALARM_ID_DG_SMALL_PRIMARY_HEATER_VOLTAGE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG small primary heater voltage out of range
|  |  | DG | Heaters Error | Call service to schedule a repair. | DG trimmer heater voltage out of range | If the trimmer voltage is out of range. | 123 | ALARM_ID_DG_TRIMMER_HEATER_VOLTAGE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG trimmer heater voltage out of range
|  |  | HD | Blood Stagnant | Blood has sat too long in the blood tubing set. No rinseback allowed. | HD end of treatment alarm (high priority - user still not acting to end treatment) | In treatment end paused state for > 5 minutes. | 124 | ALARM_ID_END_OF_TREATMENT_HIGH | ALARM_PRIORITY_HIGH | 4 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD end of treatment alarm (high priority - user still not acting to end treatment)
|  |  | HD | Blood Stagnant | Blood has sat too long in the blood tubing set. Treatment must end. | Blood sitting too long after treatment stopped by user alarm (> 5 min) | If in treatment stop sub mode when blood is in the tubing, the blood sitting counter exceeded 5 minutes. | 125 | ALARM_ID_TREATMENT_STOPPED_NO_RINSEBACK | ALARM_PRIORITY_HIGH | 5 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood sitting too long after treatment stopped by user alarm (>5 min)
|  |  | HD | Dialyzer Blood Leak Detected | A blood leak has been detected. Treatment must end. | HD blood leak detected alarm | If the blood leak is detected based on the blood leak status reported by the sensor. | 126 | ALARM_ID_HD_BLOOD_LEAK_DETECTED | ALARM_PRIORITY_HIGH | 6 | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | HD blood leak detected alarm
|  |  | HD | Venous Pressure Low | Venous pressure is low.\n- Check to be sure your blue venous connections are secure.\n- Check that your venous access is connected and secured properly. | Venous pressure too low during treatment | If the venous pressure < TREATMENT_PARAM_VEN_PRESSURE_LOW_LIMIT. | 127 | ALARM_ID_VENOUS_PRESSURE_LOW | ALARM_PRIORITY_HIGH | 8 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Venous pressure too low during treatment
|  |  | HD | Venous Air Bubble Detected | Venous air bubble detected.\n- Check for air bubbles in the venous line.\n- If bubbles are present, tap the tubing to move the bubbles up to the air trap.\n- If bubbles cannot be removed, treatment must end. | HD venous air bubble detected alarm | If bubbles are detected during the treatment. | 128 | ALARM_ID_HD_VENOUS_BUBBLE_DETECTED | ALARM_PRIORITY_HIGH | 9 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | HD venous air bubble detected alarm
|  |  | HD | Venous Air Bubble Detected During Rinseback | Venous air bubble detected.\n- Check for air bubbles in the venous line.\n- If bubbles are present, tap the tubing to move the bubbles up to the air trap.\n- If bubbles cannot be removed, treatment must end. | HD venous air bubble detected rinseback alarm | If bubble is detected in the treatment rinseback state. | 129 | ALARM_ID_HD_VENOUS_BUBBLE_DETECTED_RINSEBACK | ALARM_PRIORITY_HIGH | 10 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | HD venous air bubble detected rinseback alarm
|  |  | HD | Venous Pressure High | Venous pressure is high.\n- Check for access or patient discomfort.\n- Resolve access issue if applicable.\n- Check that there are no kinks or clots in the tubing.\n- Unkink the lines if applicable.\n- End treatment and do not return blood if clots are present. | Venous pressure too high during treatment | If the venous pressure > TREATMENT_PARAM_VEN_PRESSURE_HIGH_LIMIT. | 130 | ALARM_ID_VENOUS_PRESSURE_HIGH | ALARM_PRIORITY_HIGH | 11 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Venous pressure too high during treatment
|  |  | HD | Arterial Pressure Low | Arterial pressure is low.\n- Check that the access is properly connected and secured.\n- Resolve access issue if applicable.\n- Check to be sure the red arterial connections are secure.\n- Check that there are no kinks or clots in the tubing.\n- Unkink lines if applicable.\n- End treatment and do not return blood if clots are present. | Arterial pressure too low during treatment | If the arterial pressure < TREATMENT_PARAM_ART_PRESSURE_LOW_LIMIT. | 131 | ALARM_ID_ARTERIAL_PRESSURE_LOW | ALARM_PRIORITY_HIGH | 12 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Arterial pressure too low during treatment
|  |  | HD | Arterial Pressure High | Arterial pressure is high.\n- Check that the access is properly connected and secured.\n- Resolve access issue if applicable.\n- Check to be sure your red arterial connections are secure.\n- Check that there are no kinks or clots in the tubing.\n- Unkink lines if applicable.\n- End treatment and do not return blood if clots are present. | Arterial pressure too high during treatment | If the arterial pressure > TREATMENT_PARAM_ART_PRESSURE_HIGH_LIMIT. | 132 | ALARM_ID_ARTERIAL_PRESSURE_HIGH | ALARM_PRIORITY_HIGH | 13 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Arterial pressure too high during treatment
|  |  | DG | Fluid Leak in DG | Fluid leak detected in Dialysate Generation Delivery Device. | DG fluid leak detected alarm | If the fluid leak is detected as reported by the fluid leak detector. | 133 | ALARM_ID_DG_FLUID_LEAK_DETECTED | ALARM_PRIORITY_HIGH | 14 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | TRUE | TRUE | 0 | ALARM_ID_NO_ALARM | DG fluid leak detected alarm
|  |  | HD | Fluid Leak in HD | Fluid leak detected in Hemodialysis Delivery device.\n- Identify and correct the cause of the spill.\n- Clean up the spill with a clean paper towel.\n- Resume treatment. | HD fluid leak detected alarm | If the fluid leak is detected as reported by the fluid leak detector. | 134 | ALARM_ID_HD_FLUID_LEAK_DETECTED | ALARM_PRIORITY_HIGH | 15 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | HD fluid leak detected alarm
|  |  | HD | Excessive Acceleration Detected in the HD | Excessive movement detected with Hemodialysis Delivery device.\n- Confirm the device is stationary and level before continuing. | HD experienced large acceleration | If the system experienced excessive shock > 2.5 g. | 135 | ALARM_ID_HD_SHOCK | ALARM_PRIORITY_HIGH | 16 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD experienced large acceleration
|  |  | DG | Excessive Acceleration Detected in the DG | Excessive movement detected with Dialysis Generation device.\n- Confirm the device is stationary and level before continuing. | DG experienced large acceleration | If the system experienced excessive shock > 2.5 g. | 136 | ALARM_ID_DG_SHOCK | ALARM_PRIORITY_HIGH | 16 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG experienced large acceleration
|  |  | HD | Excessive Tilt Detected in the Hemodialysis Device | Excessive tilt detected with Hemodialysis Delivery device.\n- Confirm the device is stationary and level before continuing | HD tilt exceeds maximum | If the system experienced excessive tilt x and y axis angle > 7 deg for more than 1 sec. | 137 | ALARM_ID_HD_EXCESSIVE_TILT | ALARM_PRIORITY_HIGH | 17 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD tilt exceeds maximum
|  |  | DG | Excessive Tilt Detected in the Dialysate Generation Device | Excessive tilt detected with Dialysis Generation device.\n-Confirm the device is stationary and level before continuing. | DG tilt exceeds maximum | If the system experienced excessive tilt x and y axis angle > 7 deg for more than 1 sec. | 138 | ALARM_ID_DG_EXCESSIVE_TILT | ALARM_PRIORITY_HIGH | 17 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG tilt exceeds maximum
|  |  | HD | HD Power Lost | Power has been lost. Check all electrical connections.\n- The display and CPU will run off of the battery until power returns. | HD AC power lost alarm | The HD Software shall trigger an alarm if an AC power loss is detected while the HD system is in standby, service, or treatment parameters mode. | 139 | ALARM_ID_HD_AC_POWER_LOST | ALARM_PRIORITY_HIGH | 18 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | HD AC power lost alarm
|  |  | HD | Communication Failure with Dialysate Generator | Communication failure with Dialysis Generator.\n- Ensure all cords are securely connected. | Dialysate generator communication timeout | If DG communication timeout is detected for more than 2 seconds. | 140 | ALARM_ID_DG_COMM_TIMEOUT | ALARM_PRIORITY_HIGH | 19 | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Dialysate generator communication timeout
|  |  | HD | Air Trap Unable to Fill | Failure to fill the Air Trap.\n- Check the air trap vent line and the transducer protector.\n- If the transducer protector is wet, replace it. | Air trap fill timeout during treatment | If the air trap fill timeout during treatment. | 141 | ALARM_ID_AIR_TRAP_FILL_DURING_TREATMENT | ALARM_PRIORITY_HIGH | 20 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Air trap fill timeout during treatment
|  |  | HD | Blood Pump Occlusion Detected | Occlusion detected after the blood pump.\n- Check for clots or kinks around the dialyzer and in the venous line.\n- Resolve access issue if applicable.\n- End treatment and do not return blood if clots are present.\n- Check for access or patient discomfort.\n- Unkink the lines if applicable. | Blood pump occlusion detected | If blood occlusion is detected > OCCLUSION_THRESHOLD_OFFSET + bloodPumpOcclusionAfterCartridgeInstall | 142 | ALARM_ID_OCCLUSION_BLOOD_PUMP | ALARM_PRIORITY_HIGH | 21 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Blood pump occlusion detected
|  |  | DG | Dialysate Temperature Error | Call service to schedule a repair. | DG dialysate temperature sensors drift timeout. | If TDi and TRo have drifter for more 1 C for a certain period of time. | 143 | ALARM_ID_DG_DIALYSATE_TEMPERATURE_SENSORS_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | DG dialysate temperature sensors out of range
|  |  | DG | Cleaning Mode Inlet Water Conductivity Too High | Inlet water conductivity too high.\n- Check the incoming water for issues with cold water supply. | DG cleaning mode inlet water conductivity out of high range | If inlet water conductivity is > 2000 uS/cm. | 144 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_COND_TOO_HIGH | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG cleaning mode inlet water conductivity out of high range.
|  |  | DG | Acid Concentrate Conductivity Out of Range | Dialysate conductivity is out of range.\n- Check correct acid concentrate is connected. | DG Concentrate conductivity after adding acid out of range alarm | If the conductivity value for the acid concentrate is out of range for more than 45 secs. | 145 | ALARM_ID_ACID_CONDUCTIVITY_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Concentrate conductivity after adding acid out of range alarm
|  |  | DG | Self-Test Error | Call service to schedule a repair. | DG RTC vs. Timer accuracy error | If RTC and timer are not aligned on 1 second elapsed time. | 146 | ALARM_ID_DG_RTC_OR_TIMER_ACCURACY_FAILURE | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | DG RTC or timer accuracy failure.
|  |  | DG | Dialysate Conductivity Error | Call service to schedule a repair. | DG dialysate generation conductivity fault alarm | If dialysate conductivity > 14850 or < 13000. | 147 | ALARM_ID_CREATING_DIALYSATE_PLEASE_WAIT | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Wait for the DG to produce dialysate
|  |  | DG | Cleaning Mode Inlet Water Temperature Too High | Cleaning Mode Inlet water temperature too high.\n- Check the incoming water for issues with cold water supply. | Cleaning Mode Inlet water temperature too high | Inlet temperature > 37 C. | 148 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_TEMP_TOO_HIGH | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG Cleaning mode inlet water temperature too high
|  |  | HD | Watchdog Error | Call service to schedule a repair. | HD watchdog expired | If the watchdog signal is low while not in POST. | 149 | ALARM_ID_HD_WATCHDOG_EXPIRED | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD watchdog expired
|  |  | DG | Inlet Water Conductivity High | Inlet water conductivity is high.\n- Check the incoming water for issues. | Inlet water conductivity too high | Inlet water conductivity > 2000 uS/cm. | 150 | ALARM_ID_INLET_WATER_CONDUCTIVITY_IN_HIGH_RANGE | ALARM_PRIORITY_MEDIUM | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Inlet water conductivity in high range
|  |  | DG | Inlet Water Conductivity Low | Inlet water conductivity is low.\n- Check the incoming water for issues. | Inlet water conductivity too low | Inlet water conductivity < 200 uS/cm. | 151 | ALARM_ID_INLET_WATER_CONDUCTIVITY_IN_LOW_RANGE | ALARM_PRIORITY_MEDIUM | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Inlet water conductivity in low range
|  |  | DG | Inlet Water Pressure Low | Inlet water pressure is low.\n- Check the incoming water supply. | Inlet water pressure too low | Inlet water pressure < 25 PSI. | 152 | ALARM_ID_INLET_WATER_PRESSURE_IN_LOW_RANGE | ALARM_PRIORITY_MEDIUM | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Inlet water pressure in low range
|  |  | HD | Prime Complete | Prime of the bloodlines is complete. | HD prime completed high priority alarm | If alarm 213 is not cleared by user within 30 minutes. | 153 | ALARM_ID_PRIME_COMPLETED_HIGH | ALARM_PRIORITY_HIGH | 36 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD prime completed high priority alarm
|  |  | DG | Cleaning Mode Inlet Water Conductivity Low | Cleaning Mode Inlet water conductivity is low.\n- Check the incoming water. | Cleaning mode inlet water conductivity too low | If inlet water conductivity is < 200 uS/cm in cleaning mode. | 154 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_COND_TOO_LOW | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG cleaning mode inlet water conductivity out of low range.
|  |  | HD | No reservoir data received | Call service to schedule a repair. | HD No reservoir data message receive at least once every 2 seconds | No new reservoir data for 2 seconds. | 155 | ALARM_ID_HD_NEW_RESERVOIRS_DATA_MESSAGE_NOT_RECEIVE | ALARM_PRIORITY_LOW | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD No reservoirs data message receive at least once every 2 seconds
|  |  | HD | No DG op mode received | Call service to schedule a repair. | HD No DG op mode message receive at least once every 2 seconds | No DG op mode for 2 seconds. | 156 | ALARM_ID_HD_NEW_DG_OPERATION_MODE_MESSAGE_NOT_RECEIVE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD No DG operation mode message receive at least once every 2 seconds
|  |  | DG | Chemical Disinfection Error | Ensure the acid is connected to the bicarb line | Chemical disinfect priming the acid line timeout | If the disinfectant acid line is not primed within a certain period of time. | 157 | ALARM_ID_DG_CHEM_DISINFECT_PRIME_ACID_LINE_TIME_OUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG chemical disinfect prime acid line timeout.
|  |  | DG | Inlet water temperature low | Inlet water temperature is low.\n- Check the incoming water supply. | Inlet water temperature low | Inlet temperature < 22 °C for 5 seconds. | 158 | ALARM_ID_INLET_WATER_TEMPERATURE_IN_LOW_RANGE | ALARM_PRIORITY_MEDIUM | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Inlet water temperature is in the low range
|  |  | DG | Chemical Disinfect Flush | Remove Acid Bottle and Close Concentrate Cap. | Close concentrate line be able to flush it | Start of chemical disinfect flush. | 159 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_REMOVE_ACID | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG chemical disinfect flush remove acid and close the concentrate cap.
|  |  | HD | Blood Pump Speed High | Blood pump failure.\n- Check the pump rotor for secure connection. | Blood pump rotor speed too high | If blood pump motor speed exceeded 100.0 RPM. | 160 | ALARM_ID_BLOOD_PUMP_ROTOR_SPEED_TOO_HIGH | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood pump rotor speed too high
|  |  | DG | Machine Error | Call service to schedule a repair. | Too many bad CRCs detected on received system messages | If >= 5 bad CRCs detected on received CAN messages within 10 minutes. | 161 | ALARM_ID_DG_COMM_TOO_MANY_BAD_CRCS | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG comm too many bad CRCs
|  |  | DG | FPGA Error | Call service to schedule a repair. | FPGA clock speed failure | If the FPGA clock speed is relatively close to the processor clock for a certain period of time. | 162 | ALARM_ID_DG_FPGA_CLOCK_SPEED_CHECK_FAILURE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG FPGA clock speed check failure.
|  |  | HD | Ultrafiltration Measurement Error | Ultrafiltration error. | HD sees primary load cell for reservoir 1 change too much too fast | If the delta volume of the active reservoir exceeded 50.0 g (MAX_ACTIVE_LOAD_CELL_CHANGE_G). | 163 | ALARM_ID_HD_LOAD_CELL_ACCELERATION_RES_1_ALARM | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD sees primary load cell for reservoir 1 change too much too fast
|  |  | HD | Ultrafiltration Measurement Error | Ultrafiltration error. | HD sees primary load cell for reservoir 2 change too much too fast | If the delta volume of the active reservoir exceeded 50.0 g (MAX_ACTIVE_LOAD_CELL_CHANGE_G). | 164 | ALARM_ID_HD_LOAD_CELL_ACCELERATION_RES_2_ALARM | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD sees primary load cell for reservoir 2 change too much too fast
|  |  | HD | Rinseback Timeout | Maximum Rinseback time exceeded. Rinseback must end. | HD in treatment rinseback operation timeout | If the maximum time inrinseback operation has been exceeded. | 165 | ALARM_ID_TREATMENT_RINSEBACK_TIMEOUT_ALARM | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD in treatment rinseback operation timeout
|  |  | HD | Recirculation Timeout | Maximum Recirculation time exceeded. Treatment must end. | HD in treatment re-circ sub-mode for too long | If the maximum time in recirculate sub-mode has been exceeded. | 166 | ALARM_ID_TREATMENT_RECIRC_TIMEOUT_ALARM | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD in treatment re-circ sub-mode for too long
|  |  | HD | Door Open | Hemodialysis Delivery device door open.\n - Close open HD door\n - Press RESUME button | HD cartridge door opened alarm | If the cartridge door is opened as reported by the FPGA. | 167 | ALARM_ID_CARTRIDGE_DOOR_OPENED | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD cartridge door opened alarm
|  |  | DG | Dialysate Recirculation Error | Dialysate recirculation error. | HD active reservoir depletion timeout. | If the active reservoir's depletion time has elapsed. | 168 | ALARM_ID_HD_ACTIVE_RESERVOIR_DEPLETION_TIME_OUT | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD active reservoir depletion timeout.
|  |  | DG | Dialysate Generation Error | Call service to schedule a repair. | Dialysate flow rate out of range | If the dialysate flow rate > 2.00 L/min. | 169 | ALARM_ID_DG_DIALYSATE_FLOW_RATE_OUT_OF_MAX_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG Dialysate flow rate out of maximum range
|  |  | HD | Syringe Empty | Heparin syringe empty.\n- Option to resume treatment without heparin. | HD syringe empty alarm | If the syringe pump position is near empty position. | 170 | ALARM_ID_HD_SYRINGE_PUMP_SYRINGE_EMPTY | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe empty alarm
|  |  | HD | Heparin Occlusion | Heparin pump occlusion detected. Check heparin line is unclamped and there are no kinks in the heparin line. | HD syringe pump occlusion alarm | If the pressure sensor voltage reading > SYRINGE_FORCE_OCCLUSION_THRESHOLD_V. | 171 | ALARM_ID_HD_SYRINGE_PUMP_OCCLUSION | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump occlusion alarm
|  |  | HD | Not Enough Heparin | Not enough heparin is available for treatment. Check if syringe pump is touching syringe plunger. If syringe pump is touching syringe plunger, end treatment. If syringe pump is not touching syringe plunger, reload syringe with enough heparin. | HD syringe pump not enough Heparin alarm | If the estimated syringe volume needed for treatment is not sufficient | 172 | ALARM_ID_HD_SYRINGE_PUMP_NOT_ENOUGH_HEPARIN_ALARM | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump not enough Heparin alarm
|  |  | HD | RTC Error | Call service to schedule a repair. | HD RTC configuration error | If the RTC configuration registers report an error. | 173 | ALARM_ID_HD_RTC_CONFIG_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD RTC configuration error
|  |  | HD | Self-Test Error | Call service to schedule a repair. | HD RTC vs. Timer accuracy error | If RTC and timer are not aligned on 1 second elapsed time. | 174 | ALARM_ID_HD_RTC_OR_TIMER_ACCURACY_FAILURE | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD RTC or timer accuracy failure
|  |  | HD | Blood Pump Error | Blood pump failure. \n- Check the pump rotor for secure connection. | HD pump direction status error | If the set pump direction does not agree with the hall sensor direction. | 175 | ALARM_ID_HD_PUMP_DIRECTION_STATUS_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD pump direction status error
|  |  | HD | Dialysate Temperature Error | Call service to schedule a repair. | Dialysate temperature is out of high safety range | If the dialysate temperature is > 46 C for a certain period of time. | 176 | ALARM_ID_HD_DIALYSATE_TEMP_OUT_OF_HIGH_SAFETY_RANGE | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD dialysate temperature out of high safety range
|  |  | DG | Machine Error | Call service to schedule a repair. | DG software fault. Software found itself in an unexpected state | If the specific software faults listed in alarm management SW faults has been detected | 177 | ALARM_ID_DG_SOFTWARE_FAULT | ALARM_PRIORITY_HIGH | 799 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG software fault. Software found itself in an unexpected state
|  |  | DG | Communication Failure | Call service to schedule a repair. | HD communication timeout | If the HD stops communicating with the UI. | 178 | ALARM_ID_HD_COMM_TIMEOUT | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD communication timeout
|  |  | DG | Communication Failure | Call service to schedule a repair. | DG FPGA communication down for too long | If retries for commands exceeds limit or FPGA reports comm error. | 179 | ALARM_ID_DG_FPGA_COMM_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG FPGA communication down for too long
|  |  | DG | RO Flow Error | Call service to schedule a repair. | RO flow rate out of range | If the RO flow rate > 2.00 L/min. | 180 | ALARM_ID_DG_RO_FLOW_RATE_OUT_OF_MAX_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG RO flow out of maximum range
|  |  | DG | Ultrafiltration Measurement Error | Call service to schedule a repair. | DG load cells weight out of range for tare | If the load cell weight is out of range for tare. | 181 | ALARM_ID_DG_LOAD_CELLS_TARE_WEIGHT_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG load cells weight out of range for tare
|  |  | DG | Ultrafiltration Measurement Error | Call service to schedule a repair. | DG load cells invalid calibration | If the CRC of load cells calibration data does not match the calculated CRC. | 182 | ALARM_ID_DG_LOAD_CELLS_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG load cells invalid calibration
|  |  | DG | Heat Disinfection Error | Call service to schedule a repair. | DG invalid load cell value | If the half full reservoir is filled prior to filling the reservoir that is supposed to overflow to the half full reservoir. | 183 | ALARM_ID_DG_INVALID_LOAD_CELL_VALUE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG invalid load cell value
|  |  | DG | UV Reactor Error | Call service to schedule a repair. | Inlet or Outlet UV reactor not healthy | If the inlet or outlet reactor are not healthy. | 184 | ALARM_ID_UV_REACTOR_NOT_HEALTHY | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Inlet or Outlet UV reactor not healthy
|  |  | DG | Fan Error | Fan failure.\n- Check that the fans inlets are not blocked or obstructed. | DG fan RPM out of range | If the fan RPM is out of Min/Max range. | 185 | ALARM_ID_DG_FAN_RPM_OUT_OF_RANGE | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | DG fan RPM out of range
|  |  | DG | Dialysate Concentrate Pump Error | Call service to schedule a repair. | DG concentrate pump fault | If the concentrate pumps fault is reported by the FPGA. | 186 | ALARM_ID_DG_CONCENTRATE_PUMP_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG concentrate pump fault
|  |  | DG | Dialysate Concentrate Pump Error | Call service to schedule a repair. | DG concentrate pump CP1 speed control error | If the concentrate pump CP1 speed > 0.02 tolerance for a certain period of time. | 187 | ALARM_ID_CP1_SPEED_CONTROL_ERROR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG concentrate pump CP1 speed control error
|  |  | DG | Dialysate Concentrate Pump Error | Call service to schedule a repair. | DG concentrate pump CP2 speed control error | If the concentrate pump CP1 speed > 0.02 tolerance for a certain period of time. | 188 | ALARM_ID_CP2_SPEED_CONTROL_ERROR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG concentrate pump CP2 speed control error
|  |  | DG | Drain Pump Error | Call service to schedule a repair. | DG drain pump RPM out of range | If the drain pump RPM > 100.0 of the target RPM. | 189 | ALARM_ID_DRAIN_PUMP_RPM_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG drain pump RPM out of range
|  |  | DG | Drain Pump Error | Call service to schedule a repair. | DG drain pump off fault | If the pump is in off state and the RPM is greater than the minimum RPM of 300.0 RPM. | 190 | ALARM_ID_DRAIN_PUMP_OFF_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG drain pump off fault
|  |  | DG | RO Flow Error | Call service to schedule a repair. | DG flow rate out of upper range | If the RO flow rate > 10% from the target flow rate. | 191 | ALARM_ID_FLOW_RATE_OUT_OF_UPPER_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG flow rate out of upper range
|  |  | DG | RO Flow Error | Call service to schedule a repair. | DG flow rate out of lower range | If the RO flow rate < 10% from the target flow rate. | 192 | ALARM_ID_FLOW_RATE_OUT_OF_LOWER_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG flow rate out of lower range
|  |  | HD | Dialysate Temperature Error | Call service to schedule a repair. | Dialysate temperature is out of low safety range | If the dialysate temepertaure is < 42 C for a certain period of time. | 193 | ALARM_ID_HD_DIALYSATE_TEMP_OUT_OF_LOW_SAFETY_RANGE | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD dialysate temperature out of low safety range.
|  |  | DG | RO Pump Error | Call service to schedule a repair. | DG RO pump duty cycle out of range | If the DG RO pump duty cycle is out of range by 5% of the target duty cycle for a certain period of time. | 194 | ALARM_ID_RO_PUMP_DUTY_CYCLE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG RO pump duty cycle out of range
|  |  | DG | RO Pump Error | Call service to schedule a repair. | DG RO pump pressure out of range | If the DG RO pump pressure > 135 psi for a certain period of time. | 195 | ALARM_ID_RO_PUMP_PRESSURE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 799 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG RO pump pressure out of range
|  |  | DG | FPGA Error | Call service to schedule a repair. | DG CPi/CPo conductivity sensors FPGA fault | If the FPGA reports error for a certain period of time. | 196 | ALARM_ID_DG_CPI_CPO_SENSORS_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG CPi/CPo sensors FPGA fault
|  |  | DG | FPGA Error | Call service to schedule a repair. | DG CD1/CD2 conductivity sensors FPGA fault | If the FPGA reports error for a certain period of time. | 197 | ALARM_ID_DG_CD1_CD2_SENSORS_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG CD1/CD2 sensors FPGA fault
|  |  | DG | Heaters Error | Call service to schedule a repair. | DG primary heater while RO flow is low | If the RO flow is < 0.2 L/min for a certain period of time. | 198 | ALARM_ID_RO_FLOW_TOO_LOW_WHILE_PRIMARY_HEATER_IS_ON | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | RO flow too low while primary heater is on
|  |  | DG | Heaters Error | Call service to schedule a repair. | DG trimmer heater while dialysate flow is low | If the RO flow is < 50 mL/min for a certain period of time. | 199 | ALARM_ID_DIALYSATE_FLOW_TOO_LOW_WHILE_TRIMMER_HEATER_IS_ON | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysate flow too low while trimmer heater is on
|  |  | DG | Thermistors Error | Call service to schedule a repair. | DG thermistor or temperature sensor out of range | If thermistor > 80.0 C or < 0.0 C. | 200 | ALARM_ID_DG_THERMISTORS_TEMPERATURE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG thermistors/sensors temperature out of range
|  |  | HD | Pre-Treatment Error | Call service to schedule a repair. | HD pre-treatment wet flow test failure | If the pre-treatment wet flow test has failed. | 201 | ALARM_ID_PRE_TREATMENT_WET_FLOW_TEST_FAILURE | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD pre-treatment mode wet self-test lc vs FMD failure
|  |  | HD | Pre-Treatment Error | Call service to schedule a repair. | HD pre-treatment dry pressure normal test failure | If the pre-treatment dry pressure normla test failed. | 202 | ALARM_ID_PRE_TREATMENT_DRY_PRESSURE_NORMAL_TEST_FAILURE | ALARM_PRIORITY_LOW | 910 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD pre-treatment mode dry pressure normal self-test failure
|  |  | DG | Pressure Sensors Error | Call service to schedule a repair. | DG pressure sensors fault | If pressure sensor ADC reading = 0 or inlet pressure ADC reading > 4095 (12-bit full scale ). | 203 | ALARM_ID_PRESSURE_SENSOR_FAULT | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Pressure sensor fault
|  |  | DG | RO Rejection Error | Call service to schedule a repair. | RO rejection ratio out of range | If RO rejection ratio > 0.1 and outlet water conductivity > 15.0 uS/cm. | 204 | ALARM_ID_RO_REJECTION_RATIO_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | RO rejection ratio out of range
|  |  | DG | Conductivity Sensors Error | Call service to schedule a repair. | Conductivity sensor fault | If conductivity sensor error count reported by the FPGA > 0 for more than 5 times. | 205 | ALARM_ID_CONDUCTIVITY_SENSOR_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Conductivity sensor fault
|  |  | DG | Dialysate Production Error | Call service to schedule a repair. | DG dialysate fill runs out of time | If DG dialysate fill > 5 minutes. | 206 | ALARM_ID_DG_DIALYSATE_FILL_OUT_OF_TIME | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG dialysate fill runs out of time
|  |  | DG | Dialysate Flowrate Error | Call service to schedule a repair. | DG flow meter check failure alarm | If integrated Volume To Load Cell Reading Percent > 0.1. | 207 | ALARM_ID_DG_FLOW_METER_CHECK_FAILURE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG flow meter check failure alarm
|  |  | HD | Venous Occlusion Error | Call service to schedule a repair. | HD venous line occlusion fault | If the venous pressure is > 50 mmHg if HD is in treatment or pre-treatment mode. | 208 | ALARM_ID_HD_VENOUS_LINE_OCCLUSION | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD venous pressure occlusion alarm
|  |  | DG | Flush Mode Error | Call service to schedule a repair. | DG drain circulation line timeout | If the integrated volume < 0.1 for more than 60 secs. | 209 | ALARM_ID_DG_DRAIN_CIRCULATION_LINE_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG drain circulation line timeout
|  |  | HD | Battery Error | Battery pack error. | HD battery pack detected an error | If the battery pack status error is detected from the slave device I2C address 0x0B. | 210 | ALARM_ID_HD_BATTERY_PACK_ERROR_DETECTED | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD battery pack detected an error
|  |  | HD | Blood Stagnant | Maximum time for blood stoppage has been exceeded. | Blood sitting too long warning (> 4 min) | During treatment stop sub-mode, if the blood is in the line for more than 4 minutes. | 211 | ALARM_ID_BLOOD_SITTING_WARNING | ALARM_PRIORITY_MEDIUM | 800 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood sitting too long warning (>4 min)
|  |  | HD | Blood Stagnant | Maximum time for blood stoppage has been exceeded. | Blood sitting too long warning (> 4 min) | During treatment end paused state, if the blood is in the line for more than 4 minutes. | 212 | ALARM_ID_END_OF_TREATMENT_ALARM | ALARM_PRIORITY_MEDIUM | 801 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD end of treatment alarm (user not acting to end treatment)
|  |  | HD | Prime Complete | Prime of the bloodlines is complete. | HD prime completed medium priority alarm | If alarm 226 is not cleared by user within 90 minutes. | 213 | ALARM_ID_PRIME_COMPLETED_MEDIUM | ALARM_PRIORITY_MEDIUM | 802 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | HD prime completed medium priority alarm
|  |  | HD | Saline Bolus Error | Inability to deliver saline bolus.\n- Check for kinks or clamped lines in the saline line. | HD Saline bolus volume check failure | If the saline bolus is under or over delivering. | 214 | ALARM_ID_SALINE_BOLUS_VOLUME_CHECK_FAILURE | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD Saline bolus volume check failure
|  |  | HD | Rinseback Error | Unable to complete Rinseback.\n- Check for kinks or clamped blood lines. | HD rinseback volume check failure | If the delta between the expected rinseback volume and the measured rinseback volume > 60.0 mL. | 215 | ALARM_ID_RINSEBACK_VOLUME_CHECK_FAILURE | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD rinseback volume check failure
|  |  | HD | Treatment Ended | Treatment has ended | HD end treatment sub-mode timeout alarm | If the treatment ended successfully. | 216 | ALARM_ID_END_TREATMENT_TIMEOUT_ALARM | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD end treatment sub-mode timeout alarm
|  |  | HD | Prime Error | Blood prime error. Check for kinks or clamped blood lines. | HD blood prime volume check failure | If expected blood prime volume - blood prime volume > 60.0 mL. | 217 | ALARM_ID_BLOOD_PRIME_VOLUME_CHECK_FAILURE | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood prime volume check failure
|  |  | HD | Syringe Detected | Remove Heparin syringe. | HD syringe detected alarm | If the syringe detection switch > 2.0 volts. | 218 | ALARM_ID_HD_SYRINGE_DETECTED | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe detected alarm
|  |  | HD | Syringe Removed | Install Heparin syringe. | HD syringe pump syringe removed alarm | If the syringe detection switch < 2.0 volts. | 219 | ALARM_ID_HD_SYRINGE_PUMP_SYRINGE_REMOVED | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump syringe removed alarm
|  |  | DG | FPGA Error | Call service to schedule a repair. | DG THd sensor FPGA fault | If the FPGA does not report fresh data or indicates an error for a certain period of time. | 220 | ALARM_ID_DG_THD_SENSORS_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG THd sensors FPGA fault
|  |  | HD | Saline Bag Empty | Saline back appears empty.\n- Check saline bag and line and replace if saline bag is empty. | HD Saline bag is empty | If the measured arterial pressure < -300.0 mmHg for more than 250 msec. | 221 | ALARM_ID_EMPTY_SALINE_BAG | ALARM_PRIORITY_LOW | 900 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD Saline bag is empty
|  |  | HD | FPGA Error | Call service to schedule a repair. | HD occlusion sensor FPGA fault | If the FPGA does not report fresh data or indicates an error for a certain period of time. | 222 | ALARM_ID_HD_OCCLUSION_SENSOR_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD occlusion sensor FPGA fault
|  |  | HD | FPGA Error | Call service to schedule a repair. | HD arterial sensor FPGA fault | If the FPGA does not report fresh data or indicates an error for a certain period of time. | 223 | ALARM_ID_HD_ARTERIAL_SENSOR_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD arterial sensor FPGA fault
|  |  | HD | Treatment Stopped | Treatment paused. | Treatment stopped by user action - pressed stop button | If the user has pressed the stop button. | 224 | ALARM_ID_TREATMENT_STOPPED_BY_USER | ALARM_PRIORITY_MEDIUM | 803 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Treatment stopped by user action - pressed stop button
|  |  | HD | End of Treatment | Treatment has ended | HD end of treatment warning | If the treatment duration has exceeded the prescribed treatment time entered by the user. | 225 | ALARM_ID_END_OF_TREATMENT_WARNING | ALARM_PRIORITY_LOW | 904 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD end of treatment warning
|  |  | HD | Prime Complete | Prime of the bloodlines is complete. | HD prime completed low priority alarm | If the submodeCompleteTransitionTimeCounter > 20. | 226 | ALARM_ID_PRIME_COMPLETED_LOW_PRIORITY | ALARM_PRIORITY_LOW | 905 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | HD prime completed low priority alarm
|  |  | HD | Maximum Prime Time Exceeded | Maximum time for blood prime has been exceeded. | HD prime out of time alarm | If the priming elapsed time exceeded 10 minutes. | 227 | ALARM_ID_PRIME_OUT_OF_TIME | ALARM_PRIORITY_LOW | 909 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD prime out of time alarm
|  |  | HD | Maximum Prime Time Exceeded | Priming failure.\n- Check blood and dialysate lines for clamps and kinks. | HD time out on prime sub-mode purge air state. | During the pre-treatment prime state, if the prime saline purge air state exceeded 60 secs. | 228 | ALARM_ID_HD_PRIME_SALINE_PURGE_AIR_TIME_OUT | ALARM_PRIORITY_LOW | 909 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD time out on prime saline purge air state.
|  |  | HD | Maximum Prime Time Exceeded | Priming failure.\n- Check blood and dialysate lines for clamps and kinks. | HD prime dialysate dialyzer time out alarm | If the priming of the dialyzer dialysate flow path has exceeded the time limit. | 229 | ALARM_ID_PRIME_DIALYSATE_DIALYZER_TIME_OUT | ALARM_PRIORITY_LOW | 909 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD prime dialysate dialyzer time out alarm
|  |  | HD | Maximum Prime Time Exceeded | Priming failure.\n- Check blood and dialysate lines for clamps and kinks. | HD prime dialysate bypass time out alarm | If the priming of the dialyzer bypass flow path has exceeded the time limit. | 230 | ALARM_ID_PRIME_DIALYSATE_BYPASS_TIME_OUT | ALARM_PRIORITY_LOW | 909 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD prime dialysate bypass time out alarm
|  |  | HD | Self Test Error | Self Test Failure. Please ensure the disposable cartridge is completely installed and the handle moved to the down position. | HD pre-treatment mode dry self-tests failure | If the following dry self tests have failed: 1. handleDrySelfTestPressureSensorsState 2. handleDrySelfTestPressureSensorsNormalState | 231 | ALARM_ID_PRE_TREATMENT_DRY_PRESSURE_TEST_FAILURE | ALARM_PRIORITY_LOW | 910 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD pre-treatment mode dry pressure self-test failure
|  |  | HD | Self Test Error | Self Test Failure. Please ensure the disposable cartridge is completely installed and the handle moved to the down position. Check for closed clamps or kinks in the tubing. | HD pre-treatment mode wet self-tests failure | If the following wet self tests have failed: 1. handleWetSelfTestPrimeCheckState 2. handleWetSelfTestFirstDisplacementVerifyState 3. handleWetSelfTestSecondDisplacementVerifyState | 232 | ALARM_ID_PRE_TREATMENT_WET_LC_TEST_FAILURE | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD pre-treatment mode wet self-test LC vs. LC failure
|  |  | DG | Pressure Sensors Error | Call service to schedule a repair. | DG barometric pressure sensor bad coefficients CRC | If the following happens: 1. The stored CRC in the sensor does not match the calculated CRC. 2. The CRC and coefficients have not been received from the sensor after 20 seconds from the power up. | 233 | ALARM_ID_DG_BAROMETRIC_SENSOR_COEFFS_BAD_CRC | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG barometric sensor coefficients bad CRC
|  |  | DG | Inlet Water Pressure High | Inlet water pressure is high.\n- Check the incoming water for issues. | DG inlet water pressure is too high | If the inlet water pressure is > 80 psig. | 234 | ALARM_ID_INLET_WATER_PRESSURE_IN_HIGH_RANGE | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | DG inlet water pressure in high range.
|  |  | HD | Rinseback Complete | Rinseback has been completed. | HD in treatment stopped sub-mode after rinseback completed (no escalation) | If there is no active alarm and rinseback has completed. | 235 | ALARM_ID_TREATMENT_STOPPED_AFTER_RINSEBACK | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD in treatment stopped sub-mode after rinseback completed (no escalation)
|  |  | HD | New Cartridge Needed | New cartridge needs to be installed. | HD needs new cartridge to be installed | If the air trap sensor reported AIR and bubble sensor detected bubbles. | 236 | ALARM_ID_INSTALL_NEW_CARTRIDGE | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD needs new cartridge to be installed
|  |  | HD | Prime Error | Call service to schedule a repair. | HD prime saline dialyzer time out alarm | If timeout, just transition to HD_PRIME_RESERVOIR_TWO_FILL_COMPLETE_STATE | 237 | ALARM_ID_PRIME_SALINE_DIALYZER_TIME_OUT | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD prime saline dialyzer time out alarm
|  |  | HD | Cartridge Needed | Cartridge not loaded or loaded improperly.\n- Check that a cartridge is properly installed. | HD no cartridge loaded or installed improperly alarm | If the measured blood pump occlucsion pressure reading is < 5000 or > 24500. | 238 | ALARM_ID_NO_CARTRIDGE_LOADED | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD no cartridge loaded or installed improperly alarm
|  |  | HD | Cartridge Removal Error | Cartridge not unloaded, or unloaded improperly.\n- Check that the cartridge has been removed. | HD fail to remove cartridge alarm | In the post treatment state, if the cartridge has not been removed. | 239 | ALARM_ID_CARTRIDGE_REMOVAL_FAILURE | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD fail to remove cartridge alarm
|  |  | DG | Dialysate Conductivity Error | Dialysate conductivity out of range.\n- Check to be sure the correct bicarbonate dialysate concentrate is connected. | Bicarb conductivity out of range during bicarb pump check alarm | If the conductivity value for bicarbonate concentrate is out of range for more than 45 secs. | 240 | ALARM_ID_BICARB_CONDUCTIVITY_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Bicarb conductivity out of range during bicarb pump check alarm
|  |  | DG | Cleaning Mode Error | Call service to schedule a repair. | DG reservoir drain time out | If the DG reservoir drain is not complete in 120 secs. | 241 | ALARM_ID_DG_RESERVOIR_DRAIN_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG reservoir drain time out
|  |  | DG | Cleaning Mode Error | Call service to schedule a repair. | DG reservoir fill time out | If the DG reservoir fill is not complete in 120 secs. | 242 | ALARM_ID_DG_RESERVOIR_FILL_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG reservoir fill time out
|  |  | DG | Heat Disinfection Error | Call service to schedule a repair. | DG reservoir leak time out | If the reservoir level is drifted by 100 mL for more than 30 seconds. | 243 | ALARM_ID_DG_RESERVOIR_LEAK_TIMEOUT | ALARM_PRIORITY_HIGH | 799 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG reservoir leak time out
|  |  | DG | Cleaning Mode Error | Call service to schedule a repair. | DG temperature sensors difference out of range | If any of the temperature sensors are out of the average of the sensors by more than 3 C. | 244 | ALARM_ID_DG_CLEANING_MODE_TEMP_SENSORS_DIFF_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG cleaning mode temperature sensors difference out of range
|  |  | DG | Heat Disinfection Error | Call service to schedule a repair. | DG heat disinfect target temperature time out (could not reach to temperature) | If heating up to minimum temperature for heat disinfect failed after 4 hours. | 245 | ALARM_ID_DG_HEAT_DISINFECT_TARGET_TEMP_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG heat disinfect target temperature time out (could not reach to temperature)
|  |  | DG | Chemical Disinfection Error | Call service to schedule a repair. | DG chemical disinfect inlet conducitivty out of range | If CPo and CD2 conductivity values are above 100 uS/cm. | 246 | ALARM_ID_DG_CLEANING_MODE_COND_SENSORS_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG cleaning mode conductivity sensors out of range
|  |  | DG | Heat Disinfection Active Cool Error | Call service to schedule a repair. | DG heat disinfect active cool time out | If THd > 44 C after 90 minutes of cooling. | 247 | ALARM_ID_DG_HEAT_DISINFECT_ACTIVE_COOL_TIME_OUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG heat disinfect active cool time tour
|  |  | DG | Chemical Disinfection Error | Call service to schedule a repair. | DG chemical disinfect out of target temperature | If TPo > 50 C after 1 second. | 248 | ALARM_ID_DG_CHEM_DISINFECT_TARGET_TEMP_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG chemical disinfect could not reach to target temperature out of range
|  |  | DG | Chemical Disinfection Error | Call service to schedule a repair. | DG chemical disinfect out of target conductivity | If CD2 < 150 uS/cm or CD2 > 650 uS/cm after 10 seconds. | 249 | ALARM_ID_DG_CHEM_DISINFECT_TARGET_COND_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG chemical disinfect target conductivity out of range
|  |  | DG | Chemical Disinfect | Insert the Acid Bottle | Insert the acid to start chemical disinfect | Start of chemical disinfect | 250 | ALARM_ID_DG_CHEM_DISINFECT_INSERT_ACID | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG chemical disinfect insert acid and remove the concentrate cap
|  |  | HD | Machine Error | Call service to schedule a repair. | HD invalid system record | If the calculated CRC of the HD system record image does not match the expected CRC. | 251 | ALARM_ID_HD_INVALID_SYSTEM_RECORD_CRC | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD invalid system record
|  |  | HD | Machine Error | Call service to schedule a repair. | HD invalid service record | If the calculated CRC of the HD service record image does not match the expected CRC. | 252 | ALARM_ID_HD_INVALID_SERVICE_RECORD_CRC | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD invalid service record
|  |  | DG | Machine Error | Call service to schedule a repair. | DG invalid system record | If the calculated CRC of the DG system record image does not match the expected CRC. | 253 | ALARM_ID_DG_INVALID_SYSTEM_RECORD_CRC | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG invalid system record
|  |  | DG | Machine Error | Call service to schedule a repair. | DG invalid service record | If the calculated CRC of the DG service record image does not match the expected CRC. | 254 | ALARM_ID_DG_INVALID_SERVICE_RECORD_CRC | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG invalid service record
|  |  | HD | Machine Error | Call service to schedule a repair. | HD and UI software builds are not compatible | If the compatibility revision field of the UI is not equal to SW_COMPATIBILITY_REV. | 255 | ALARM_ID_HD_UI_COMPATIBILITY_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD and UI software builds are not compatible
|  |  | N/A | N/A | N/A | N/A | N/A | 256 | ALARM_ID_AVAILABLE_2 | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Alarm ID available for use
|  |  | DG | Machine Error | Call service to schedule a repair. | DG FPGA power out timeout | If the FPGA power gpio status pin has a logic '0' for more than 2 seconds. | 257 | ALARM_ID_DG_FPGA_POWER_OUT_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG FPGA power out timeout
|  |  | HD | Machine Error | Call service to schedule a repair. | HD FPGA power out timeout | If the FPGA power gpio status pin has a logic '0' for more than 2 seconds. | 258 | ALARM_ID_HD_FPGA_POWER_OUT_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD FPGA power out timeout
|  |  | HD | Temperature Sensors Error | Call service to schedule a repair. | HD temperatures out of range | If thermistor reports temperature > 80.0 C or < 0.0 C for a certain period of time. | 259 | ALARM_ID_HD_TEMPERATURES_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD temperatures out of range
|  |  | UI | Machine Error | Call service to schedule a repair. | UI POST Application Integrity (Sha256Sum) failure | If the UI POST reports a shasum faillure. | 260 | ALARM_ID_UI_POST_FAILURE_SHASUM | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST Application Integrity (Sha256Sum) failure
|  |  | UI | Machine Error | Call service to schedule a repair. | UI POST CANBus failure | If the UI POST reports a can bus failure. | 261 | ALARM_ID_UI_POST_FAILURE_CANBUS | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST CANBus failure
|  |  | UI | Machine Error | Call service to schedule a repair. | UI POST Display failure | If the UI POST reports a display failure. | 262 | ALARM_ID_UI_POST_FAILURE_DISPLAY | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST Display failure
|  |  | UI | Machine Error | Call service to schedule a repair. | UI POST Touch failure | If the UI POST reports a touch failure. | 263 | ALARM_ID_UI_POST_FAILURE_TOUCH | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST Touch failure
|  |  | UI | Machine Error | SD-Card Failure \n- Power cycle the system.\n- Call service to schedule a repair if problem persist. | UI POST SD-Card failure | If the UI POST reports SD card failure. | 264 | ALARM_ID_UI_POST_FAILURE_SDCARD | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST SD-Card failure
|  |  | UI | Machine Error | Call service to schedule a repair. | UI POST RTC failure | If the UI POST reports RTC failure. | 265 | ALARM_ID_UI_POST_FAILURE_RTC | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST RTC failure
|  |  | UI | Machine Error | Call service to schedule a repair. | UI POST WiFi failure | If the UI POST reports WiFi failure. | 266 | ALARM_ID_UI_POST_FAILURE_WIFI | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST WiFi failure
|  |  | UI | Machine Error | Call service to schedule a repair. | UI POST Bluetooth failure | If the UI POST reports bluetooth failure. | 267 | ALARM_ID_UI_POST_FAILURE_BLUETOOTH | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST Bluetooth failure
|  |  | UI | Machine Error | Call service to schedule a repair. | UI POST Ethernet failure | If the UI POST reports ethernet failure. | 268 | ALARM_ID_UI_POST_FAILURE_ETHERNET | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST Ethernet failure
|  |  | UI | Machine Error | Call service to schedule a repair. | UI POST Sound failure | If the UI POST reports sound failure. | 269 | ALARM_ID_UI_POST_FAILURE_SOUND | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST Sound failure
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD POST Safety Shutdown failure | If the 24 volts power source is greater than 5.0 volts during safety shutdown self test state. If the 24 volts power source is less than 22.6 volts during safety shutdown self test recover state. | 270 | ALARM_ID_HD_SAFETY_SHUTDOWN_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD POST Safety Shutdown failure
|  |  | DG | Self Test Error | Call service to schedule a repair. | DG POST Safety Shutdown failure | If the 24 volts power source is greater than 5.0 volts during safety shutdown self test state. If the 24 volts power source is less than 22.6 volts during safety shutdown self test recover state. | 271 | ALARM_ID_DG_SAFETY_SHUTDOWN_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG POST Safety Shutdown failure
|  |  | HD | Fans Error | Fan failure.\n- Check that the fans inlets are not blocked or obstructed. | HD Fan RPM out of range | If the fan RPM is out of Min/Max range. | 272 | ALARM_ID_HD_FAN_RPM_OUT_OF_RANGE | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | HD Fan RPM out of range
|  |  | HD | Blood Pump Error | Call service to schedule a repair. | HD measured blood pump flow rate is out of range | If the blood pump flow rate is > 1320.0 mL/min or < 1320.0 mL/min. | 273 | ALARM_ID_HD_BLOOD_FLOW_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD measured blood pump flow rate is out of range
|  |  | HD | Dialysate Pump Error | Call service to schedule a repair. | HD measured dialysate inlet pump flow rate is out of range | If the DPi pump flow rate is > 1320.0 mL/min or < 1320.0 mL/min. | 274 | ALARM_ID_HD_DIAL_IN_FLOW_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD measured dialysate inlet pump flow rate is out of range
|  |  | HD | Arterial Pressure Error | Call service to schedule a repair. | HD arterial pressure sensor is reading out of range | If the arterial pressure > 2000.0 mmHg or < -1500.0 mmHg. | 275 | ALARM_ID_HD_ARTERIAL_PRESSURE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD arterial pressure sensor is reading out of range
|  |  | HD | Venous Pressure Error | Call service to schedule a repair. | HD venous pressure sensor is reading out of range | If the venous pressure > 2000.0 mmHg or < -1500.0 mmHg. | 276 | ALARM_ID_HD_VENOUS_PRESSURE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD venous pressure sensor is reading out of range
|  |  | HD | Occlusion Sensor Error | Call service to schedule a repair. | HD BP occlusion sensor is reading out of range | If the blood pump occlusion count is < 2000. | 277 | ALARM_ID_HD_BP_OCCLUSION_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD BP occlusion sensor is reading out of range
|  |  | HD | Dialysate Production Error | Ultrafiltration error. | HD active reservoir weight out of range. | If the active reservoir weight is > 1950.0 g. | 278 | ALARM_ID_HD_ACTIVE_RESERVOIR_WEIGHT_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD active reservoir weight out of range.
|  |  | DG | Dialysate Drain Pump Error | Call service to schedule a repair. | DG dialysate drain time out | If the DG dialysate drain exceeded 120 seconds. | 279 | ALARM_ID_DG_DIALYSATE_DRAIN_TIME_OUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG dialysate drain time out
|  |  | HD | Drain Pump Error | Call service to schedule a repair. | DG drain pump invalid direction | If the drain pump 's direction is not 1 when draining. | 280 | ALARM_ID_DG_DRAIN_PUMP_DIRECTION_INVALID | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG drain pump direction invalid
|  |  | DG | Acid Concentrate Volume Low | Acid concentrate low.\n- Replace acid concentrate bottle. | DG Acid concentration bottle low volume alarm | If the acid concentration bottle volume > 3240 (3340 - 100). | 281 | ALARM_ID_DG_ACID_BOTTLE_LOW_VOLUME | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Acid concentration bottle low volume alarm
|  |  | DG | Bicarbonate Concentrate Volume Low | Bicarbonate concentrate low.\n- Replace bicarbonate concentrate bottle. | DG Bicarbonate concentration bottle low volume alarm | If the bicarbonate concentration bottle volume > 3680 (3780 - 100). | 282 | ALARM_ID_DG_BICARB_BOTTLE_LOW_VOLUME | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Bicarbonate concentration bottle low volume alarm
|  |  | DG | Load Cells Error | Call service to schedule a repair. | DG load cells weight out of range | If the weight of the DG load cells > 4500.0 g or < 0.0 g. | 283 | ALARM_ID_DG_LOAD_CELL_WEIGHT_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG load cells weight out of range
|  |  | DG | Load Cells Error | Call service to schedule a repair. | DG load cells primary/back up drift out of range | If redundant load cells deviate by more than 60g after tare for 5 seconds. | 284 | ALARM_ID_DG_LOAD_CELL_PRIMARY_BACKUP_DRIFT_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG load cells primary/back up drift out of range
|  |  | HD | Dialysate Recirculation Error | Dialysate recirculation error. | HD treatment recirculate timeout warning | If the maximum time in recirculate sub-mode has been exceeded. | 285 | ALARM_ID_HD_TREATMENT_RECIRC_TIMEOUT_WARNING | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | HD treatment recirculate timeout warning
|  |  | HD | Rinseback Complete | Rinseback complete. | HD treatment rinseback complete timeout warning | If the rise back time counter checks have failed when full blood volume has been delivered back to the patient. | 286 | ALARM_ID_HD_TREATMENT_RINSEBACK_TIMEOUT_WARNING | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | HD treatment rinseback complete timeout warning
|  |  | HD | Machine Error | Call service to schedule a repair. | HD processor clock speed checks against FPGA clock failure | If during initialization mode, the difference between the processor clock speed and the FPGA clock speed > 1. | 287 | ALARM_ID_HD_FPGA_CLOCK_SPEED_CHECK_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD processor clock speed checks against FPGA clock failure
|  |  | HD | Load Cells Error | Call service to schedule a repair. | HD load cell inconsistency | If redundant load cells deviate from start to end of a reservoir use during treatment by more than 10g. | 288 | ALARM_ID_HD_LOAD_CELL_PRIMARY_BACKUP_DRIFT_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD load cells primary/back up drift out of range
|  |  | DG | Cleaning Mode Error | Check dialysate/concentrate caps. | DG cap in wrong state | If the dialysate and concentrate caps are in the right position in the cleaning modes. | 289 | ALARM_ID_DG_DIALYSATE_OR_CONC_CAP_NOT_IN_PROPER_POSITION | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG dialysate or concentrate caps not closed
|  |  | HD | FPGA Error | Call service to schedule a repair. | HD venous sensor FPGA fault | If the FPGA does not report fresh data or indicates an error for a certain period of time. | 290 | ALARM_ID_HD_VENOUS_SENSOR_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD venous pressure sensor FPGA fault
|  |  | HD | Cleaning Modes Start Alarm | No Instructions | DG flush started | If DG flush successfully started from the UI. | 291 | ALARM_ID_HD_DISINFECT_FLUSH | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD processor is in Disinfect Flush mode
|  |  | HD | Cleaning Modes Start Alarm | No Instructions | DG heat disinfect started | If DG heat disinfect successfully started from the UI. | 292 | ALARM_ID_HD_DISINFECT_HEAT | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD processor is in Disinfect Heat mode
|  |  | HD | Cleaning Modes Start Alarm | No Instructions | DG chemical disinfect started | If DG chemical disinfect successfully started from the UI. | 293 | ALARM_ID_HD_DISINFECT_CHEM | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD processor is in Disinfect Chemical mode
|  |  | HD | Cleaning Modes Start Alarm | No Instructions | DG chemical disinfect instruction 1 | If DG chemical disinfect instruction has been called. | 294 | ALARM_ID_HD_DISINFECT_CHEM_INSTRUCTION1 | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD processor Disinfect Chemical mode instructions to user
|  |  | UI | Machine Error | Call service to schedule a repair. | UI cloud sync POST failure | If the UI POST reports cloud sync failure. | 295 | ALARM_ID_UI_POST_FAILURE_CLOUDSYNC | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST CloudSync failure
|  |  | DG | FPGA Error | Call service to schedule a repair. | DG TDi sensor FPGA fault | If the FPGA does not report fresh data or indicates an error for a certain period of time. | 296 | ALARM_ID_DG_TDI_SENSORS_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG TDi sensors FPGA fault
|  |  | DG | FPGA Error | Call service to schedule a repair. | DG TRo sensor FPGA fault | If the FPGA does not report fresh data or indicates an error for a certain period of time. | 297 | ALARM_ID_DG_TRO_SENSORS_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG TRo sensors FPGA fault
|  |  | DG | FPGA Error | Call service to schedule a repair. | DG barometric sensor FPGA fault | If the FPGA does not report fresh data or indicates an error for a certain period of time. | 298 | ALARM_ID_DG_BARO_SENSOR_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG baro sensor FPGA fault
|  |  | DG | Self Test Error | Call service to schedule a repair. | DG invalid serial number | If the DG serial number is empty and does not have any characters. | 299 | ALARM_ID_DG_INVALID_SERIAL_NUMBER | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG invalid serial number
|  |  | HD | Self Test Error | Call service to schedule a repair. | HD invalid serial number | If the HD serial number is empty and does not have any characters. | 300 | ALARM_ID_HD_INVALID_SERIAL_NUMBER | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD invalid serial number
|  |  | DG | Heaters Error | Call service to schedule a repair. | DG primary heater power line voltage out of range | If the DG primary heater power line voltage is off from 24V by 10%. | 301 | ALARM_ID_DG_PRIMARY_HEATER_POWER_VOLTAGE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG primary heater power voltage out of range
|  |  | DG | Processor RAM error | Call service to schedule a repair. | DG_CPU RAM error | If the RAM status indicates any of the following: Double bit error, Address decode failure, Address decode logic element failure, Read address parity failure, Write address parity failure. | 302 | ALARM_ID_DG_CPU_RAM_ERROR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG processor RAM error
|  |  | DG | FPGA Error | Call service to schedule a repair. | DG drain pump direction FPGA fault | If the FPGA indicates an error for a certain period of time. | 303 | ALARM_ID_DG_DRAIN_PUMP_DIRECTION_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG drain pump direction FPGA fault
|  |  | HD | Self-Test Error | Call service to schedule a repair. | HD invalid usage reocord | If the CRC of the HD usage record does not match the calculated CRC. | 304 | ALARM_ID_HD_INVALID_USAGE_RECORD_CRC | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD invalid usage record
|  |  | HD | Processor RAM error | Call service to schedule a repair. | HD_CPU RAM error | If the RAM status indicates any of the following: Double bit error, Address decode failure, Address decode logic element failure, Read address parity failure, Write address parity failure. | 305 | ALARM_ID_HD_CPU_RAM_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD processor RAM error
|  |  | HD | Air Trap Error | Air Trap Error.\n- Check air trap fill level. | HD air trap overfill fault | If the air trap sensor reports upper level while HD is in treatment mode for a certain period of time. | 306 | ALARM_ID_HD_AIR_TRAP_OVERFILL | ALARM_PRIORITY_HIGH | 3 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD air trap over-fill alarm
|  |  | HD | HD AC power loss | HD power has been lost. Check all electrical connections.\n- The display and CPU will run off of the battery until power returns. | HD AC power lost alarm | The HD Software shall trigger an alarm if an AC power loss is detected while the HD system is in pre-treatment, treatment, or post-treatment mode. | 307 | ALARM_ID_HD_AC_POWER_LOST_IN_TREATMENT | ALARM_PRIORITY_HIGH | 18 | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | HD AC power lost second alarm
|  |  | DG | Conductivity Sensors Error | Call service to schedule a repair. | DG CPi conductivity sensor transmits invalid characters | If the CPi conductivity sensor transmits invalid characters for a certain period of time. | 308 | ALARM_ID_DG_CPI_COND_SENSOR_INVALID_CHAR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG CPi conductivity sensor invalid character received
|  |  | DG | Conductivity Sensors Error | Call service to schedule a repair. | DG CPo conductivity sensor transmits invalid characters | If the CPo conductivity sensor transmits invalid characters for a certain period of time. | 309 | ALARM_ID_DG_CPO_COND_SENSOR_INVALID_CHAR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG CPo conductivity sensor invalid character received
|  |  | DG | Conductivity Sensors Error | Call service to schedule a repair. | DG CD1 conductivity sensor transmits invalid characters | If the CD1 conductivity sensor transmits invalid characters for a certain period of time. | 310 | ALARM_ID_DG_CD1_COND_SENSOR_INVALID_CHAR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG CD1 conductivity sensor invalid character received
|  |  | DG | Conductivity Sensors Error | Call service to schedule a repair. | DG CD2 conductivity sensor transmits invalid characters | If the CD2 conductivity sensor transmits invalid characters for a certain period of time. | 311 | ALARM_ID_DG_CD2_COND_SENSOR_INVALID_CHAR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG CD2 conductivity sensor invalid character received
|  |  | HD | Dialysate Flowrate Error | Call service to schedule a repair. | The dialysate inlet flow is too low. | If the dialysate inlet flow is too low. | 312 | ALARM_ID_HD_DIAL_IN_FLOW_CHECK_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD Dialysate In flow rate to too low.
|  |  | DG | Cleaning Mode Inlet Water Temperature Too Low | Inlet water temperature too low.\n- Check the incoming water for issues. | DG cleaning mode inlet water temperature too low | If inlet water temperature is < 24 C. | 313 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_TEMP_TOO_LOW | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG cleaning mode inlet water temperature too low.
|  |  | DG | Cleaning Mode Inlet Water Pressure Too High | Inlet water pressure too high.\n- Check the incoming water for issues. | DG cleaning mode inlet water pressure too high | If inlet water temperature is > 80 psig. | 314 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_PRESSURE_TOO_HIGH | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG cleaning mode inlet water pressure too high.
|  |  | DG | Cleaning Mode Inlet Water Pressure Too Low | Inlet water pressure too low.\n- Check the incoming water for issues. | DG cleaning mode inlet water pressure too low | If inlet water pressure is < 14 psig. | 315 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_PRESSURE_TOO_LOW | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG cleaning mode inlet water pressure too low.
|  |  | DG | Acid Concentrate Pump Park Fault | Call service to schedule a repair. | Acid concentrate pump did not park successfully | If the park took more than 1 second. | 316 | ALARM_ID_DG_ACID_CONCENTRATE_PUMP_PARK_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG acid concentrate pump park fault.
|  |  | DG | Bicarb Concentrate Pump Park Fault | Call service to schedule a repair. | Bicarb concentrate pump did not park successfully | If the park took more than 1 second. | 317 | ALARM_ID_DG_BICARB_CONCENTRATE_PUMP_PARK_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG bicarb concentrate pump part fault.
|  |  | HD | Air Pump Operation Timeout | Call service to schedule a repair. | HD Air Pump operation timeout | If air pump has been turned on for greater than 10s. | 318 | ALARM_ID_HD_AIR_PUMP_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD Air Pump timeout alarm
|  |  | DG | Chemical Disinfect Flush | Collect Flush Sample. | Collect sample to check the chemical disinfectant acid | Chemical disinfect flush sample alarm is prompted. | 319 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_FLUSH_SAMPLE | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG chemical disinfect flush flush sample.
|  |  | DG | Chemical Disinfect Flush | Call service to schedule a repair. | The user did not proceed with collecting chemical disinfect flush sample for more than 30 minutes | Chemical disinfect flush sample was not collected form more than 30 minutes. | 320 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_SAMPLE_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG chemical disinfect flush sample timeout.

## Property Definitions
Property Name | Definition | Column | Description
AlarmID | Unique alarm identification number | PRS ID | PRS requirement ID associated with the alarm.
Alarm Name | Name of the alarm | HD SRS ID | HD SRS requirement ID (if applicable) associated with the alarm.
Priority | Priority of the alarm (High, Medium, Low) | DG SRS ID | DG SRS requirement ID (if applicable) associated with the alarm.
Rank | Within the same priority, higher rank alarms have higher priority | Device | Sub-system that detects the alarm.
Fault | The alarm activation causes an HD fault | Description | Detailed description of the alarm.
DGFault | The alarm activation causes an DG fault | Displayed Title | Alarm title displayed prominantly at top of alarm dialog when the alarm is top priority.
Stop | The alarm activation causes a controlled stop | Displayed Instructions | Instructions displayed in alarm dialog when the alarm is top priority.
Clear Immediate | Alarm condition is cleared immediately after alarm activation | Trigger Conditions | Condition(s) required to trigger the alarm.
Non Recoverable | Alarm cannot be cleared (unrecoverable) | AlarmID | Unique alarm ID #.
No Resume | Alarm activation prevents treatment resume | Alarm Name | Enumeration name (code reference) for the alarm.
No Rinseback | Alarm activation prevents rinseback | Priority | Priority of the alarm (high, medium or low per 60601-1-8.
No End Treatment | Alarm activation prevents ending treatment | Rank | Rank of the alarm (used to determine which alarm is top priority) (low rank=high priority).
No New Treatment | Alarm activation prevents new treatment | Fault | Does HD f/w consider the alarm to be a system fault.
No Dialysate Recirc | Alarm activation prevents dialysate re-circulation | DGFault | Does DG f/w consider the alarm to be a system fault (will not prevent HD from some activities).
Clear Top Only | User acknowledgment clears only displayed alarm | Stop | Does the HD f/w need to pause treatment (or perhaps other activities) when alarm is active.
Escalate In | Alarm will escalate if not cleared within this time (zero indicates no escalation) | Clear Immediate | Should the f/w immediately clear the alarm condition after trigger (not clear the alarm itself).
Escalate To | ID of the alarm that this alarm will be escalated to | Non Recoverable | Is the alarm recoverable (can user clear the alarm).
|  | No Resume | Should treatment resume be prevented when the alarm is active.
|  | No Rinseback | Should rinseback be prevented when the alarm is active.
|  | No End Treatment | Should ending treatment be prevented when the alarm is active.
|  | No Blood Recirc | Should blood recirculation be prevented when the alarm is active (treatment is paused). Dialyzer is bypassed regardless.
|  | No Dialysate Recirc | Should dialysate recirculation be prevented when the alarm is active (treatment is paused). Dialyzer is bypassed regardless.
|  | Clear Top Only | If user acknowledges alarm, should all active alarms be cleared or just the alarm top (displayed) alarm.
|  | TreatmentLog | Should alarm be included in treatment log.
|  | Escalate In | Duration (in seconds) until alarm escalates to a higher priority alarm (0 indicates no escalation).
|  | Escalate To | Higher priority alarm to escalate to when escalation period expires (None indicates no escalation).
|  | Doxygen Title | Used for design documentation purposes.