## Alarms
Action | PRS ID2 | SA | HD SRS ID | DG SRS ID | UI SRS ID | Device | Displayed Title | Char count | Displayed Instructions | Char Count2 | Description | Trigger Conditions | AlarmID | Alarm Name | Alarm ID Check | Priority | Rank | Fault | DGFault | Stop | Clear Immediate | Non Recoverable | No Resume | No Rinseback | No End Treatment | Block Until SB Rinseback | Block Until SB End Tx | No Blood Recirc | No Dialysate Recirc | Clear Top Only | TreatmentLog | Doxygen Title
N/A | N/A |  | N/A | N/A |  | N/A | Indication for when no alarms is a possible situation |  | N/A |  | Indication for when no alarms is a possible situation | N/A | 0 | ALARM_ID_NO_ALARM |  | ALARM_PRIORITY_NONE | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | Indication for when no alarms is a possible situation | N/A | `
x | PRS 398 | SA 213 | SRSHD 652 |  |  | Stop or On button pressed during startup\n-Release button and power cycle\n-Call service to schedule a repair if problem persists | Stuck button POST failure. Stop or Off button detected to be pressed for at least 1 second during test shortly after power up |  | Stop or Off button are pressed for more than 1 sec during POST. |  | Stuck button POST failure. Stop or Off button detected to be pressed for at least 1 second during test shortly after power up | Stop or Off button are pressed for more than 1 sec during POST. | 1 | ALARM_ID_HD_STUCK_BUTTON_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD stuck button POST failure. Stop or Off button detected to be pressed for at least 1 second during test shortly after power up | HD | HD POST: Stuck Button
x | PRS 398 | SA 207 | SRSHD 995 SRSHD 1357 |  |  | Call service to schedule a repair. | HD FPGA POST failure |  | Minimum HD FPGA major revision is not compatible with this version of HD firmware or the FPGA Id is not correct. |  | HD FPGA POST failure | Minimum HD FPGA major revision is not compatible with this version of HD firmware or the FPGA Id is not correct. | 2 | ALARM_ID_HD_FPGA_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD FPGA POST failure | HD | HD POST: FPGA Compatibility
x | PRS 863 | SA 287 |  | SRSDG 602 SRSDG 814 SRSDG 1061 |  | Call service to schedule a repair. | DG FPGA POST failure |  | Minimum DG FPGA major and minor revision are not equal to the FPGA major and minor revision in FPGA header |  | DG FPGA POST failure | Minimum DG FPGA major and minor revision are not equal to the FPGA major and minor revision in FPGA header | 3 | ALARM_ID_DG_FPGA_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG FPGA POST failure | DG | DG POST: FPGA Compatibility
Sean - check if DG and HD POST test is the same in code. Consider updating trigger conditions | PRS 398 | SA 208 SA 209 SA 210 SA 211 | SRSHD 650 |  |  | Call service to schedule a repair. | HD Watchdog POST failure |  | If the following are not observed after 500ms of not petting the WD: 1. WD exp pin=low 2. 24V <= 5V 3. audio current >= 200mA If the following are not observed after 500ms of resumed WD petting at proper timing: 1. WD exp pin=high 2. 24V >= 22.6V 3. audio current <= 10 mA |  | HD Watchdog POST failure | If the following are not observed after 500ms of not petting the WD: 1. WD exp pin=low 2. 24V <= 5V 3. audio current >= 200mA If the following are not observed after 500ms of resumed WD petting at proper timing: 1. WD exp pin=high 2. 24V >= 22.6V 3. audio current <= 10 mA | 4 | ALARM_ID_HD_WATCHDOG_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD Watchdog POST failure | HD | HD POST: WD Safestate
Sean - check if DG and HD POST test is the same in code. Consider updating trigger conditions | PRS 863 | SA 288 SA 290 SA 289 |  | SRSDG 53 |  | Call service to schedule a repair. | DG Watchdog POST failure |  | If the following are not observed after 500ms of not petting the WD: 1. WD exp pin=low 2. 24V <= 5V If the following are not observed after 750ms of resumed WD petting at proper timing: 1. WD exp pin=high 2. 24V >= 22.6V |  | DG Watchdog POST failure | If the following are not observed after 500ms of not petting the WD: 1. WD exp pin=low 2. 24V <= 5V If the following are not observed after 750ms of resumed WD petting at proper timing: 1. WD exp pin=high 2. 24V >= 22.6V | 5 | ALARM_ID_DG_WATCHDOG_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG Watchdog POST failure | DG | DG POST: WD Safestate
x | PRS 398 | SA 204 | SRSHD 1484 |  |  | Call service to schedule a repair. | UI communication POST failure. UI failed to communicate within a reasonable time after power up |  | If the UI failed to checkin via the CAN bus within 40 seconds |  | UI communication POST failure. UI failed to communicate within a reasonable time after power up | If the UI failed to checkin via the CAN bus within 40 seconds | 6 | ALARM_ID_HD_UI_COMM_POST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI communication POST failure. UI failed to communicate within a reasonable time after power up | HD | HD POST: Proc-UI Comm
x | PRS 872 | SA 457 | SRSHD 1595 |  |  | Call service to schedule a repair. | Real-time clock configuration error |  | If the RTC is not functional as indicated by the bits in the first 3 control registers. |  | Real-time clock configuration error | If the RTC is not functional as indicated by the bits in the first 3 control registers. | 7 | ALARM_ID_HD_RTC_BATTERY_LOW |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD RTC battery low | HD | HD Alarm (Low): RTC Bat Low
x | PRS 398 | SA 471 | SRSHD 632 |  |  | Call service to schedule a repair. | HD accelerometer failed POST |  | If the calculated vector length of the accelerometer x, y, z is not the nominal length. |  | HD accelerometer failed POST | If the calculated vector length of the accelerometer x, y, z is not the nominal length. | 8 | ALARM_ID_HD_ACCELEROMETER_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD accelerometer failed POST | HD | HD POST: Acc Vector
x | PRS 863 | SA 472 |  | SRSDG 323 |  | Call service to schedule a repair. | DG accelerometer failed POST |  | If the calculated vector length of the accelerometer x, y, z is not the nominal length. |  | DG accelerometer failed POST | If the calculated vector length of the accelerometer x, y, z is not the nominal length. | 9 | ALARM_ID_DG_ACCELEROMETER_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG accelerometer failed POST | DG | DG POST: Acc Vector
x | PRS 398 | SA 480 | SRSHD 1478 |  |  | Call service to schedule a repair. | HD blood leak sensor embedded mode communication failed |  | If the blood leak sensor embedded mode is unresponsive. |  | HD blood leak sensor embedded mode communication failed | If the blood leak sensor embedded mode is unresponsive. | 10 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_EMBEDDED_MODE_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD blood leak sensor setting embedded mode failure. | HD | HD POST: BLD Config Emb Mode
Dialysate Temp | [5 empty] | Dialysate temperature is low\n- Treatment is paused and will be ready to resume when \ndialysate is within target range. | Dialysate temperature is below the target temperature |  | If the dialysate temperature is below the target temperature for a certain period of time. |  | Dialysate temperature is below the target temperature | If the dialysate temperature is below the target temperature for a certain period of time. | 11 | ALARM_ID_HD_DIALYSATE_TEMP_BELOW_TARGET_TEMP |  | ALARM_PRIORITY_MEDIUM | 402 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD dialysate temperature below target temperature or below safety temp | HD | Dial. Gen: Dialysate Temp Low
x | PRS 398 | SA 195 | SRSHD 658 |  |  | Call service to schedule a repair. | HD firmware image integrity POST test failed |  | If the calculated CRC of the HD image does not match the expected CRC. |  | HD firmware image integrity POST test failed | If the calculated CRC of the HD image does not match the expected CRC. | 12 | ALARM_ID_HD_INTEGRITY_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD firmware image integrity POST test failed | HD | HD POST: HD Image
x | PRS 863 | SA 282 |  | SRSDG 813 |  | Call service to schedule a repair. | DG firmware image integrity POST test failed |  | If the calculated CRC of the DG firmware image does not matched the expected CRC |  | DG firmware image integrity POST test failed | If the calculated CRC of the DG firmware image does not matched the expected CRC | 13 | ALARM_ID_DG_INTEGRITY_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG firmware image integrity POST test failed | DG | DG POST: DG Image
x | PRS 863 | SA 284 |  | SRSDG 957 |  | Call service to schedule a repair. | DG non-volatile usage information group invalid record CRC |  | If the DG usage information record's CRC does not match the calculated CRC in POST. |  | DG non-volatile usage information group invalid record CRC | If the DG usage information record's CRC does not match the calculated CRC in POST. | 14 | ALARM_ID_DG_INVALID_USAGE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid usage record CRC | DG | DG POST: Usage Rec NVRAM CRC
Dara - remove Alarm ID 15 (Duplicate with Alarm ID 304 - SRS 1587) | [5 empty] | Dialysate temperature is high\n-Treatment is paused and will be ready to resume when \ndialysate is within target range. | Dialysate temperature is out of high safety range |  | If the dialysate temperature is > 46 C for a certain period of time. |  | Dialysate temperature is out of high safety range | If the dialysate temperature is > 46 C for a certain period of time. | 15 | ALARM_ID_HD_DIALYSATE_TEMP_ABOVE_SAFETY_TEMP |  | ALARM_PRIORITY_MEDIUM | 401 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD dialysate temperature above high safety. | HD | Dial. Gen: Dialysate Temp High
x | PRS 398 | SA 212 | SRSHD 656 |  |  | Call service to schedule a repair. | HD alarm audio failed POST |  | If alarm audio current is NOT in expected range for than 1 second while alarm tone is being output |  | HD alarm audio failed POST | If alarm audio current is NOT in expected range for than 1 second while alarm tone is being output | 16 | ALARM_ID_HD_ALARM_AUDIO_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD alarm audio failed POST | HD | HD POST: Primary Alarm Current
Sean to review and update |  |  | SRSHD 1677 |  |  | Call service to schedule a repair. | HD UI POST failed |  | If UI reports POST failed but did not trigger fault OR UI has not sent POST results within 2 seconds. |  | HD UI POST failed | If UI reports POST failed but did not trigger fault OR UI has not sent POST results within 2 seconds. | 17 | ALARM_ID_HD_UI_POST_TIMEOUT |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI POST failed | HD | HD POST: UI POST Not Complete
[6 empty] | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | HD to DG comm error. |  | HD message to DG that requires acknowledge was not acknowledged within a certain time. |  | HD to DG comm error. | HD message to DG that requires acknowledge was not acknowledged within a certain time. | 18 | ALARM_ID_HD_CAN_MESSAGE_NOT_ACKED_BY_DG |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD didn't get ACK on message to DG that required acknowledgment | HD | Loss of DG Comm (DG ACK)
Dialysate Temp |  |  | SRSHD 901 |  |  | Dialysate temperature is high\n- Treatment is paused and will be ready to resume when \ndialysate is within target range. | Dialysate temperature is above the target temperature |  | If the dialysate temperature is above the target temperature for a certain period of time. |  | Dialysate temperature is above the target temperature | If the dialysate temperature is above the target temperature for a certain period of time. | 19 | ALARM_ID_HD_DIALYSATE_TEMP_ABOVE_TARGET_TEMP |  | ALARM_PRIORITY_MEDIUM | 401 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD dialysate temperature above target temperature. | HD | Dial. Gen: Dialysate Temp High
x | PRS 863 | SA 369 |  | SRSDG 16 |  | Call service to schedule a repair. | DG conductivity sensors invalid calibration record |  | If the calibration data that was received from NVDataMgmt is NOT legitimate |  | DG conductivity sensors invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 20 | ALARM_ID_DG_COND_SENSORS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG conductivity sensors invalid calibration record | DG | DG POST: Conductivity Cal
x | PRS 863 | SA 369 |  | SRSDG 936 |  | Call service to schedule a repair. | DG drain line volume invalid calibration record |  | If the calibration data that was received from NVDataMgmt is NOT legitimate |  | DG drain line volume invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 21 | ALARM_ID_DG_DRAIN_LINE_VOLUME_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain line volume invalid calibration record | DG | DG POST: Drain Line Vol Cal
x | PRS 863 | SA 369 |  | SRSDG 935 |  | Call service to schedule a repair. | DG reservoirs invalid calibration record |  | If the calibration data that was received from NVDataMgmt is NOT legitimate |  | DG reservoirs invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 22 | ALARM_ID_DG_RESERVOIRS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG reservoirs invalid calibration record | DG | DG POST: Reservoir Vol Cal
x | PRS 863 | SA 369 |  | SRSDG 934 |  | Call service to schedule a repair. | DG acid concentrate calibration record |  | If the non-volatile data is valid and if not raise the alarm |  | DG acid concentrate calibration record | If the non-volatile data is valid and if not raise the alarm | 23 | ALARM_ID_DG_ACID_CONCENTRATE_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG acid concentrate calibration record | DG | DG POST: Acid Conc Cal
x | PRS 863 | SA 369 |  | SRSDG 933 |  | Call service to schedule a repair. | DG bicarb concentrate calibration record |  | If the non-volatile data is valid and if not raise the alarm |  | DG bicarb concentrate calibration record | If the non-volatile data is valid and if not raise the alarm | 24 | ALARM_ID_DG_BICARB_CONCENTRATE_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG bicarb concentrate calibration record | DG | DG POST: Bicarb Conc Cal
x | PRS 863 | SA 369 |  | SRSDG 817 |  | Call service to schedule a repair. | DG accelerometer invalid calibration record |  | If the calibration data that was received from NVDataMgmt is NOT legitimate |  | DG accelerometer invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 25 | ALARM_ID_DG_ACCELEROMETERS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG accelerometer invalid calibration record | DG | DG POST: Acc Cal
x | PRS 398 | SA 368 | SRSHD 1576 |  |  | Call service to schedule a repair. | HD accelerometer invalid calibration record |  | If the accel calibration data that was received from NVDataMgmt is NOT legitimate. |  | HD accelerometer invalid calibration record | If the accel calibration data that was received from NVDataMgmt is NOT legitimate. | 26 | ALARM_ID_HD_ACCELEROMETERS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD accelerometer invalid calibration record | HD | HD POST: Acc Cal
Move to wet self-test - procedural |  |  | SRSHD 1428 |  |  | Self-test detected has a problem with the blood leak sensor\n- Check cartridge installation on to machine\n- If problem persists, terminate pre-treatment and restart \nwith a new cartridge. | HD BLD sensor failed zero sequence |  | BLD fails to complete zero sequence. |  | HD BLD sensor failed zero sequence | BLD fails to complete zero sequence. | 27 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_ZERO_SEQUENCE_FAILED |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD blood leak sensor zero and self test sequence failed | HD | Process: Pre-Tx Wet Self-Test
x | PRS 863 | SA 319 SA 320 |  | SRSDG 864 |  | Call service to schedule a repair. | DG two wire sensors FPGA fault |  | If the FPGA does not report fresh data for a certain period of time |  | DG two wire sensors FPGA fault | If the FPGA does not report fresh data for a certain period of time | 28 | ALARM_ID_DG_TWO_WIRE_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG two wire sensors FPGA fault | DG | DG Fault: Conductivity RTD Comm
x | PRS 398 | SA 368 | SRSHD 1482 |  |  | Call service to schedule a repair. | HD heparin force sensor invalid calibration record |  | If the heparin force sensor calibration data that was received from NVDataMgmt is NOT legitimate. |  | HD heparin force sensor invalid calibration record | If the heparin force sensor calibration data that was received from NVDataMgmt is NOT legitimate. | 29 | ALARM_ID_HD_HEPARIN_FORCE_SENSOR_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD heparin force sensor invalid calibration record | HD | HD POST: SP DAC Cal
x | PRS 494 | SA 416 | SRSHD 1704 |  |  | Call service to schedule a repair. | HD Software fault. Software found itself in an unexpected state |  | If the specific software faults listed in alarm management SW faults has been detected |  | HD Software fault. Software found itself in an unexpected state | If the specific software faults listed in alarm management SW faults has been detected | 30 | ALARM_ID_HD_SOFTWARE_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD Software fault. Software found itself in an unexpected state | HD | HD Fault: SW Invalid State
x | PRS 494 | SA 241 SA 242 | SRSHD 515 |  |  | Call service to schedule a repair. | Blood pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running |  | If the measured MC current is not greater than BP_MAX_CURR_WHEN_STOPPED_MA or is not greater than BP_MAX_CURR_WHEN_RUNNING_MA. |  | Blood pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | If the measured MC current is not greater than BP_MAX_CURR_WHEN_STOPPED_MA or is not greater than BP_MAX_CURR_WHEN_RUNNING_MA. | 31 | ALARM_ID_HD_BLOOD_PUMP_MC_CURRENT_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | HD | HD Fault: BP Current
x | PRS 494 | SA 239 | SRSHD 1087 |  |  | Call service to schedule a repair. | Blood pump failed motor off check. Measured speed while commanded off |  | If the measured motor speed is greater than BP_MAX_MOTOR_SPEED_WHILE_OFF_RPM |  | Blood pump failed motor off check. Measured speed while commanded off | If the measured motor speed is greater than BP_MAX_MOTOR_SPEED_WHILE_OFF_RPM | 32 | ALARM_ID_HD_BLOOD_PUMP_OFF_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed motor off check. Measured speed while commanded off | HD | HD Fault: BP Speed
x | PRS 494 | SA 240 SA 244 | SRSHD 1079 |  |  | Call service to schedule a repair. | Blood pump failed motor direction check. Measured vs. commanded |  | If the set direction versus hall sensors direction versus sign of motor speed DOES NOT agree. |  | Blood pump failed motor direction check. Measured vs. commanded | If the set direction versus hall sensors direction versus sign of motor speed DOES NOT agree. | 33 | ALARM_ID_HD_BLOOD_PUMP_MC_DIRECTION_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed motor direction check. Measured vs. commanded | HD | HD Fault: BP Direction
x | PRS 494 | SA 246 | SRSHD 1081 |  |  | Blood pump rotor failure.\n- Check the pump rotor for secure connection. | Blood pump failed rotor speed check. Mismatch with rotor and motor speeds |  | If the delta of the measured motor speed and the commanded motor speed is not within tolerance. |  | Blood pump failed rotor speed check. Mismatch with rotor and motor speeds | If the delta of the measured motor speed and the commanded motor speed is not within tolerance. | 34 | ALARM_ID_HD_BLOOD_PUMP_ROTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed rotor speed check. Mismatch with rotor and motor speeds | HD | HD Fault: BP Rotor
x | PRS 494 | SA 250 SA 251 | SRSHD 542 |  |  | Call service to schedule a repair. | Dialysis inlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running |  | If the Motor Controller measured DPi pump current > 150.0 mA for more than 250.0 msec. |  | Dialysis inlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | If the Motor Controller measured DPi pump current > 150.0 mA for more than 250.0 msec. | 35 | ALARM_ID_HD_DIAL_IN_PUMP_MC_CURRENT_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD dialysis inlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | HD | HD Fault: DPi Current
x | PRS 494 | SA 248 | SRSHD 1088 |  |  | Call service to schedule a repair. | Dialysis inlet pump failed motor off check. Measured speed while commanded off |  | If during command off state, the DPi pump motor speed is measured > 100.0 RMP. |  | Dialysis inlet pump failed motor off check. Measured speed while commanded off | If during command off state, the DPi pump motor speed is measured > 100.0 RMP. | 36 | ALARM_ID_HD_DIAL_IN_PUMP_OFF_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD dialysis inlet pump failed motor off check. Measured speed while commanded off | HD | HD Fault: DPi Speed
x | PRS 494 | SA 249 SA 253 | SRSHD 1050 |  |  | Call service to schedule a repair. | Dialysis inlet pump failed motor direction check. Measured vs. commanded |  | If the set direction versus hall sensors direction DO NOT agree for more than 250.0 msec. |  | Dialysis inlet pump failed motor direction check. Measured vs. commanded | If the set direction versus hall sensors direction DO NOT agree for more than 250.0 msec. | 37 | ALARM_ID_HD_DIAL_IN_PUMP_MC_DIRECTION_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD dialysis inlet pump failed motor direction check. Measured vs. commanded | HD | HD Fault: DPi Direction
x | PRS 494 | SA 255 | SRSHD 1082 |  |  | Dialysate in pump rotor failure.\n- Check the pump rotor for secure connection. | Dialysis inlet pump failed rotor speed check. Mismatch with rotor and motor speeds |  | If the measured motor speed versus the measured motor speed while controlling the target > 5.0 for more than 12 seconds. |  | Dialysis inlet pump failed rotor speed check. Mismatch with rotor and motor speeds | If the measured motor speed versus the measured motor speed while controlling the target > 5.0 for more than 12 seconds. | 38 | ALARM_ID_HD_DIAL_IN_PUMP_ROTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD dialysis inlet pump failed rotor speed check. Mismatch with rotor and motor speeds | HD | HD Fault: DPi Rotor
x | PRS 494 | SA 258 SA 259 | SRSHD 554 |  |  | Call service to schedule a repair. | Dialysis outlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running |  | If the DPo pump current > 150.0 mA while stopped or Dpo current > 2000.0 mA for more than 2 seconds while running. |  | Dialysis outlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | If the DPo pump current > 150.0 mA while stopped or Dpo current > 2000.0 mA for more than 2 seconds while running. | 39 | ALARM_ID_HD_DIAL_OUT_PUMP_MC_CURRENT_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD dialysis outlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | HD | HD Fault: DPo Current
x | PRS 494 | SA 256 | SRSHD 1089 |  |  | Call service to schedule a repair. | Dialysis outlet pump failed motor off check. Measured speed while commanded off |  | If the Dpo pump current > 100.0 RPM while commanded off for more than 5000 msec. |  | Dialysis outlet pump failed motor off check. Measured speed while commanded off | If the Dpo pump current > 100.0 RPM while commanded off for more than 5000 msec. | 40 | ALARM_ID_HD_DIAL_OUT_PUMP_OFF_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD dialysis outlet pump failed motor off check. Measured speed while commanded off | HD | HD Fault: DPo Speed
x | PRS 494 | SA 257 SA 261 | SRSHD 1051 |  |  | Call service to schedule a repair. | Dialysis outlet pump failed motor direction check. Measured vs. commanded |  | If the DPo pump set direction does not agree with hall sensors or set direction does not agree with the sign of the motor speed for more than 25 msec. |  | Dialysis outlet pump failed motor direction check. Measured vs. commanded | If the DPo pump set direction does not agree with hall sensors or set direction does not agree with the sign of the motor speed for more than 25 msec. | 41 | ALARM_ID_HD_DIAL_OUT_PUMP_MC_DIRECTION_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD dialysis outlet pump failed motor direction check. Measured vs. commanded | HD | HD Fault: DPo Direction
x | PRS 494 | SA 263 | SRSHD 1083 |  |  | Dialysate out pump rotor failure.\n- Check the pump rotor for secure connection. | Dialysis outlet pump failed rotor speed check. Mismatch with rotor and motor speeds |  | If DPo pump measured rotor speed does not agreed with the measured/commanded motor speed |  | Dialysis outlet pump failed rotor speed check. Mismatch with rotor and motor speeds | If DPo pump measured rotor speed does not agreed with the measured/commanded motor speed | 42 | ALARM_ID_HD_DIAL_OUT_PUMP_ROTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD dialysis outlet pump failed rotor speed check. Mismatch with rotor and motor speeds | HD | HD Fault: DPo Rotor
x | PRS 863 | SA 357 |  | SRSDG 954 |  | Call service to schedule a repair. | Main primary heater FPGA fault |  | If the FPGA does not report fresh data for a certain period of time |  | Main primary heater FPGA fault | If the FPGA does not report fresh data for a certain period of time | 43 | ALARM_ID_DG_MAIN_PRIMARY_HEATER_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG main primary heater FPGA fault | DG | DG Fault: Iso ADC Comm
x | PRS 494 | SA 226 | SRSHD 1001 |  |  | Call service to schedule a repair. | User interface communication timeout |  | If the sub-system communication timeout has occurred. UI is 5 seconds. DG is 2 seconds. |  | User interface communication timeout | If the sub-system communication timeout has occurred. UI is 5 seconds. DG is 2 seconds. | 44 | ALARM_ID_HD_UI_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI communication timeout | HD | HD Fault: UI Comm
x | PRS 494 | SA 226 SA 477 | SRSHD 967 |  |  | Call service to schedule a repair. | Too many bad CRCs detected on received system messages |  | If too many bad messages CRCs are received within a set period of time |  | Too many bad CRCs detected on received system messages | If too many bad messages CRCs are received within a set period of time | 45 | ALARM_ID_HD_COMM_TOO_MANY_BAD_CRCS |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD too many bad communications CRC | HD | HD Fault: CAN CRC Com
x | PRS 494 | SA 226 SA 477 | SRSHD 964 |  |  | Call service to schedule a repair. | System message that required acknowledgment was not acknowledged |  | If no retries left in the pending acknowlegement list. |  | System message that required acknowledgment was not acknowledged | If no retries left in the pending acknowlegement list. | 46 | ALARM_ID_HD_CAN_MESSAGE_NOT_ACKED_BY_UI |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD didn't get ACK on message to UI that required acknowledgment | HD | HD Fault: CAN ACK UI Com
x | PRS 494 | SA 476 | SRSHD 907 | [9 empty] | 47 | ALARM_ID_AVAILABLE_16 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | Available for use
x | PRS 494 | SA 452 | SRSHD 1029 |  |  | Call service to schedule a repair. | Ultrafiltration volume accuracy error during treatment |  | If the UF volume (refUFvolume - measureVolume) > 100.0 mL. |  | Ultrafiltration volume accuracy error during treatment | If the UF volume (refUFvolume - measureVolume) > 100.0 mL. | 48 | ALARM_ID_HD_UF_VOLUME_ACCURACY_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD ultrafiltration volume accuracy error during treatment | HD | HD Fault: UF Rate Tare Err
x | PRS 494 | SA 224 | SRSHD 1356 SRSHD 1490 |  |  | Call service to schedule a repair. | HD FPGA communication down for too long |  | If there are more than 3 comm errors in a 1 min period. |  | HD FPGA communication down for too long | If there are more than 3 comm errors in a 1 min period. | 49 | ALARM_ID_HD_FPGA_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD FPGA communication down for too long | HD | HD Fault: FPGA Com
x | PRS 863 | SA 481 |  | SRSDG 26 |  | Call service to schedule a repair. | DG FPGA not accepting commanded valve states |  | If the read back FPGA valve states does not match the commanded valve states for a certain amount of time. |  | DG FPGA not accepting commanded valve states | If the read back FPGA valve states does not match the commanded valve states for a certain amount of time. | 50 | ALARM_ID_DG_VALVE_CONTROL_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG FPGA not accepting commanded valve states | DG | DG Fault: Valve Control
x | PRS 494 | SA 239 SA 243 | SRSHD 1098 |  |  | Call service to schedule a repair. | Blood pump failed motor speed check. Measured vs. commanded |  | If measured motor speed versus commanded rotor speed exceeded limits. |  | Blood pump failed motor speed check. Measured vs. commanded | If measured motor speed versus commanded rotor speed exceeded limits. | 51 | ALARM_ID_HD_BLOOD_PUMP_MOTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed motor speed check. Measured vs. commanded | HD | HD Fault: BP Speed
x | PRS 494 | SA 252 | SRSHD 1099 |  |  | Call service to schedule a repair. | Dialysate inlet pump failed motor speed check. Measured vs. commanded |  | If measured motor speed versus commanded motor speed exceeded limits for more than 5 secs. |  | Dialysate inlet pump failed motor speed check. Measured vs. commanded | If measured motor speed versus commanded motor speed exceeded limits for more than 5 secs. | 52 | ALARM_ID_HD_DIAL_IN_PUMP_MOTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD dialysate inlet pump failed motor speed check. Measured vs. commanded | HD | HD Fault: DPi Speed
x | PRS 494 | SA 260 SA 256 | SRSHD 1198 SRSHD 1431 SRSHD 1670 |  |  | Call service to schedule a repair. | Dialysate outlet pump failed motor speed check. Measured vs. commanded |  | If measured motor speed versus commanded motor speed exceeded limits for more than 5 secs. |  | Dialysate outlet pump failed motor speed check. Measured vs. commanded | If measured motor speed versus commanded motor speed exceeded limits for more than 5 secs. | 53 | ALARM_ID_HD_DIAL_OUT_PUMP_MOTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD dialysate outlet pump failed motor speed check. Measured vs. commanded | HD | HD Fault: DPo Speed
x | PRS 494 | SA 501 | SRSHD 1705 |  |  | Call service to schedule a repair. | HD critical data integrity check failed |  | If the critical data failed range check. |  | HD critical data integrity check failed | If the critical data failed range check. | 54 | ALARM_ID_HD_CRITICAL_DATA_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD critical data integrity check failed | HD | HD Fault: Critical Data Err
| PRS 863 | SA 500 |  | SRSDG 921 |  | Call service to schedule a repair. | DG critical data integrity check failed |  | If the critical data failed range check. |  | DG critical data integrity check failed | If the critical data failed range check. | 55 | ALARM_ID_DG_CRITICAL_DATA_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG critical data integrity check failed | DG | DG Fault: Critical Data Err
x | PRS 494 | SA 232 | SRSHD 1526 |  |  | Call service to schedule a repair. | HD accelerometer error (no readings or FPGA reports error) |  | If no fresh sample of accelerometer data has been detected for more than 100 msec. |  | HD accelerometer error (no readings or FPGA reports error) | If no fresh sample of accelerometer data has been detected for more than 100 msec. | 56 | ALARM_ID_HD_ACCELEROMETER_FAILURE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD accelerometer error (no readings or FPGA reports error) | HD | HD Fault: Acc Com
x | PRS 863 | SA 346 |  | SRSDG 869 SRSDG 870 |  | Call service to schedule a repair. | DG accelerometer error (no readings or FPGA reports error) |  | If no fresh sample of accelerometer data has been detected for more than 100 msec. |  | DG accelerometer error (no readings or FPGA reports error) | If no fresh sample of accelerometer data has been detected for more than 100 msec. | 57 | ALARM_ID_DG_ACCELEROMETER_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG accelerometer error (no readings or FPGA reports error) | DG | DG Fault: Acc Com
x | PRS 494 | SA 384 SA 385 SA 386 SA 387 | SRSHD 1340 |  |  | Call service to schedule a repair. | HD valve homing failed |  | If the number of failed homing > = 3 counts. |  | HD valve homing failed | If the number of failed homing > = 3 counts. | 58 | ALARM_ID_HD_VALVE_HOMING_FAILED |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD valve homing failed | HD | HD Fault: Valve Home
x | PRS 494 | SA 264 SA 266 SA 268 SA 270 | SRSHD 1338 |  |  | Call service to schedule a repair. | HD valve transition time out |  | if the valve's transition time has timed out (3000msec) |  | HD valve transition time out | if the valve's transition time has timed out (3000msec) | 59 | ALARM_ID_HD_VALVE_TRANSITION_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD valve transition time out | HD | HD Fault: Valve Transition Err
x | PRS 494 | SA 502 SA 505 SA 503 SA 504 | SRSHD 1515 SRSHD 1516 SRSHD 1632 |  |  | Call service to schedule a repair. | HD valve not functional |  | If the valve bit masks are not as expected in different valve control mode. |  | HD valve not functional | If the valve bit masks are not as expected in different valve control mode. | 60 | ALARM_ID_HD_VALVE_NOT_FUNCTIONAL |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD valve not functional | HD | HD Fault: Valve Control
x | PRS 494 | SA 265 SA 267 SA 269 SA 271 | SRSHD 1339 |  |  | Call service to schedule a repair. | HD valve current out of range |  | If the valve current is over the threshold of 1.0 A for 1000 msec. |  | HD valve current out of range | If the valve current is over the threshold of 1.0 A for 1000 msec. | 61 | ALARM_ID_HD_VALVE_CURRENT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD valve current out of range | HD | HD Fault: Valve Current
x | PRS 494 | SA 264 SA 266 SA 268 SA 270 | SRSHD 1336 |  |  | Call service to schedule a repair. | HD valve position out of target |  | If the current position has deviated from the expected position for more than 1 second. |  | HD valve position out of target | If the current position has deviated from the expected position for more than 1 second. | 62 | ALARM_ID_HD_VALVE_POSITION_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD valve position out of target | HD | HD Fault: Valve Position
x | PRS 494 | SA 453 | SRSHD 1586 |  |  | Call service to schedule a repair. | FPGA reports error |  | If the FPGA error counter increments for a certain period of time. |  | FPGA reports error | If the FPGA error counter increments for a certain period of time. | 63 | ALARM_ID_HD_BLOOD_LEAK_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD blood leak FPGA fault | HD | HD Fault: BLD Com
x | PRS 863 | SA 350 |  | SRSDG 874 |  | Call service to schedule a repair. | Barometric pressure is out of range |  | If the barometric pressure is > 10.1 psi or < 15.4 psi for a certain period of time. |  | Barometric pressure is out of range | If the barometric pressure is > 10.1 psi or < 15.4 psi for a certain period of time. | 64 | ALARM_ID_DG_BARO_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG barometric pressure sensor out of range | DG | DG Fault: Baro Range
x | PRS 400 |  | SRSHD 1531 |  |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | HD requests DG command with invalid parameter fault |  | If the latest DG command response indicated invalid parameter |  | HD requests DG command with invalid parameter fault | If the latest DG command response indicated invalid parameter | 65 | ALARM_ID_HD_DG_COMMAND_INVALID_PARAMETER_FAULT |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD DG requests DG command with invalid parameter fault | HD | Loss of DG Comm (Reject)
x | PRS 398 | SA 480 | SRSHD 1427 SRSHD 1429 SRSHD 1631 |  |  | Call service to schedule a repair. | HD blood leak sensor set point set failure |  | If the set point of the blood leak sensor does not match the set point in the calibration record. |  | HD blood leak sensor set point set failure | If the set point of the blood leak sensor does not match the set point in the calibration record. | 66 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_SET_POINT_SET_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD blood leak sensor set point set failure | HD | HD POST: BLD Config Set Point
No cartridge self-test (procedural) |  |  | SRSHD 1511 |  |  | Cartridge has been detected.\n- Open door\n- Remove cartridge\n- Close door\n- Resume | HD blood pump occlusion self-test failure alarm |  | If the blood pump pressure > 5000 |  | HD blood pump occlusion self-test failure alarm | If the blood pump pressure > 5000 | 67 | ALARM_ID_HD_BP_OCCLUSION_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 700 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD blood pump occlusion self-test failure alarm | HD | Process: Pre-Tx Remove Cartridge
agree to change rank and re-assess similar to temperature alarm Blaine |  |  | SRSHD 1426 |  |  | Dialysate is not yet ready.\n- Treatment is paused and will be ready to resume when \ndialysate is ready for use. | HD active reservoir recirculation out of range |  | If the recirculation level has exceeded the max allowed |  | HD active reservoir recirculation out of range | If the recirculation level has exceeded the max allowed | 68 | ALARM_ID_HD_ACTIVE_RESERVOIR_RECIRCULATION_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 421 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD active reservoir recirculation out of range | HD | Dial. Gen: Dialysate not ready
x | PRS 398 | SA 368 | SRSHD 532 |  |  | Call service to schedule a repair. | HD blood leak sensor invalid calibration record |  | If the BLD data that was received from NVDataMgmt is NOT legitimate. |  | HD blood leak sensor invalid calibration record | If the BLD data that was received from NVDataMgmt is NOT legitimate. | 69 | ALARM_ID_HD_BLOOD_LEAK_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD blood leak sensor invalid calibration record | HD | HD POST: BLD Cal
This should be in-treatment alarm? Not machine alarm? See alarm ID 275 for machine alarm (valid range check) agree - make this in-treatment alarms not faults Blaine to update - not a fault Sean to update the range to what is in the SRS |  |  | SRSHD 737 |  |  | Cartridge has been detected.\n- Open door\n- Remove cartridge\n- Close door\n- Resume | HD arterial pressure self-test failure alarm |  | If arterial pressure < -300 or > 100 mmHg |  | HD arterial pressure self-test failure alarm | If arterial pressure < -300 or > 100 mmHg | 70 | ALARM_ID_HD_ARTERIAL_PRESSURE_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 700 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD arterial pressure self-test failure alarm | HD | Process: Pre-Tx Remove Cartridge
This should be in-treatment alarm? Not machine alarm? See alarm ID 275 for machine alarm (valid range check) agree - make this in-treatment alarms not faults Blaine to update - not a fault Sean to update the range to what is in the SRS |  |  | SRSHD 738 |  |  | Cartridge has been detected.\n- Open door\n- Remove cartridge\n- Close door\n- Resume | HD venous pressure self-test failure alarm |  | If the venous pressure < -100 or > 600 mmHg |  | HD venous pressure self-test failure alarm | If the venous pressure < -100 or > 600 mmHg | 71 | ALARM_ID_HD_VENOUS_PRESSURE_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 700 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD venous pressure self-test failure alarm | HD | Process: Pre-Tx Remove Cartridge
x | PRS 400 |  | SRSHD 1126 |  |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | HD No load cell data message receive at least once every 2 seconds |  | No new load cell data for 2 seconds. |  | HD No load cell data message receive at least once every 2 seconds | No new load cell data for 2 seconds. | 72 | ALARM_ID_HD_NEW_LOAD_CELL_DATA_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD No load cell data message receive at least once every 2 seconds | HD | Loss of DG Comm (LC)
x | PRS 400 |  | SRSHD 1573 |  |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | HD No dialysate temperature data message received at least once every 2 seconds |  | No new dialysate temperature data for 2 seconds. |  | HD No dialysate temperature data message received at least once every 2 seconds | No new dialysate temperature data for 2 seconds. | 73 | ALARM_ID_HD_NEW_DIALYSATE_TEMP_DATA_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD No dialysate temperature data message receive at least once every 2 seconds | HD | Loss of DG Comm (TDI/TRO)
x | PRS 863 | SA 364 |  | SRSDG 600 |  | Call service to schedule a repair. | DG inlet UV reactor is on while there is no flow |  | If the inlet UV reactor is on while VPi is closed for a certain period of time. |  | DG inlet UV reactor is on while there is no flow | If the inlet UV reactor is on while VPi is closed for a certain period of time. | 74 | ALARM_ID_DG_INLET_UV_REACTOR_ON_WITH_NO_FLOW |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG inlet UV reactor is on with no flow | DG | DG Fault: UVi On No Flow
x | PRS 494 | SA 469 | SRSHD 1488 |  |  | Call service to schedule a repair. | HD syringe pump self-test failure alarm |  | If the calibration data that was received from NVDataMgmt and the DAC voltage with syringe retracted are not within tol. |  | HD syringe pump self-test failure alarm | If the calibration data that was received from NVDataMgmt and the DAC voltage with syringe retracted are not within tol. | 75 | ALARM_ID_HD_SYRINGE_PUMP_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD syringe pump self-test failure alarm | HD | HD Fault: Syringe DAC
Sean to update code | PRS 494 | SA 217 SA 218 SA 219 | SRSHD 635 SRSHD 1678 SRSHD 1681 |  |  | Call service to schedule a repair. | HD monitored voltage is out of range |  | If the HD monitored voltages are out of range for a certain period of time. |  | HD monitored voltage is out of range | If the HD monitored voltages are out of range for a certain period of time. | 76 | ALARM_ID_HD_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD monitored voltage is out of range | HD | HD Fault: Voltage Monitor
Sean to update code | PRS 863 | SA 295 SA 296 SA 297 |  | SRSDG 46 SRSDG 866 SRSDG 1088 |  | Call service to schedule a repair. | DG monitored voltage is out of range |  | If the DG monitored voltages are out of range for a certain period of time. |  | DG monitored voltage is out of range | If the DG monitored voltages are out of range for a certain period of time. | 77 | ALARM_ID_DG_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG monitored voltage is out of range | DG | DG Fault: Voltage Monitor
x | PRS 494 | SA 275 | SRSHD 1324 |  |  | Call service to schedule a repair. | HD syringe pump direction (from encoder) error |  | If the measured encoder direction of the syringe pump does not agree with the expected direction |  | HD syringe pump direction (from encoder) error | If the measured encoder direction of the syringe pump does not agree with the expected direction | 78 | ALARM_ID_HD_SYRINGE_PUMP_ENCODER_DIRECTION_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD syringe pump direction (from encoder) error | HD | HD Fault: Syringe Pump Enc Dir
x | PRS 494 | SA 489 | SRSHD 1507 |  |  | Call service to schedule a repair. | HD syringe pump direction (from controller) error |  | If the measured direction of the syringe pump motor controller does not agree with the expected direction |  | HD syringe pump direction (from controller) error | If the measured direction of the syringe pump motor controller does not agree with the expected direction | 79 | ALARM_ID_HD_SYRINGE_PUMP_CONTROLLER_DIRECTION_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD syringe pump direction (from controller) error | HD | HD Fault: Syringe Pump Drv Dir
x | PRS 494 | SA 272 | SRSHD 1317 |  |  | Call service to schedule a repair. | HD syringe pump fault reported by FPGA |  | If the FPGA reported the syringe pump fault |  | HD syringe pump fault reported by FPGA | If the FPGA reported the syringe pump fault | 80 | ALARM_ID_HD_SYRINGE_PUMP_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD syringe pump fault reported by FPGA | HD | HD Fault: Syringe Pump Driver
x | PRS 494 | SA 488 | SRSHD 1330 |  |  | Call service to schedule a repair. | HD syringe pump over travel error |  | If the syringe pump position is deemed to have overtraveled |  | HD syringe pump over travel error | If the syringe pump position is deemed to have overtraveled | 81 | ALARM_ID_HD_SYRINGE_PUMP_OVER_TRAVEL_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD syringe pump over travel error | HD | HD Fault: Syringe Pump OT
x | PRS 494 | SA 277 | SRSHD 1506 |  |  | Call service to schedule a repair. | HD syringe pump DAC write failure |  | If the syringe pump DAC write error is reported by the FPGA |  | HD syringe pump DAC write failure | If the syringe pump DAC write error is reported by the FPGA | 82 | ALARM_ID_HD_SYRINGE_PUMP_DAC_WRITE_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD syringe pump DAC write failure | HD | HD Fault: Syringe Pump DAC Err
x | PRS 494 | SA 415 | SRSHD 1328 |  |  | Call service to schedule a repair. | HD syringe pump is running while the BP is off |  | If the syringe pump is on while the blood pump is off |  | HD syringe pump is running while the BP is off | If the syringe pump is on while the blood pump is off | 83 | ALARM_ID_HD_SYRINGE_PUMP_RUNNING_WHILE_BP_OFF_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD syringe pump is running while the BP is off | HD | HD Fault: Syringe Pump On Err
x | PRS 863 | SA 286 |  | SRSDG 815 |  | Call service to schedule a repair. | DG RTC set year is invalid |  | If DG RTC set year is < 2021. |  | DG RTC set year is invalid | If DG RTC set year is < 2021. | 84 | ALARM_ID_DG_SET_RTC_YEAR_INVALID |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG set RTC year is invalid | DG | DG POST: HD RTC Time
revisit - PRS has this as low priority - why did SRS make this high priority? Probably needs to be high priority as blood pump needs to stop - update PRS to make high priority | [5 empty] | Cartridge Lock Lever is not closed.\n- Open front door and push lever all the way down.\n- Close door and press RESUME button to continue. | HD pump track lever is not fully latched |  | If the pump track lever is detected as unlatched when it should be. |  | HD pump track lever is not fully latched | If the pump track lever is detected as unlatched when it should be. | 85 | ALARM_ID_HD_PUMP_TRACK_LATCH_OPENED |  | ALARM_PRIORITY_HIGH | 209 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD pump track latch opened alarm | HD | Blood Stop: Cartridge Lever Open
x | PRS 398 | SA 199 | SRSHD 1479 |  |  | Call service to schedule a repair. | HD RTC set year is invalid |  | If HD RTC set year is < 2021. |  | HD RTC set year is invalid | If HD RTC set year is < 2021. | 86 | ALARM_ID_HD_SET_RTC_YEAR_INVALID |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD set RTC year is invalid | HD | HD POST: HD RTC Time
x | PRS 863 | SA 369 |  | SRSDG 929 |  | Call service to schedule a repair. | DG heating parameters invalid calibration record |  | If the heating paramters' non-volatile record's CRC does not match the calculated CRC. |  | DG heating parameters invalid calibration record | If the heating paramters' non-volatile record's CRC does not match the calculated CRC. | 87 | ALARM_ID_DG_HEATING_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG heating invalid calibration record | DG | DG POST: Heat Cal
[6 empty] | Call service to schedule a repair. | DG concentrate pump hall sensor out of range. |  | Hall sensor pulse to pulse timing is < 1ms (speed too fast). |  | DG concentrate pump hall sensor out of range. | Hall sensor pulse to pulse timing is < 1ms (speed too fast). | 88 | ALARM_ID_DG_CONC_PUMP_HALL_SENSOR_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG concentrate pumps hall sensor out of range | DG | DG Fault: CP1 CP2 Speed Range
x | PRS 863 | SA 365 |  | SRSDG 599 |  | Call service to schedule a repair. | DG outlet UV reactor is on while there is no flow |  | If the outlet UV reactor is on while the RO flow is less than 200 mL/min. |  | DG outlet UV reactor is on while there is no flow | If the outlet UV reactor is on while the RO flow is less than 200 mL/min. | 89 | ALARM_ID_DG_OUTLET_UV_REACTOR_ON_WITH_NO_FLOW |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG outlet UV reactor on with no flow | DG | DG Fault: UVo On No Flow
x | PRS 863 | SA 307 SA 308 |  | SRSDG 604 |  | Call service to schedule a repair. | DG load cells A1/B1 FPGA fault |  | If the FPGA does not report fresh data for a certain period of time. |  | DG load cells A1/B1 FPGA fault | If the FPGA does not report fresh data for a certain period of time. | 90 | ALARM_ID_DG_LOAD_CELL_A1_B1_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells A1/B1 FPGA fault | DG | DG Fault: LC A1/B1 Comm
x | PRS 400 |  | SRSHD 1408 |  |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | HD No dialysate flow data received in the last 3 seconds |  | If not receiving new dialysate flow readings in expected interval (3 seconds) |  | HD No dialysate flow data received in the last 3 seconds | If not receiving new dialysate flow readings in expected interval (3 seconds) | 91 | ALARM_ID_HD_DIALYSATE_FLOW_DATA_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD No dialysate flow data receive in the last 3 seconds | HD | Loss of DG Comm (FMD)
Remove (redundant with SRSDG 913) | [12 empty] | 92
[13 empty] | 92 | ALARM_ID_AVAILABLE_1 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Available for use
x | PRS 863 | SA 380 |  | SRSDG 430 |  | Call service to schedule a repair. | DG outlet primary conductivity sensor out of range |  | If the outlet primary conductivity sensor, CPo is greater than 30 uS/cm for a certain period of time. |  | DG outlet primary conductivity sensor out of range | If the outlet primary conductivity sensor, CPo is greater than 30 uS/cm for a certain period of time. | 93 | ALARM_ID_DG_OUTLET_PRIMARY_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG outlet primary conductivity out of range | DG | DG Fault: RO Permeate Cond High
x | PRS 863 | SA 303 SA 304 SA 305 SA 306 |  | SRSDG 629 |  | Call service to schedule a repair. | DG pressure sensor out of range |  | If either of the pressure sensors is < 0 psia or > 200 psia for DG pressure sensors. |  | DG pressure sensor out of range | If either of the pressure sensors is < 0 psia or > 200 psia for DG pressure sensors. | 94 | ALARM_ID_DG_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG pressure out of range | DG | DG Fault: Pressure Sensor Range
x | PRS 863 | SA 486 |  | SRSDG 907 |  | Call service to schedule a repair. | DG watchdog expired |  | If the watchdog signal is low while not in POST. |  | DG watchdog expired | If the watchdog signal is low while not in POST. | 95 | ALARM_ID_DG_WATCHDOG_EXPIRED |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG watchdog expired | DG | DG Fault: WD Error
x | PRS 406 |  |  | SRSDG 410 |  | Inlet water temperature is high.\n- Treatment is paused\n- Check the incoming water supply. | DG inlet water temperature out of range |  | If the inlet water temperature is > 37 C. |  | DG inlet water temperature out of range | If the inlet water temperature is > 37 C. | 96 | ALARM_ID_DG_INLET_WATER_TEMPERATURE_IN_HIGH_RANGE |  | ALARM_PRIORITY_MEDIUM | 432 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water temperature in high range | DG | Dial. Gen: Inlet Water Temp High
Check that software behavior and PRS behavior are aligned | PRS 543 |  |  | SRSDG 397 |  | Dialysate conductivity is out of range.\n- Check acid and bicarb concentrate is correct and not empty. | DG fill conductivity out of range |  | If the average acid or bicard conductivity is > 5% of the normal acid and bicarb conductivity values. |  | DG fill conductivity out of range | If the average acid or bicard conductivity is > 5% of the normal acid and bicarb conductivity values. | 97 | ALARM_ID_DG_FILL_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 420 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | DG fill conductivity out of range | DG | Dial. Gen: Dialysate Cond.
x | PRS 872 | SA 227 SA 228 | SRSHD 1395 |  |  | Call service to schedule a repair. | HD battery or battery charger communication fault |  | If battery I2C communication fault persistence count > 5. |  | HD battery or battery charger communication fault | If battery I2C communication fault persistence count > 5. | 98 | ALARM_ID_HD_BATTERY_COMM_FAULT |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD battery communication fault | HD | HD Alarm (Low): Battery Com
x | PRS 494 | SA 274 | SRSHD 1683 |  |  | Call service to schedule a repair. | HD syringe pump stall alarm |  | If syringe pump measured rate < stall speed threshold (0.05) and the ramp up has been restarted 3 times after stall has been detected. |  | HD syringe pump stall alarm | If syringe pump measured rate < stall speed threshold (0.05) and the ramp up has been restarted 3 times after stall has been detected. | 99 | ALARM_ID_HD_SYRINGE_PUMP_STALL |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD syringe pump stall alarm | HD | HD Fault: Syringe Pump Stall
x | PRS 494 | SA 416 | SRSHD 1679 |  |  | Call service to schedule a repair. | HD no cartridge self-test timeout |  | If during the no cartridge self-tests state, the elapsed self-test time > maximum no cartridge self test time (30secs) |  | HD no cartridge self-test timeout | If during the no cartridge self-tests state, the elapsed self-test time > maximum no cartridge self test time (30secs) | 100 | ALARM_ID_HD_NO_CART_SELF_TEST_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD no cartridge self-test timeout | HD | HD Fault: Timeout No Cart ST
[6 empty] | Call service to schedule a repair. | UI POST FW Communication failure |  | If the HD is not communicating with the UI during POST. |  | UI POST FW Communication failure | If the HD is not communicating with the UI during POST. | 101 | ALARM_ID_UI_POST_HD_COMM |  | ALARM_PRIORITY_HIGH | 0 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | UI POST HD communication failure | UI | HD POST: UI-Proc Comm
cleaning | PRS 863 | SA 482 |  | SRSDG 856 SRSDG 930 |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | DG heat disinfection temperature gradient out of range |  | If temperature gradient between Tpo and Thd exceeds maximum 15°C. |  | DG heat disinfection temperature gradient out of range | If temperature gradient between Tpo and Thd exceeds maximum 15°C. | 102 | ALARM_ID_DG_HEAT_DISINFECT_TEMP_GRAD_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG heat disinfect temperature gradient out of range | DG | Cleaning: Tpo THD Temp Delta
x | PRS 398 | SA 196 | SRSHD 1059 |  |  | Call service to schedule a repair. | HD non-volatile calibration group invalid record CRC |  | If the computed CRC of the calibration group record is not equal to the record CRC. |  | HD non-volatile calibration group invalid record CRC | If the computed CRC of the calibration group record is not equal to the record CRC. | 103 | ALARM_ID_HD_NVDATAMGMT_CAL_GROUP_RECORD_CRC_INVALID |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD invalid calibration CRC | HD | HD POST: Cal NVRAM CRC
x | PRS 494 | SA 478 | SRSHD 1115 |  |  | Call service to schedule a repair. | Air trap level sensors reporting illegal combination of air/fluid |  | If the air trap level readings are not valid for more than 2 seconds. |  | Air trap level sensors reporting illegal combination of air/fluid | If the air trap level readings are not valid for more than 2 seconds. | 104 | ALARM_ID_HD_AIR_TRAP_ILLEGAL_LEVELS |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD air trap level sensors reporting illegal combination of air/fluid | HD | HD Fault: Air Trap Sensor
x | PRS 863 | SA 283 |  | SRSDG 608 |  | Call service to schedule a repair. | DG non-volatile system group invalid record CRC |  | If the DG NVDATAMGMT CAL group calculated CRC does not mached the expected CRC. |  | DG non-volatile system group invalid record CRC | If the DG NVDATAMGMT CAL group calculated CRC does not mached the expected CRC. | 105 | ALARM_ID_DG_NVDATAMGMT_CAL_GROUP_RECORD_CRC_INVALID |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid calibration CRC | DG | DG POST: Cal NVRAM CRC
x | PRS 863 | SA 369 |  | SRSDG 820 SRSDG 821 SRSDG 1087 |  | Call service to schedule a repair. | DG flow sensors calibration error |  | If the calculated CRC of the flow sensors calibration record does not match the calibration records' CRC. |  | DG flow sensors calibration error | If the calculated CRC of the flow sensors calibration record does not match the calibration records' CRC. | 106 | ALARM_ID_DG_FLOW_SENSORS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG dialysate flow sensor invalid calibration record | DG | DG POST: Flow Sensor Cal
No SRS ID - need to add then assess Move to DG machine alarm? (allow properties table to be aligned) Sean - to revisit code to assess | [5 empty] | Call service to schedule a repair. | HD reports DG restarted fault |  | If DG reports being in any mode other than idle, drain, fill or fault during pre-tx, tx, or post-tx modes - means DG restarted unexpectedly and is no longer in sync w/ HD. |  | HD reports DG restarted fault | If DG reports being in any mode other than idle, drain, fill or fault during pre-tx, tx, or post-tx modes - means DG restarted unexpectedly and is no longer in sync w/ HD. | 107 | ALARM_ID_HD_DG_RESTARTED_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD reports DG restarted fault | HD | DG Fault: DG Restarted
x | PRS 494 | SA 276 | SRSHD 1318 |  |  | Call service to schedule a repair. | HD syringe pump ADC error |  | If the ADC read is not fresh after POST init and the ADC read counter has not changed. |  | HD syringe pump ADC error | If the ADC read is not fresh after POST init and the ADC read counter has not changed. | 108 | ALARM_ID_HD_SYRINGE_PUMP_FPGA_ADC_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD syringe pump ADC FPGA fault | HD | HD Fault: Syringe Pump ADC Com
x | PRS 494 | SA 490 | SRSHD 580 |  |  | Call service to schedule a repair. | HD syringe pump volume check error |  | If the volume delivered is off by more than 5% or 0.1mL. |  | HD syringe pump volume check error | If the volume delivered is off by more than 5% or 0.1mL. | 109 | ALARM_ID_HD_SYRINGE_PUMP_VOLUME_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD syringe pump volume check error | HD | HD Fault: Syringe Pump Vol
x | PRS 494 | SA 274 | SRSHD 579 |  |  | Call service to schedule a repair. | HD syringe pump speed check error |  | If the syringe measured rate is off by more than 5% or 0.1mL. |  | HD syringe pump speed check error | If the syringe measured rate is off by more than 5% or 0.1mL. | 110 | ALARM_ID_HD_SYRINGE_PUMP_SPEED_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD syringe pump speed check error | HD | HD Fault: Syringe Pump Speed
x | PRS 494 | SA 274 | SRSHD 1333 |  |  | Call service to schedule a repair. | HD syringe pump not stopped in off state error |  | If the syringe pump position is not equal to its last position. |  | HD syringe pump not stopped in off state error | If the syringe pump position is not equal to its last position. | 111 | ALARM_ID_HD_SYRINGE_PUMP_NOT_STOPPED_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD syringe pump not stopped in off state error | HD | HD Fault: Syringe Pump Speed
x | PRS 863 | SA 312 SA 313 |  | SRSDG 953 |  | Call service to schedule a repair. | DG load cells A2/B2 FPGA fault |  | If the FPGA does not report fresh data for a certain period of time. |  | DG load cells A2/B2 FPGA fault | If the FPGA does not report fresh data for a certain period of time. | 112 | ALARM_ID_DG_LOAD_CELL_A2_B2_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells A1/B1 FPGA fault | DG | DG Fault: LC A2/B2 Comm
x | PRS 863 | SA 341 |  | SRSDG 890 |  | Call service to schedule a repair. | DG drain pump current out of range |  | If the DG drain pump current > 2.2 A when it is in open loop or if the current > 0.1 A when it is off. |  | DG drain pump current out of range | If the DG drain pump current > 2.2 A when it is in open loop or if the current > 0.1 A when it is off. | 113 | ALARM_ID_DG_DRAIN_PUMP_CURRENT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG Drain pump current out of range | DG | DG Fault: DRP Current
move to wet self-test? | PRS 494 | SA 479 | SRSHD 1532 |  |  | Self-test has detected a problem with the air bubble sensor\n- Check cartridge installation on to machine\n- If problem persists, terminate pre-treatment and \nrestart with a new cartridge. | HD venous air bubble detector self-test failure |  | If the venous air bubble self test did not detect bubble for more than 500 msec. |  | HD venous air bubble detector self-test failure | If the venous air bubble self test did not detect bubble for more than 500 msec. | 114 | ALARM_ID_HD_VENOUS_BUBBLE_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD venous air bubble detector self-test failure | HD | Process: Pre-Tx Wet Self-Test
Sean - update code to remove temperatures from code that are no longer needed (see updated SRSDG 224) | PRS 863 | SA 321 SA 322 SA 323 SA 324 SA 328 SA 332 SA 336 SA 298 SA 460 |  | SRSDG 224 |  | Call service to schedule a repair. | DG temperature sensor out of range |  | If the measured temperature < 0 C or > 80 C for the non-fluid inline sensors and > 120 C for the fluid sensors. |  | DG temperature sensor out of range | If the measured temperature < 0 C or > 80 C for the non-fluid inline sensors and > 120 C for the fluid sensors. | 115 | ALARM_ID_DG_TEMPERATURE_SENSOR_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG temperature sensor out of range | DG | DG Fault: Temp Sensor Range
[6 empty] | SD-Card Failure \n- Power cycle the system.\n- Call service to schedule a repair if problem persists. | SD card is non-operational |  | SD card is non-operational |  | SD card is non-operational | SD card is non-operational | 116 | ALARM_ID_HD_UI_SDCARD_FAILURE |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD UI SD card failure | UI HD | HD Alarm (Low): SD-Card Error
|  |  |  | SRSDG 909 |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | System message that required acknowledgment was not acknowledged |  | If no retries left in the pending acknowlegement list. |  | System message that required acknowledgment was not acknowledged | If no retries left in the pending acknowlegement list. | 117 | ALARM_ID_DG_CAN_MESSAGE_NOT_ACKED |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CAN message not acked | DG | Loss of HD Comm (HD ACK)
x | PRS 863 | SA 458 |  | SRSDG 43 |  | Call service to schedule a repair. | DG RTC configuration error |  | If the RTC configuration registers report an error. |  | DG RTC configuration error | If the RTC configuration registers report an error. | 118 | ALARM_ID_DG_RTC_CONFIG_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RTC config error | DG | DG Fault: RTC Config
x | PRS 874 | SA 499 |  | SRSDG 908 |  | Call service to schedule a repair. | DG RTC battery low |  | If the RTC battery bit report a low alarm. |  | DG RTC battery low | If the RTC battery bit report a low alarm. | 119 | ALARM_ID_DG_RTC_BATTERY_LOW |  | ALARM_PRIORITY_LOW | 851 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | DG RTC battery low error | DG | DG Alarm (Low): RTC Bat Low
|  |  | SRSHD 777 |  |  | Self-test has detected a bubble in the venous bubble detector\n- The cartridge should be primed\n- Check cartridge installation on to machine\n- If problem persists, terminate pre-treatment and \nrestart with a new cartridge. | HD pre-treatment wet prime test failure |  | If the HD pre-treatment wet self test failed. |  | HD pre-treatment wet prime test failure | If the HD pre-treatment wet self test failed. | 120 | ALARM_ID_HD_PRE_TREATMENT_WET_PRIME_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode wet self-test prime check failure | HD | Process: Pre-Tx Wet Self-Test
x | PRS 863 | SA 363 |  | SRSDG 611 |  | Call service to schedule a repair. | DG main primary heater voltage out of range |  | If the main primary heater's voltage is out of range. |  | DG main primary heater voltage out of range | If the main primary heater's voltage is out of range. | 121 | ALARM_ID_DG_MAIN_PRIMARY_HEATER_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG main primary heater voltage out of range | DG | DG Fault: Prim Heater Lg Control
x | PRS 863 | SA 362 |  | SRSDG 612 |  | Call service to schedule a repair. | DG small primary heater voltage out of range |  | If the small primary heater's voltage is out of range. |  | DG small primary heater voltage out of range | If the small primary heater's voltage is out of range. | 122 | ALARM_ID_DG_SMALL_PRIMARY_HEATER_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG small primary heater voltage out of range | DG | DG Fault: Prim Heater Sm Control
x | PRS 863 | SA 361 |  | SRSDG 613 |  | Call service to schedule a repair. | DG trimmer heater voltage out of range |  | If the trimmer voltage is out of range. |  | DG trimmer heater voltage out of range | If the trimmer voltage is out of range. | 123 | ALARM_ID_DG_TRIMMER_HEATER_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG trimmer heater voltage out of range | DG | DG Fault: Trimmer Heater Control
x | PRS 390 |  | SRSHD 954 |  |  | Blood pump has been off for too long.\n- Treatment must be terminated. Blood clot risk.\n- NO RINSEBACK ALLOWED. | HD end of treatment alarm (high priority - user still not acting to end treatment) |  | In treatment end paused state for > 5 minutes. |  | HD end of treatment alarm (high priority - user still not acting to end treatment) | In treatment end paused state for > 5 minutes. | 124 | ALARM_ID_HD_END_OF_TREATMENT_HIGH |  | ALARM_PRIORITY_HIGH | 11 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | HD end of treatment alarm (high priority - user still not acting to end treatment) | HD | Blood Stop: Blood Clot Risk
x | PRS 390 |  | SRSHD 948 |  |  | Blood pump has been off for too long.\n- Treatment must be terminated. Blood clot risk.\n- NO RINSEBACK ALLOWED. | Blood sitting too long after treatment stopped by user alarm (> 5 min) |  | If in treatment stop sub mode when blood is in the tubing, the blood sitting counter exceeded 5 minutes. |  | Blood sitting too long after treatment stopped by user alarm (> 5 min) | If in treatment stop sub mode when blood is in the tubing, the blood sitting counter exceeded 5 minutes. | 125 | ALARM_ID_HD_TREATMENT_STOPPED_NO_RINSEBACK |  | ALARM_PRIORITY_HIGH | 11 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood sitting too long after treatment stopped by user alarm (>5 min) | HD | Blood Stop: Blood Clot Risk
x | PRS 387 |  | SRSHD 925 |  |  | A blood leak has been detected. Treatment must end. | HD blood leak detected alarm |  | If the blood leak is detected based on the blood leak status reported by the sensor. |  | HD blood leak detected alarm | If the blood leak is detected based on the blood leak status reported by the sensor. | 126 | ALARM_ID_HD_BLOOD_LEAK_DETECTED |  | ALARM_PRIORITY_HIGH | 150 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | HD blood leak detected alarm | HD | Blood Stop: Dialyzer Blood Leak
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 911 |  |  | Venous pressure is low.\n- Check to be sure your blue venous connections are secure.\n- Check that your venous access is connected and secured \nproperly. | Venous pressure too low during treatment |  | If the venous pressure < TREATMENT_PARAM_VEN_PRESSURE_LOW_LIMIT. Or expanded pressure limits based on mode of operation. |  | Venous pressure too low during treatment | If the venous pressure < TREATMENT_PARAM_VEN_PRESSURE_LOW_LIMIT. Or expanded pressure limits based on mode of operation. | 127 | ALARM_ID_HD_VENOUS_PRESSURE_LOW |  | ALARM_PRIORITY_HIGH | 200 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | HD venous pressure too low during treatment | HD | Blood Stop: Venous Pres. Low
x | PRS 494 | SA 506 | SRSHD 919 |  |  | Venous air bubble detected indicating failure of air trap \ncontrol system.\n- Treatment must be terminated without option for rinseback.\n- Check access for dislodgement. | HD venous air bubble detected alarm |  | If bubbles are detected during the treatment. |  | HD venous air bubble detected alarm | If bubbles are detected during the treatment. | 128 | ALARM_ID_HD_VENOUS_BUBBLE_DETECTED |  | ALARM_PRIORITY_HIGH | 12 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | HD venous air bubble detected alarm | HD | Blood Stop: Venous Air Bubble
x | PRS 494 | SA 507 | SRSHD 921 | [9 empty] | 129 | ALARM_ID_AVAILABLE_17 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | Available for use
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 913 |  |  | Venous pressure is high.\n- Check for access or patient discomfort.\n- Resolve access issue if applicable.\n- Check that there are no kinks or clots in the tubing.\n- Unkink the lines if applicable.\n- End treatment and do not return blood if clots are present. | Venous pressure too high during treatment |  | If the venous pressure > TREATMENT_PARAM_VEN_PRESSURE_HIGH_LIMIT. Or expanded pressure limits based on mode of operation. |  | Venous pressure too high during treatment | If the venous pressure > TREATMENT_PARAM_VEN_PRESSURE_HIGH_LIMIT. Or expanded pressure limits based on mode of operation. | 130 | ALARM_ID_HD_VENOUS_PRESSURE_HIGH |  | ALARM_PRIORITY_HIGH | 201 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | HD venous pressure too high during treatment | HD | Blood Stop: Venous Pres. High
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 915 |  |  | Arterial pressure is low.\n- Check that the access is properly connected and secured.\n- Resolve access issue if applicable.\n- Check to be sure the red arterial connections are secure.\n- Check that there are no kinks or clots in the tubing.\n- Unkink lines if applicable.\n- End treatment and do not return blood if clots are present. | Arterial pressure too low during treatment |  | If the arterial pressure < TREATMENT_PARAM_ART_PRESSURE_LOW_LIMIT. Or expanded pressure limits based on mode of operation. |  | Arterial pressure too low during treatment | If the arterial pressure < TREATMENT_PARAM_ART_PRESSURE_LOW_LIMIT. Or expanded pressure limits based on mode of operation. | 131 | ALARM_ID_HD_ARTERIAL_PRESSURE_LOW |  | ALARM_PRIORITY_HIGH | 202 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | HD arterial pressure too low during treatment | HD | Blood Stop: Arterial Pres. Low
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 917 |  |  | Arterial pressure is high.\n- Check that the access is properly connected and secured.\n- Resolve access issue if applicable.\n- Check to be sure your red arterial connections are secure.\n- Check that there are no kinks or clots in the tubing.\n- Unkink lines if applicable.\n- End treatment and do not return blood if clots are present. | Arterial pressure too high during treatment |  | If the arterial pressure > TREATMENT_PARAM_ART_PRESSURE_HIGH_LIMIT. Or expanded pressure limits based on mode of operation. |  | Arterial pressure too high during treatment | If the arterial pressure > TREATMENT_PARAM_ART_PRESSURE_HIGH_LIMIT. Or expanded pressure limits based on mode of operation. | 132 | ALARM_ID_HD_ARTERIAL_PRESSURE_HIGH |  | ALARM_PRIORITY_HIGH | 203 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | HD arterial pressure too high during treatment | HD | Blood Stop: Arterial Pres. High
x | PRS 863 | SA 419 |  | SRSDG 328 |  | Fluid leak detected in Dialysate Generation Delivery Device.\n-Treatment must be terminated.\n- Shut off water source and disconnect water and \nelectrical connections from system.\n- Call service to schedule a repair. | DG fluid leak detected alarm |  | If the fluid leak is detected as reported by the fluid leak detector. |  | DG fluid leak detected alarm | If the fluid leak is detected as reported by the fluid leak detector. | 133 | ALARM_ID_DG_FLUID_LEAK_DETECTED |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | DG fluid leak detected alarm | DG | DG Fault: Fluid Leak in DG
x | PRS 409 |  | SRSHD 927 |  |  | Fluid leak detected in Hemodialysis Delivery device.\n- If in treatment, BLOOD STOPPED.\n-Identify and correct the cause of the spill.\n- Clean up the spill and dry leak sensor.\n- If in treatment, resume treatment. | HD fluid leak detected alarm |  | If the fluid leak is detected as reported by the fluid leak detector. |  | HD fluid leak detected alarm | If the fluid leak is detected as reported by the fluid leak detector. | 134 | ALARM_ID_HD_FLUID_LEAK_DETECTED |  | ALARM_PRIORITY_HIGH | 208 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | HD fluid leak detected alarm | HD | Blood Stop: Fluid Leak in HD
x | PRS 411 |  | SRSHD 938 |  |  | Excessive movement detected with Hemodialysis Delivery \ndevice.\n- Confirm the device is stationary and level before continuing. | HD experienced large acceleration |  | If the system experienced excessive shock > 2.5 g. |  | HD experienced large acceleration | If the system experienced excessive shock > 2.5 g. | 135 | ALARM_ID_HD_SHOCK |  | ALARM_PRIORITY_HIGH | 210 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD experienced large acceleration | HD | Blood Stop: HD Shock Detected
x | PRS 686 |  |  | SRSDG 327 |  | Excessive movement detected with Dialysis Generation \ndevice.\n- Treatment is paused.\n- Confirm the device is stationary and level before continuing. | DG experienced large acceleration |  | If the system experienced excessive shock > 2.5 g. |  | DG experienced large acceleration | If the system experienced excessive shock > 2.5 g. | 136 | ALARM_ID_DG_SHOCK |  | ALARM_PRIORITY_MEDIUM | 440 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG experienced large acceleration | DG | Dial. Gen: Shock-Vibration
x | PRS 412 |  | SRSHD 936 |  |  | Excessive tilt detected with Hemodialysis Delivery device.\n- Confirm the device is stationary and level before continuing. | HD tilt exceeds maximum |  | If the system experienced excessive tilt x and y axis angle > 7 deg for more than 1 sec. |  | HD tilt exceeds maximum | If the system experienced excessive tilt x and y axis angle > 7 deg for more than 1 sec. | 137 | ALARM_ID_HD_EXCESSIVE_TILT |  | ALARM_PRIORITY_HIGH | 211 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD tilt exceeds maximum | HD | Blood Stop: HD Tilted
x | PRS 687 |  |  | SRSDG 326 |  | Excessive tilt detected with Dialysis Generation device.\n- Treatment is paused.\n-Confirm the device is stationary and level before continuing. | DG tilt exceeds maximum |  | If the system experienced excessive tilt x and y axis angle > 7 deg for more than 1 sec. |  | DG tilt exceeds maximum | If the system experienced excessive tilt x and y axis angle > 7 deg for more than 1 sec. | 138 | ALARM_ID_DG_EXCESSIVE_TILT |  | ALARM_PRIORITY_MEDIUM | 441 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG tilt exceeds maximum | DG | Dial. Gen: Tilt
x | PRS 823 |  | SRSHD 1593 |  |  | Power has been lost. Check all electrical connections.\n- The display and CPU will run off of the battery until \npower returns. | HD AC power lost alarm |  | The HD Software shall trigger an alarm if an AC power loss is detected while the HD system is in standby, service, or treatment parameters mode. |  | HD AC power lost alarm | The HD Software shall trigger an alarm if an AC power loss is detected while the HD system is in standby, service, or treatment parameters mode. | 139 | ALARM_ID_HD_AC_POWER_LOST |  | ALARM_PRIORITY_HIGH | 852 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD AC power lost alarm | HD | Loss of AC Power
x | PRS 400 |  | SRSHD 1124 |  |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | Dialysate generator communication timeout |  | If DG communication timeout is detected for more than 2 seconds. |  | Dialysate generator communication timeout | If DG communication timeout is detected for more than 2 seconds. | 140 | ALARM_ID_HD_DG_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD DG communication timeout | HD | Loss of DG Comm (Timeout)
|  |  | SRSHD 1169 |  |  | Failure to fill the Air Trap.\n- Treatment must be terminated without option for rinseback.\n- Check access for dislodgement. | Air trap fill timeout during treatment |  | If the air trap fill timeout during treatment. |  | Air trap fill timeout during treatment | If the air trap fill timeout during treatment. | 141 | ALARM_ID_HD_AIR_TRAP_FILL_DURING_TREATMENT |  | ALARM_PRIORITY_HIGH | 13 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | HD air trap fill timeout during treatment | HD | Blood Stop: Air Trap Fill
x | PRS 496 |  | SRSHD 1361 SRSHD 1362 |  |  | Occlusion detected after the blood pump.\n- Check for clots or kinks around the dialyzer and in the \nvenous line.\n- Resolve access issue if applicable.\n- End treatment and do not return blood if clots are present.\n- Check for access or patient discomfort.\n- Unkink the lines if applicable. | Blood pump occlusion detected |  | If blood occlusion is detected > OCCLUSION_THRESHOLD_OFFSET + bloodPumpOcclusionAfterCartridgeInstall |  | Blood pump occlusion detected | If blood occlusion is detected > OCCLUSION_THRESHOLD_OFFSET + bloodPumpOcclusionAfterCartridgeInstall | 142 | ALARM_ID_HD_OCCLUSION_BLOOD_PUMP |  | ALARM_PRIORITY_HIGH | 206 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | HD blood pump occlusion detected | HD | Blood Stop: Blood Pump Occlusion
x | PRS 863 | SA 338 |  | SRSDG 912 |  | Call service to schedule a repair. | DG dialysate temperature sensors drift timeout. |  | If TDi and TRo have drifted for more 1 C for a certain period of time. |  | DG dialysate temperature sensors drift timeout. | If TDi and TRo have drifted for more 1 C for a certain period of time. | 143 | ALARM_ID_DG_DIALYSATE_TEMPERATURE_SENSORS_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG dialysate temperature sensors out of range | DG | DG Fault: TRo TDi Delta
x | PRS 853 |  |  | SRSDG 1066 SRSDG 1054 SRSDG 970 |  | Cleaning Mode Inlet water conductivity too high.\n- Cleaning mode process has terminated \n- Check inlet water before restarting cleaning mode. | DG cleaning mode inlet water conductivity out of high range |  | Inlet water conductivity > 2000 uS/cm. Inlet water conductivity > 100 uS/cm (PURE WATER MODE). In cleaning mode |  | DG cleaning mode inlet water conductivity out of high range | Inlet water conductivity > 2000 uS/cm. Inlet water conductivity > 100 uS/cm (PURE WATER MODE). In cleaning mode | 144 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_COND_TOO_HIGH |  | ALARM_PRIORITY_MEDIUM | 603 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water conductivity out of high range. | DG | Cleaning: Inlet Water Cond High
x | PRS 394 |  |  | SRSDG 846 |  | Acid concentrate conductivity during prime is out of range.\n- Check correct acid concentrate is connected and full.\n- Resume to reprime concentrate. | DG Concentrate conductivity after adding acid out of range alarm |  | If the conductivity value for the acid concentrate is out of range for more than 45 secs. |  | DG Concentrate conductivity after adding acid out of range alarm | If the conductivity value for the acid concentrate is out of range for more than 45 secs. | 145 | ALARM_ID_DG_ACID_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 410 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG concentrate conductivity after adding acid out of range alarm | DG | Dial. Gen: Acid Conc. Cond.
x | PRS 863 | SA 285 |  | SRSDG 312 |  | Call service to schedule a repair. | DG RTC vs. Timer accuracy error |  | If RTC and timer are not aligned on 1 second elapsed time. |  | DG RTC vs. Timer accuracy error | If RTC and timer are not aligned on 1 second elapsed time. | 146 | ALARM_ID_DG_RTC_OR_TIMER_ACCURACY_FAILURE |  | ALARM_PRIORITY_MEDIUM | 421 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | DG RTC or timer accuracy failure. | DG | DG POST: DG RTC Timer
x | PRS 543 |  |  | SRSDG 398 |  | Dialysate is not yet ready.\n- Treatment is paused and will be ready to resume when \ndialysate is ready for use. | DG dialysate generation conductivity fault alarm |  | If dialysate conductivity is not with 5% of the TCD |  | DG dialysate generation conductivity fault alarm | If dialysate conductivity is not with 5% of the TCD | 147 | ALARM_ID_DG_CREATING_DIALYSATE_PLEASE_WAIT |  | ALARM_PRIORITY_MEDIUM | 421 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | DG wait for the DG to produce dialysate | DG | Dial. Gen: Dialysate not ready
x | PRS 557 |  |  | SRSDG 1065 SRSDG 1053 SRSDG 969 |  | Cleaning Mode Inlet water temperature too high.\n- Cleaning mode process has terminated.\n- Check inlet water before restarting cleaning mode. | Cleaning Mode Inlet water temperature too high |  | Inlet temperature > 37 C. |  | Cleaning Mode Inlet water temperature too high | Inlet temperature > 37 C. | 148 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_TEMP_TOO_HIGH |  | ALARM_PRIORITY_MEDIUM | 601 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG Cleaning mode inlet water temperature too high | DG | Cleaning: Inlet Water Temp High
x | PRS 494 | SA 475 | SRSHD 646 |  |  | Call service to schedule a repair. | HD watchdog expired |  | If the watchdog signal is low while not in POST. |  | HD watchdog expired | If the watchdog signal is low while not in POST. | 149 | ALARM_ID_HD_WATCHDOG_EXPIRED |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD watchdog expired | HD | HD Fault: WD Error
x | PRS 404 |  |  | SRSDG 391 |  | Inlet water conductivity is high.\n- Treatment is paused.\n- Check the incoming water supply. | Inlet water conductivity too high |  | Inlet water conductivity > 2000 uS/cm. Inlet water conductivity > 100 uS/cm (PURE WATER MODE). |  | Inlet water conductivity too high | Inlet water conductivity > 2000 uS/cm. Inlet water conductivity > 100 uS/cm (PURE WATER MODE). | 150 | ALARM_ID_DG_INLET_WATER_CONDUCTIVITY_IN_HIGH_RANGE |  | ALARM_PRIORITY_MEDIUM | 434 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water conductivity in high range | DG | Dial. Gen: Inlet Water Cond High
x | PRS 403 |  |  | SRSDG 826 |  | Inlet water conductivity is low.\n- Treatment is paused.\n- Check the incoming water supply. | Inlet water conductivity too low |  | Inlet water conductivity < 200 uS/cm. Inlet water conductivity low disabled (PURE WATER MODE). |  | Inlet water conductivity too low | Inlet water conductivity < 200 uS/cm. Inlet water conductivity low disabled (PURE WATER MODE). | 151 | ALARM_ID_DG_INLET_WATER_CONDUCTIVITY_IN_LOW_RANGE |  | ALARM_PRIORITY_MEDIUM | 435 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water conductivity in low range | DG | Dial. Gen: Inlet Water Cond Low
x | PRS 401 |  |  | SRSDG 417 |  | Inlet water pressure is low.\n- Treatment is paused.\n- Check the incoming water supply. | Inlet water pressure too low |  | Inlet water pressure < 25 PSI. |  | Inlet water pressure too low | Inlet water pressure < 25 PSI. | 152 | ALARM_ID_DG_INLET_WATER_PRESSURE_IN_LOW_RANGE |  | ALARM_PRIORITY_MEDIUM | 431 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water pressure in low range | DG | Dial. Gen: Inlet Water Pres Low
x | PRS 529 |  | SRSHD 1122 |  |  | Pre-Treatment timeout.\n- Patient was not connected to device\n- Pre-Treatment has been terminated\n- Proceed to end treatment to remove cartridge from system. | HD prime completed high priority alarm |  | In Pre-Treatment Recirculation for 120 minutes (escalation of alarm 213) |  | HD prime completed high priority alarm | In Pre-Treatment Recirculation for 120 minutes (escalation of alarm 213) | 153 | ALARM_ID_HD_PRIME_COMPLETED_HIGH |  | ALARM_PRIORITY_HIGH | 292 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | HD prime completed high priority alarm | HD | Process: Pre-Tx Terminated
x | PRS 854 |  |  | SRSDG 1067 SRSDG 1055 SRSDG 971 |  | Cleaning Mode Inlet water conductivity too low.\n- Cleaning mode process has terminated \n- Check inlet water before restarting cleaning mode. | Cleaning mode inlet water conductivity too low |  | Inlet water conductivity < 200 uS/cm. Inlet water conductivity low disabled (PURE WATER MODE). In cleaning mode |  | Cleaning mode inlet water conductivity too low | Inlet water conductivity < 200 uS/cm. Inlet water conductivity low disabled (PURE WATER MODE). In cleaning mode | 154 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_COND_TOO_LOW |  | ALARM_PRIORITY_MEDIUM | 604 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water conductivity out of low range. | DG | Cleaning: Inlet Water Cond Low
x | PRS 400 |  | SRSHD 1574 |  |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | HD No reservoir data message receive at least once every 2 seconds |  | No new reservoir data for 2 seconds. |  | HD No reservoir data message receive at least once every 2 seconds | No new reservoir data for 2 seconds. | 155 | ALARM_ID_HD_NEW_RESERVOIRS_DATA_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD No reservoirs data message receive at least once every 2 seconds | HD | Loss of DG Comm (Reservoir)
x | PRS 400 |  | SRSHD 1575 |  |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | HD No DG op mode message receive at least once every 2 seconds |  | No DG op mode for 2 seconds. |  | HD No DG op mode message receive at least once every 2 seconds | No DG op mode for 2 seconds. | 156 | ALARM_ID_HD_DG_NEW_OPERATION_MODE_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD No DG operation mode message receive at least once every 2 seconds | HD | Loss of DG Comm (Op Mode)
cleaning |  |  |  | SRSDG 988 |  | Unable to sense proper chemical disinfectant.\n- Cleaning mode process has terminated.\n- Check that proper chemical disinfectant is connected to the \nbicarb port.\n- Check chemical disinfectant bottle is not empty. | Chemical disinfect priming the acid line timeout |  | If the disinfectant acid line is not primed within a certain period of time. |  | Chemical disinfect priming the acid line timeout | If the disinfectant acid line is not primed within a certain period of time. | 157 | ALARM_ID_DG_CHEM_DISINFECT_PRIME_ACID_LINE_TIME_OUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG chemical disinfect prime acid line timeout. | DG | Cleaning: Chem Disinf PAA Prime
x | PRS 405 |  |  | SRSDG 408 |  | Inlet water temperature is low.\n- Treatment is paused.\n- Check the incoming water supply. | Inlet water temperature low |  | Inlet temperature < 24 °C for 5 seconds. |  | Inlet water temperature low | Inlet temperature < 24 °C for 5 seconds. | 158 | ALARM_ID_DG_INLET_WATER_TEMPERATURE_IN_LOW_RANGE |  | ALARM_PRIORITY_MEDIUM | 433 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water temperature is in the low range | DG | Dial. Gen: Inlet Water Temp Low
cleaning |  |  |  | SRSDG 1044 |  | Remove chemical disinfectant connection and close cap. | Close concentrate line be able to flush it |  | Start of chemical disinfect flush. |  | Close concentrate line be able to flush it | Start of chemical disinfect flush. | 159 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_REMOVE_ACID |  | ALARM_PRIORITY_LOW | 751 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG chemical disinfect flush remove acid and close the concentrate cap. | DG | Cleaning: Chem Disinf Connect Cap
x | PRS 494 | SA 239 | SRSHD 972 |  |  | Call service to schedule a repair. | Blood pump rotor speed too high |  | If blood pump motor speed exceeded 100.0 RPM. |  | Blood pump rotor speed too high | If blood pump motor speed exceeded 100.0 RPM. | 160 | ALARM_ID_HD_BLOOD_PUMP_ROTOR_SPEED_TOO_HIGH |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD blood pump rotor speed too high | HD | HD Fault: BP Rotor Max Speed
x | PRS 863 | SA 497 |  | SRSDG 122 |  | Call service to schedule a repair. | Too many bad CRCs detected on received system messages |  | If >= 5 bad CRCs detected on received CAN messages within 10 minutes. |  | Too many bad CRCs detected on received system messages | If >= 5 bad CRCs detected on received CAN messages within 10 minutes. | 161 | ALARM_ID_DG_COMM_TOO_MANY_BAD_CRCS |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG comm too many bad CRCs | DG | DG Fault: HD-DG Comm
x | PRS 863 | SA 300 |  | SRSDG 867 |  | Call service to schedule a repair. | FPGA clock speed failure |  | If the FPGA clock speed is relatively close to the processor clock for a certain period of time. |  | FPGA clock speed failure | If the FPGA clock speed is relatively close to the processor clock for a certain period of time. | 162 | ALARM_ID_DG_FPGA_CLOCK_SPEED_CHECK_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG FPGA clock speed check failure. | DG | DG Fault: Proc-FPGA Clock
revisit - determine if this is blood stopped or dialysate may need to add to PRS |  |  | SRSHD 1260 |  |  | Ultrafiltration control detected and incorrect rate.\n- This can be caused by excessive movement of system\n(shock/vibration)\n- Check system is stable and level\n- Resume and If problem persists, call service to schedule a\nrepair. | HD sees primary load cell for reservoir 1 change too much too fast |  | If the delta volume of the active reservoir exceeded 50.0 g (MAX_ACTIVE_LOAD_CELL_CHANGE_G). |  | HD sees primary load cell for reservoir 1 change too much too fast | If the delta volume of the active reservoir exceeded 50.0 g (MAX_ACTIVE_LOAD_CELL_CHANGE_G). | 163 | ALARM_ID_HD_LOAD_CELL_ACCELERATION_RES_1_ALARM |  | ALARM_PRIORITY_MEDIUM | 442 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD sees primary load cell for reservoir 1 change too much too fast | HD | Dial. Gen: UF Rate Check
revisit - determine if this is blood stopped or dialysate may need to add to PRS |  |  | SRSHD 1601 |  |  | Ultrafiltration control detected and incorrect rate.\n- This can be caused by excessive movement of system \n(shock/vibration).\n- Check system is stable and level\n- Resume and If problem persists, call service to schedule a \nrepair. | HD sees primary load cell for reservoir 2 change too much too fast |  | If the delta volume of the active reservoir exceeded 50.0 g (MAX_ACTIVE_LOAD_CELL_CHANGE_G). |  | HD sees primary load cell for reservoir 2 change too much too fast | If the delta volume of the active reservoir exceeded 50.0 g (MAX_ACTIVE_LOAD_CELL_CHANGE_G). | 164 | ALARM_ID_HD_LOAD_CELL_ACCELERATION_RES_2_ALARM |  | ALARM_PRIORITY_MEDIUM | 442 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD sees primary load cell for reservoir 2 change too much too fast | HD | Dial. Gen: UF Rate Check
sean remove from code 15min limint (SRS 1556 has been updated) |  |  | SRSHD 1228 SRSHD 1399 SRSHD 1556 SRSHD 1709 |  |  | Blood pump has been off for too long.\n- Treament must be terminated. Blood clot risk.\n- NO RINSEBACK ALLOWED. | HD in treatment rinseback operation timeout |  | If the maximum time inrinseback operation has been exceeded. |  | HD in treatment rinseback operation timeout | If the maximum time inrinseback operation has been exceeded. | 165 | ALARM_ID_HD_TREATMENT_RINSEBACK_TIMEOUT_ALARM |  | ALARM_PRIORITY_HIGH | 11 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | HD in treatment rinseback operation timeout | HD | Blood Stop: Blood Clot Risk
[13 empty] | 166 | ALARM_ID_AVAILABLE_3 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | Available for use
needs to be high priority due to blood stopped | [5 empty] | Device door is open.\n - Close door and press RESUME button to continue. | HD cartridge door opened alarm |  | If the cartridge door is opened as reported by the FPGA. |  | HD cartridge door opened alarm | If the cartridge door is opened as reported by the FPGA. | 167 | ALARM_ID_HD_CARTRIDGE_DOOR_OPENED |  | ALARM_PRIORITY_HIGH | 209 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD cartridge door opened alarm | HD | Blood Stop: Door Open
[13 empty] | 168 | ALARM_ID__AVAILABLE_70 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | Available for use
x | PRS 863 | SA 493 |  | SRSDG 643 |  | Call service to schedule a repair. | Dialysate flow rate out of range |  | If the dialysate flow rate > 2.00 L/min. |  | Dialysate flow rate out of range | If the dialysate flow rate > 2.00 L/min. | 169 | ALARM_ID_DG_DIALYSATE_FLOW_RATE_OUT_OF_MAX_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG Dialysate flow rate out of maximum range | DG | DG Fault: Dialysate Flow Sens Rg
Currently in CBIT section - should be somewhere else? | [5 empty] | Heparin syringe empty.\n- Treatment will continue\n- If additional heparin is required to complete treatment, \nadminister manually. | HD syringe empty alarm |  | If the syringe pump position is near empty position. |  | HD syringe empty alarm | If the syringe pump position is near empty position. | 170 | ALARM_ID_HD_SYRINGE_PUMP_SYRINGE_EMPTY |  | ALARM_PRIORITY_LOW | 730 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | HD syringe empty alarm | HD | Heparin Syringe Empty
|  |  | SRSHD 1525 |  |  | Heparin pump occlusion detected.\n- Check heparin line is unclamped.\n- Check there are no kinks in the heparin line. | HD syringe pump occlusion alarm |  | If the pressure sensor voltage reading > SYRINGE_FORCE_OCCLUSION_THRESHOLD_V. |  | HD syringe pump occlusion alarm | If the pressure sensor voltage reading > SYRINGE_FORCE_OCCLUSION_THRESHOLD_V. | 171 | ALARM_ID_HD_SYRINGE_PUMP_OCCLUSION |  | ALARM_PRIORITY_LOW | 716 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD syringe pump occlusion alarm | HD | Process: Pre-Tx Hep. Pump Occ.
|  |  | SRSHD 1383 |  |  | Not enough heparin is available for treatment.\n- Check if syringe pump is touching syringe plunger.\n- If syringe pump is touching syringe plunger, end treatment.\n- If syringe pump is not touching syringe plunger, reload \nsyringe with enough heparin. | HD syringe pump not enough Heparin alarm |  | If the estimated syringe volume needed for treatment is not sufficient |  | HD syringe pump not enough Heparin alarm | If the estimated syringe volume needed for treatment is not sufficient | 172 | ALARM_ID_HD_SYRINGE_PUMP_NOT_ENOUGH_HEPARIN_ALARM |  | ALARM_PRIORITY_LOW | 715 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD syringe pump not enough Heparin alarm | HD | Process: Pre-Tx Not Enough Hep.
x | PRS 494 | SA 474 | SRSHD 659 |  |  | Call service to schedule a repair. | HD RTC configuration error |  | If the RTC configuration registers report an error. |  | HD RTC configuration error | If the RTC configuration registers report an error. | 173 | ALARM_ID_HD_RTC_CONFIG_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD RTC configuration error | HD | HD Fault: RTC Config
x | PRS 398 | SA 198 | SRSHD 630 |  |  | Call service to schedule a repair. | HD RTC vs. Timer accuracy error |  | If RTC and timer are not aligned on 1 second elapsed time. |  | HD RTC vs. Timer accuracy error | If RTC and timer are not aligned on 1 second elapsed time. | 174 | ALARM_ID_HD_RTC_OR_TIMER_ACCURACY_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD RTC or timer accuracy failure | HD | HD POST: HD RTC Timer
x | PRS 494 | SA 245 SA 254 SA 262 | SRSHD 1502 SRSHD 1503 SRSHD 1568 |  |  | Call service to schedule a repair. | HD pump commutation error |  | If the set pump direction does not agree with the hall sensor direction. |  | HD pump commutation error | If the set pump direction does not agree with the hall sensor direction. | 175 | ALARM_ID_HD_PUMP_DIRECTION_STATUS_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD pump direction status error | HD | HD Fault: Pump Commutation Err
Dialysate Temp |  |  | SRSHD 895 | [9 empty] | 176 | ALARM_ID_AVAILABLE_4 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD dialysate temperature above high safety range
x | PRS 863 | SA 417 |  | SRSDG 919 |  | Call service to schedule a repair. | DG software fault. Software found itself in an unexpected state |  | If the specific software faults listed in alarm management SW faults has been detected |  | DG software fault. Software found itself in an unexpected state | If the specific software faults listed in alarm management SW faults has been detected | 177 | ALARM_ID_DG_SOFTWARE_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG software fault. Software found itself in an unexpected state | DG | DG Fault: Software Invalid State
x | PRS 494 | SA 225 |  |  | SRSUI 1066 | Call service to schedule a repair. | HD communication timeout |  | If the HD stops communicating with the UI. |  | HD communication timeout | If the HD stops communicating with the UI. | 178 | ALARM_ID_HD_COMM_TIMEOUT | OK | ALARM_PRIORITY_HIGH | 0 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD communication timeout | UI | HD Fault: UI-Proc Comm
|  |  |  | SRSDG 605 |  | Call service to schedule a repair. | DG FPGA communication down for too long |  | If retries for commands exceeds limit or FPGA reports comm error. |  | DG FPGA communication down for too long | If retries for commands exceeds limit or FPGA reports comm error. | 179 | ALARM_ID_DG_FPGA_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG FPGA communication down for too long | DG | DG Fault: FPGA Com
x | PRS 863 | SA 494 |  | SRSDG 883 |  | Call service to schedule a repair. | RO flow rate out of range |  | If the RO flow rate > 2.00 L/min. |  | RO flow rate out of range | If the RO flow rate > 2.00 L/min. | 180 | ALARM_ID_DG_RO_FLOW_RATE_OUT_OF_MAX_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO flow out of maximum range | DG | DG Fault: RO Flow Sensor Range
|  |  |  | SRSDG 849 SRSDG 838 |  | Call service to schedule a repair. | DG load cells weight out of range for tare |  | If the load cell weight is out of range for tare. |  | DG load cells weight out of range for tare | If the load cell weight is out of range for tare. | 181 | ALARM_ID_DG_LOAD_CELLS_TARE_WEIGHT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells weight out of range for tare | DG | DG Fault: Tare Out of Range
x | PRS 863 | SA 369 |  | SRSDG 162 |  | Call service to schedule a repair. | DG load cells invalid calibration |  | If the CRC of load cells calibration data does not match the calculated CRC. |  | DG load cells invalid calibration | If the CRC of load cells calibration data does not match the calculated CRC. | 182 | ALARM_ID_DG_LOAD_CELLS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells invalid calibration | DG | DG POST: LC Cal
cleaning |  |  |  | SRSDG 851 |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | DG invalid load cell value |  | If the half full reservoir is filled prior to filling the reservoir that is supposed to overflow to the half full reservoir. |  | DG invalid load cell value | If the half full reservoir is filled prior to filling the reservoir that is supposed to overflow to the half full reservoir. | 183 | ALARM_ID_DG_INVALID_LOAD_CELL_VALUE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG invalid load cell value | DG | Cleaning: Invalid Resv Transfer
x | PRS 863 | SA 355 SA 356 |  | SRSDG 559 |  | Call service to schedule a repair. | Inlet reactor not healthy |  | If the inlet reactor are not healthy. |  | Inlet reactor not healthy | If the inlet reactor are not healthy. | 184 | ALARM_ID_DG_INLET_UV_REACTOR_NOT_HEALTHY |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG inlet UV reactor not healthy | DG | DG Fault: UVi Error
x | PRS 874 | SA 299 |  | SRSDG 628 |  | Fan failure.\n- Check that the fans inlets are not blocked or obstructed. | DG fan RPM out of range |  | If the fan RPM is out of Min/Max range. |  | DG fan RPM out of range | If the fan RPM is out of Min/Max range. | 185 | ALARM_ID_DG_FAN_RPM_OUT_OF_RANGE |  | ALARM_PRIORITY_LOW | 851 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | DG fan RPM out of range | DG | DG Alarm (Low): Fan Speed
x | PRS 863 | SA 351 SA 353 |  | SRSDG 891 |  | Call service to schedule a repair. | DG concentrate pump fault |  | If the concentrate pumps fault is reported by the FPGA. |  | DG concentrate pump fault | If the concentrate pumps fault is reported by the FPGA. | 186 | ALARM_ID_DG_CONCENTRATE_PUMP_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG concentrate pump fault | DG | DG Fault: CP1/CP2 Fault
x | PRS 863 | SA 352 |  | SRSDG 905 |  | Call service to schedule a repair. | DG concentrate pump CP1 speed control error |  | If the concentrate pump CP1 speed > 0.02 tolerance for a certain period of time. |  | DG concentrate pump CP1 speed control error | If the concentrate pump CP1 speed > 0.02 tolerance for a certain period of time. | 187 | ALARM_ID_DG_CP1_SPEED_CONTROL_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG concentrate pump CP1 speed control error | DG | DG Fault: CP1 Speed Fault
x | PRS 863 | SA 354 |  | SRSDG 906 |  | Call service to schedule a repair. | DG concentrate pump CP2 speed control error |  | If the concentrate pump CP1 speed > 0.02 tolerance for a certain period of time. |  | DG concentrate pump CP2 speed control error | If the concentrate pump CP1 speed > 0.02 tolerance for a certain period of time. | 188 | ALARM_ID_DG_CP2_SPEED_CONTROL_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG concentrate pump CP2 speed control error | DG | DG Fault: CP2 Speed Fault
x | PRS 863 | SA 339 SA 343 |  | SRSDG 448 |  | Call service to schedule a repair. | DG drain pump RPM out of range |  | If the drain pump RPM > 100.0 of the target RPM. |  | DG drain pump RPM out of range | If the drain pump RPM > 100.0 of the target RPM. | 189 | ALARM_ID_DG_DRAIN_PUMP_RPM_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain pump RPM out of range | DG | DG Fault: DRP Speed Fault
[6 empty] | Call service to schedule a repair. | DG drain pump off fault |  | If the pump is in off state and the RPM is greater than the minimum RPM of 300.0 RPM. |  | DG drain pump off fault | If the pump is in off state and the RPM is greater than the minimum RPM of 300.0 RPM. | 190 | ALARM_ID_DG_DRAIN_PUMP_OFF_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain pump off fault | DG | DG Fault: DRP Speed Not Off
x | PRS 863 | SA 498 |  | SRSDG 885 |  | Call service to schedule a repair. | DG flow rate out of upper range |  | If the RO flow rate > 10% from the target flow rate. |  | DG flow rate out of upper range | If the RO flow rate > 10% from the target flow rate. | 191 | ALARM_ID_DG_FLOW_RATE_OUT_OF_UPPER_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG flow rate out of upper range | DG | DG Fault: RO Flow High
x | PRS 863 | SA 498 |  | SRSDG 945 |  | Call service to schedule a repair. | DG flow rate out of lower range |  | If the RO flow rate < 10% from the target flow rate. |  | DG flow rate out of lower range | If the RO flow rate < 10% from the target flow rate. | 192 | ALARM_ID_DG_FLOW_RATE_OUT_OF_LOWER_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG flow rate out of lower range | DG | DG Fault: RO Flow Low
Dialysate Temp |  |  | SRSHD 897 | [9 empty] | 193 | ALARM_ID_AVAILABLE_5 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD dialysate temperature out of low safety range.
x | PRS 863 | SA 360 |  | SRSDG 881 |  | Call service to schedule a repair. | DG RO pump duty cycle out of range |  | If the DG RO pump duty cycle is out of range by 5% of the target duty cycle for a certain period of time. |  | DG RO pump duty cycle out of range | If the DG RO pump duty cycle is out of range by 5% of the target duty cycle for a certain period of time. | 194 | ALARM_ID_DG_RO_PUMP_DUTY_CYCLE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO pump duty cycle out of range | DG | DG Fault: RO Driver Feedback
x | PRS 863 | SA 487 |  | SRSDG 925 |  | Call service to schedule a repair. | DG RO pump pressure out of range |  | If the DG RO pump pressure > 135 psi for a certain period of time. |  | DG RO pump pressure out of range | If the DG RO pump pressure > 135 psi for a certain period of time. | 195 | ALARM_ID_DG_RO_PUMP_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO pump pressure out of range | DG | DG Fault: RO Pump High Pressure
x | PRS 863 | SA 348 |  | SRSDG 955 |  | Call service to schedule a repair. | DG CPi/CPo conductivity sensors FPGA fault |  | If the FPGA reports error for a certain period of time. |  | DG CPi/CPo conductivity sensors FPGA fault | If the FPGA reports error for a certain period of time. | 196 | ALARM_ID_DG_CPI_CPO_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CPi/CPo sensors FPGA fault | DG | DG Fault: CPi/CPo Comm
x | PRS 863 | SA 347 |  | SRSDG 956 |  | Call service to schedule a repair. | DG CD1/CD2 conductivity sensors FPGA fault |  | If the FPGA reports error for a certain period of time. |  | DG CD1/CD2 conductivity sensors FPGA fault | If the FPGA reports error for a certain period of time. | 197 | ALARM_ID_DG_CD1_CD2_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CD1/CD2 sensors FPGA fault | DG | DG Fault: CD1/CD2 Comm
x | PRS 863 | SA 367 |  | SRSDG 872 |  | Call service to schedule a repair. | DG primary heater while RO flow is low |  | If the RO flow is < 0.2 L/min for a certain period of time. |  | DG primary heater while RO flow is low | If the RO flow is < 0.2 L/min for a certain period of time. | 198 | ALARM_ID_DG_RO_FLOW_TOO_LOW_WHILE_PRIMARY_HEATER_IS_ON |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO flow too low while primary heater is on | DG | DG Fault: HP On With No Flow
x | PRS 863 | SA 366 |  | SRSDG 871 |  | Call service to schedule a repair. | DG trimmer heater while dialysate flow is low |  | If the FMD flow is ≤ 0 mL/min (raw) for a certain period of time. |  | DG trimmer heater while dialysate flow is low | If the FMD flow is ≤ 0 mL/min (raw) for a certain period of time. | 199 | ALARM_ID_DG_DIALYSATE_FLOW_TOO_LOW_WHILE_TRIMMER_HEATER_IS_ON |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG Dialysate flow too low while trimmer heater is on | DG | DG Fault: HR On With No Flow
x | PRS 863 | SA 293 SA 358 SA 359 |  | SRSDG 178 |  | Call service to schedule a repair. | DG thermistor or temperature sensor out of range |  | If thermistor > 80.0 C or < 0.0 C. |  | DG thermistor or temperature sensor out of range | If thermistor > 80.0 C or < 0.0 C. | 200 | ALARM_ID_DG_THERMISTORS_TEMPERATURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG thermistors/sensors temperature out of range | DG | DG Fault: Thermistor Range
|  |  | SRSHD 746 SRSHD 747 |  |  | Self-test has detected a problem in the dialysate flow \nmeasurement or there is a leak in the dialysate in/out lines.\n- The cartridge should be primed\n- Check cartridge installation on to machine\n- If problem persists, terminate pre-treatment and restart \nwith a new cartridge. | HD pre-treatment wet flow test failure |  | If the pre-treatment wet flow test has failed. |  | HD pre-treatment wet flow test failure | If the pre-treatment wet flow test has failed. | 201 | ALARM_ID_HD_PRE_TREATMENT_WET_FLOW_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode wet self-test lc vs FMD failure | HD | Process: Pre-Tx Wet Self-Test
[6 empty] | Self-test has detected blood circuit pressure cannot be \nrelieved.\n- Check PBo connection to HD for kinks\n -Check PBo transducer protector for occlusion\n- If problem persists, install a new cartridge and \nPBo transducer protector. | HD pre-treatment dry pressure normal test failure |  | If the pre-treatment dry pressure normal test failed. |  | HD pre-treatment dry pressure normal test failure | If the pre-treatment dry pressure normal test failed. | 202 | ALARM_ID_HD_PRE_TREATMENT_DRY_PRESSURE_NORMAL_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode dry pressure normal self-test failure | HD | Process: Pre-Tx Dry Self-Test
[13 empty] | 203 | ALARM_ID_AVAILABLE_6 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Available for use
x | PRS 863 | SA 373 |  | SRSDG 429 |  | Call service to schedule a repair. | RO rejection ratio out of range |  | If RO rejection ratio > 0.1 |  | RO rejection ratio out of range | If RO rejection ratio > 0.1 | 204 | ALARM_ID_DG_RO_REJECTION_RATIO_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO rejection ratio out of range | DG | DG Fault: RO Rejection
|  |  |  | SRSDG 913 |  | Call service to schedule a repair. | Conductivity sensor fault |  | If conductivity sensor reports 0x11 for more than 2 consecutive reads. |  | Conductivity sensor fault | If conductivity sensor reports 0x11 for more than 2 consecutive reads. | 205 | ALARM_ID_DG_CONDUCTIVITY_SENSOR_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG conductivity sensor fault | DG | DG Fault: Cond Sensors Error
|  |  |  | SRSDG 437 |  | Call service to schedule a repair. | DG dialysate fill runs out of time |  | If DG dialysate fill > 5 minutes. |  | DG dialysate fill runs out of time | If DG dialysate fill > 5 minutes. | 206 | ALARM_ID_DG_DIALYSATE_FILL_OUT_OF_TIME |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG dialysate fill runs out of time | DG | DG Fault: Reservoir Fill Timeout
|  |  |  | SRSDG 240 |  | Call service to schedule a repair. | DG flow meter check failure alarm |  | If integrated Volume To Load Cell Reading Percent > 0.1. |  | DG flow meter check failure alarm | If integrated Volume To Load Cell Reading Percent > 0.1. | 207 | ALARM_ID_DG_FLOW_METER_CHECK_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG flow meter check failure alarm | DG | DG Fault: FMP Flow Rate Error
blaine to revisit redundant with venous pressure alarm limit? | [5 empty] | Venous pressure is high.\n- Check for access or patient discomfort.\n- Resolve access issue if applicable.\n- Check that there are no kinks or clots in the tubing.\n- Unkink the lines if applicable.\n- End treatment and do not return blood if clots are present. | HD venous line occlusion fault |  | If the venous pressure is > set limit + 50 mmHg if HD is in treatment or pre-treatment mode. |  | HD venous line occlusion fault | If the venous pressure is > set limit + 50 mmHg if HD is in treatment or pre-treatment mode. | 208 | ALARM_ID_HD_VENOUS_LINE_OCCLUSION |  | ALARM_PRIORITY_HIGH | 201 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | HD venous pressure occlusion alarm | HD | Blood Stop: Venous Pres. High
|  |  |  | SRSDG 980 |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | DG drain circulation line timeout |  | If the integrated volume < 0.1 for more than 60 secs. |  | DG drain circulation line timeout | If the integrated volume < 0.1 for more than 60 secs. | 209 | ALARM_ID_DG_DRAIN_CIRCULATION_LINE_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG drain circulation line timeout | DG | Cleaning: Flush Circ Error
x | PRS 872 | SA 229 | SRSHD 1397 |  |  | Call service to schedule a repair. | HD battery pack detected an error |  | If the battery pack status error is detected from the slave device I2C address 0x0B. |  | HD battery pack detected an error | If the battery pack status error is detected from the slave device I2C address 0x0B. | 210 | ALARM_ID_HD_BATTERY_PACK_ERROR_DETECTED |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD battery pack detected an error | HD | HD Alarm (Low): Battery Fault
need to check with sean how this will escalate with other alarm active that will affect rank | PRS 389 |  | SRSHD 946 |  |  | Blood pump is stopped.\n- Resume treatment before countdown timer expires. | Blood sitting too long warning (> 4 min) |  | During treatment stop sub-mode, if the blood is in the line for more than 4 minutes. |  | Blood sitting too long warning (> 4 min) | During treatment stop sub-mode, if the blood is in the line for more than 4 minutes. | 211 | ALARM_ID_HD_BLOOD_SITTING_WARNING |  | ALARM_PRIORITY_HIGH | 290 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD blood sitting too long warning (>4 min) | HD | Blood Stopped: Warning
need to check with sean how this will escalate with other alarm active that will affect rank | PRS 389 |  | SRSHD 952 |  |  | Blood pump is stopped.\n- Resume treatment before countdown timer expires. | Blood sitting too long warning (> 4 min) |  | During treatment end paused state, if the blood is in the line for more than 4 minutes. |  | Blood sitting too long warning (> 4 min) | During treatment end paused state, if the blood is in the line for more than 4 minutes. | 212 | ALARM_ID_HD_END_OF_TREATMENT_ALARM |  | ALARM_PRIORITY_HIGH | 290 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD end of treatment alarm (user not acting to end treatment) | HD | Blood Stopped: Warning
sean to advise on properties | PRS 530 |  | SRSHD 1120 |  |  | Prime of the cartridge is complete.\n- Proceed to connect to patient. | HD prime completed medium priority alarm |  | In Pre-Treatment Recirculation for 90 minutes (escalation of alarm 226) |  | HD prime completed medium priority alarm | In Pre-Treatment Recirculation for 90 minutes (escalation of alarm 226) | 213 | ALARM_ID_HD_PRIME_COMPLETED_MEDIUM |  | ALARM_PRIORITY_MEDIUM | 590 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | HD prime completed medium priority alarm | HD | Process: Pre-Tx Prime Complete
|  |  | SRSHD 909 | [9 empty] | 214 | ALARM_ID_AVAILABLE_30 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | Available for use
|  |  | SRSHD 826 | [9 empty] | 215 | ALARM_ID_AVAILABLE_31 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | Available for use
x | PRS 393 |  | SRSHD 1405 |  |  | Dialysis Treatment has completed\n- Proceed to end treatment and rinseback. | HD end treatment sub-mode timeout alarm |  | If the treatment ended successfully. |  | HD end treatment sub-mode timeout alarm | If the treatment ended successfully. | 216 | ALARM_ID_HD_END_TREATMENT_TIMEOUT_ALARM |  | ALARM_PRIORITY_HIGH | 291 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD end treatment sub-mode timeout alarm | HD | Process: End of Dialysis Tx
|  |  | SRSHD 1454 | [9 empty] | 217 | ALARM_ID_AVAILABLE_32 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | Available for use
|  |  | SRSHD 1294 |  |  | Syringe not removed from machine.\n- Remove heparin syringe. | HD syringe detected alarm |  | If the syringe detection switch > 2.0 volts. |  | HD syringe detected alarm | If the syringe detection switch > 2.0 volts. | 218 | ALARM_ID_HD_SYRINGE_DETECTED |  | ALARM_PRIORITY_LOW | 745 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD syringe detected alarm | HD | Process: Remove Syringe
Currently in CBIT section - should be somewhere else? | [5 empty] | Heparin is needed for treatment but syringe has been removed \nfrom syringe pump.\n- Install Heparin syringe. | HD syringe pump syringe removed alarm |  | If the syringe detection switch < 2.0 volts. |  | HD syringe pump syringe removed alarm | If the syringe detection switch < 2.0 volts. | 219 | ALARM_ID_HD_SYRINGE_PUMP_SYRINGE_REMOVED |  | ALARM_PRIORITY_HIGH | 207 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | HD syringe pump syringe removed alarm | HD | Blood Stop: Syringe Not Detected
x | PRS 863 | SA 334 SA 335 |  | SRSDG 947 |  | Call service to schedule a repair. | DG THd sensor FPGA fault |  | If the FPGA does not report fresh data for a certain period of time |  | DG THd sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 220 | ALARM_ID_DG_THD_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG THd sensors FPGA fault | DG | DG Fault: THD RTD Comm
|  |  | SRSHD 1306 SRSHD 1307 SRSHD 1308 |  |  | Saline bag appears empty.\n- Check saline bag and line and replace if saline bag is \nempty. | HD Saline bag is empty |  | If the measured arterial pressure < -300.0 mmHg for more than 250 msec. |  | HD Saline bag is empty | If the measured arterial pressure < -300.0 mmHg for more than 250 msec. | 221 | ALARM_ID_HD_EMPTY_SALINE_BAG |  | ALARM_PRIORITY_HIGH | 204 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | HD Saline bag is empty | HD | Blood Stop: Saline Bag Empty
x | PRS 494 | SA 230 | SRSHD 1585 |  |  | Call service to schedule a repair. | HD occlusion sensor FPGA fault |  | If the FPGA does not report fresh data or indicates an error for a certain period of time. |  | HD occlusion sensor FPGA fault | If the FPGA does not report fresh data or indicates an error for a certain period of time. | 222 | ALARM_ID_HD_OCCLUSION_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD occlusion sensor FPGA fault | HD | HD Fault: BP Occ Com
x | PRS 494 | SA 236 | SRSHD 1583 |  |  | Call service to schedule a repair. | HD arterial sensor FPGA fault |  | If the FPGA does not report fresh data for a certain period of time |  | HD arterial sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 223 | ALARM_ID_HD_ARTERIAL_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD arterial sensor FPGA fault | HD | HD Fault: Art. Pres. Com
x | PRS 388 |  | SRSHD 944 |  |  | Treatment paused by user. | Treatment stopped by user action - pressed stop button |  | If the user has pressed the stop button. |  | Treatment stopped by user action - pressed stop button | If the user has pressed the stop button. | 224 | ALARM_ID_HD_TREATMENT_STOPPED_BY_USER |  | ALARM_PRIORITY_MEDIUM | 300 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD treatment stopped by user action - pressed stop button | HD | Blood Stopped: Treatment Pause
revisit priority - should be low | PRS 391 |  | SRSHD 810 |  |  | Dialysis Treatment has completed\n- Proceed to end treatment and rinseback. | HD end of treatment warning |  | If the treatment duration has exceeded the prescribed treatment time entered by the user. |  | HD end of treatment warning | If the treatment duration has exceeded the prescribed treatment time entered by the user. | 225 | ALARM_ID_HD_END_OF_TREATMENT_WARNING |  | ALARM_PRIORITY_LOW | 739 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD end of treatment warning | HD | Process: End of Treatment
sean to advise on properties | PRS 531 |  | SRSHD 1118 |  |  | Prime of the cartridge is complete.\n- Proceed to connect to patient. | HD prime completed low priority alarm |  | Start of Pre-Treatment Recirculation - notifiy user to proceed to patient connection |  | HD prime completed low priority alarm | Start of Pre-Treatment Recirculation - notifiy user to proceed to patient connection | 226 | ALARM_ID_HD_PRIME_COMPLETED_LOW_PRIORITY |  | ALARM_PRIORITY_LOW | 729 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | HD prime completed low priority alarm | HD | Process: Pre-Tx Priming Complete
[13 empty] | 227 | ALARM_ID_AVAILABLE_15 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | Available for use
|  |  | SRSHD 1571 |  |  | Blood line prime with saline timeout.\n- Check blood and dialysate lines for clamps, kinks or leaks. | HD time out on prime sub-mode purge air state. |  | During the pre-treatment prime state, if the prime saline purge air state exceeded 60 secs. |  | HD time out on prime sub-mode purge air state. | During the pre-treatment prime state, if the prime saline purge air state exceeded 60 secs. | 228 | ALARM_ID_HD_PRIME_SALINE_PURGE_AIR_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD time out on prime saline purge air state. | HD | Process: Pre-Tx Prime Timeout
|  |  | SRSHD 1189 |  |  | Dialyzer prime timeout.\n- Check blood and dialysate lines for clamps, kinks or leaks. | HD prime dialysate dialyzer time out alarm |  | If the priming of the dialyzer dialysate flow path has exceeded the time limit. |  | HD prime dialysate dialyzer time out alarm | If the priming of the dialyzer dialysate flow path has exceeded the time limit. | 229 | ALARM_ID_HD_PRIME_DIALYSATE_DIALYZER_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD prime dialysate dialyzer time out alarm | HD | Process: Pre-Tx Prime Timeout
|  |  | SRSHD 1190 |  |  | Dialyzer prime timeout.\n- Check blood and dialysate lines for clamps, kinks or leaks. | HD prime dialysate bypass time out alarm |  | If the priming of the dialyzer bypass flow path has exceeded the time limit. |  | HD prime dialysate bypass time out alarm | If the priming of the dialyzer bypass flow path has exceeded the time limit. | 230 | ALARM_ID_HD_PRIME_DIALYSATE_BYPASS_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD prime dialysate bypass time out alarm | HD | Process: Pre-Tx Prime Timeout
|  |  | SRSHD 771 SRSHD 1599 SRSHD 1579 SRSHD 1580 SRSHD 772 |  |  | Self-test has detected a leak in the circuit\n- Check blood line connections\n- Check dialysate line connections\n- Check PBo and PBa transducer protectors for occlusion.\n- Check clamps.\n- If problem persists, install a new cartridge. | HD pre-treatment mode dry self-tests failure |  | If the following dry self tests have failed: 1. handleDrySelfTestPressureSensorsState 2. handleDrySelfTestPressureSensorsNormalState |  | HD pre-treatment mode dry self-tests failure | If the following dry self tests have failed: 1. handleDrySelfTestPressureSensorsState 2. handleDrySelfTestPressureSensorsNormalState | 231 | ALARM_ID_HD_PRE_TREATMENT_DRY_PRESSURE_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode dry pressure self-test failure | HD | Process: Pre-Tx Dry Self-Test
|  |  | SRSHD 1534 SRSHD 1535 |  |  | Self-test has detected a problem in the load cell \nmeasurement or there is a leak in the dialysate in/out lines\n- The cartridge should be primed\n- Check cartridge installation onto machine\n- If problem persists, terminate pre-treatment and restart \nwith a new cartridge. | HD pre-treatment mode wet self-tests failure |  | If the following wet self tests have failed: 1. handleWetSelfTestPrimeCheckState 2. handleWetSelfTestFirstDisplacementVerifyState 3. handleWetSelfTestSecondDisplacementVerifyState |  | HD pre-treatment mode wet self-tests failure | If the following wet self tests have failed: 1. handleWetSelfTestPrimeCheckState 2. handleWetSelfTestFirstDisplacementVerifyState 3. handleWetSelfTestSecondDisplacementVerifyState | 232 | ALARM_ID_HD_PRE_TREATMENT_WET_LC_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode wet self-test LC vs. LC failure | HD | Process: Pre-Tx Wet Self-Test
x | PRS 863 | SA 461 |  | SRSDG 964 SRSDG 949 SRSDG 1089 |  | Call service to schedule a repair. | DG barometric pressure sensor bad coefficients CRC |  | If the following happens: 1. The stored CRC in the sensor does not match the calculated CRC. 2. The CRC and coefficients have not been received from the sensor after 20 seconds from the power up. |  | DG barometric pressure sensor bad coefficients CRC | If the following happens: 1. The stored CRC in the sensor does not match the calculated CRC. 2. The CRC and coefficients have not been received from the sensor after 20 seconds from the power up. | 233 | ALARM_ID_DG_BAROMETRIC_SENSOR_COEFFS_BAD_CRC |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG barometric sensor coefficients bad CRC | DG | DG Fault: Baro Cal CRC
x | PRS 841 |  |  | SRSDG 1037 |  | Inlet water pressure is high.\n- Treatment is paused\n- Check the incoming water supply. | DG inlet water pressure is too high |  | If the inlet water pressure is > 80 psig. |  | DG inlet water pressure is too high | If the inlet water pressure is > 80 psig. | 234 | ALARM_ID_DG_INLET_WATER_PRESSURE_IN_HIGH_RANGE |  | ALARM_PRIORITY_MEDIUM | 430 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water pressure in high range. | DG | Dial. Gen: Inlet Water Pres High
[6 empty] | Rinseback has been completed. | HD in treatment stopped sub-mode after rinseback completed (no escalation) |  | If there is no active alarm and rinseback has completed. |  | HD in treatment stopped sub-mode after rinseback completed (no escalation) | If there is no active alarm and rinseback has completed. | 235 | ALARM_ID_HD_TREATMENT_STOPPED_AFTER_RINSEBACK |  | ALARM_PRIORITY_LOW | 738 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | HD in treatment stopped sub-mode after rinseback completed (no escalation) | HD | Process: Rinseback Complete
|  |  | SRSHD 1205 |  |  | Self-test has detected fluid in the cartridge.\n- Cartridges are single-use only\n- Install new cartridge to resume. | HD needs new cartridge to be installed |  | If the air trap sensor reported AIR and bubble sensor detected bubbles. |  | HD needs new cartridge to be installed | If the air trap sensor reported AIR and bubble sensor detected bubbles. | 236 | ALARM_ID_HD_INSTALL_NEW_CARTRIDGE |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD needs new cartridge to be installed | HD | Process: Pre-Tx Dry Self-Test
|  |  | SRSHD 1188 |  |  | Dialyzer prime timeout.\n- Check blood and dialysate lines for clamps, kinks or leaks. | HD prime saline dialyzer time out alarm |  | If timeout, just transition to HD_PRIME_RESERVOIR_TWO_FILL_COMPLETE_STATE |  | HD prime saline dialyzer time out alarm | If timeout, just transition to HD_PRIME_RESERVOIR_TWO_FILL_COMPLETE_STATE | 237 | ALARM_ID_HD_PRIME_SALINE_DIALYZER_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD prime saline dialyzer time out alarm | HD | Process: Pre-Tx Prime Timeout
|  |  | SRSHD 1359 |  |  | Self-test has detected cartridge is not installed.\n- Check that a cartridge is properly installed. | HD no cartridge loaded or installed improperly alarm |  | If the measured blood pump occlucsion pressure reading is < 5000 or > 24500. |  | HD no cartridge loaded or installed improperly alarm | If the measured blood pump occlucsion pressure reading is < 5000 or > 24500. | 238 | ALARM_ID_HD_NO_CARTRIDGE_LOADED |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | HD no cartridge loaded or installed improperly alarm | HD | Process: Pre-Tx Dry Self-Test
|  |  | SRSHD 1542 |  |  | Cartridge not unloaded, or unloaded improperly.\n- Check that the cartridge has been removed. | HD fail to remove cartridge alarm |  | In the post treatment state, if the cartridge has not been removed. |  | HD fail to remove cartridge alarm | In the post treatment state, if the cartridge has not been removed. | 239 | ALARM_ID_HD_CARTRIDGE_REMOVAL_FAILURE |  | ALARM_PRIORITY_LOW | 746 | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | HD fail to remove cartridge alarm | HD | Process: Post-Tx Remove Cartridge
x | PRS 542 |  |  | SRSDG 482 |  | Bicarbonate conductivity during prime is out of range.\n- Check correct concentrate is connected and full.\n- Resume to reprime concentrate. | Bicarb conductivity out of range during bicarb pump check alarm |  | If the conductivity value for bicarbonate concentrate is out of range for more than 45 secs. |  | Bicarb conductivity out of range during bicarb pump check alarm | If the conductivity value for bicarbonate concentrate is out of range for more than 45 secs. | 240 | ALARM_ID_DG_BICARB_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 411 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | DG bicarb conductivity out of range during bicarb pump check alarm | DG | Dial. Gen: Bicarb Conc. Cond.
x | PRS 873 |  |  | SRSDG 848 SRSDG 90 SRSDG 561 SRSDG 577 SRSDG 590 SRSDG 591 SRSDG 336 SRSDG 879 SRSDG 340 SRSDG 496 SRSDG 501 SRSDG 505 SRSDG 532 SRSDG 535 SRSDG 556 SRSDG 557 SRSDG 1048 SRSDG 1049 SRSDG 651 SRSDG 655 SRSDG 744 SRSDG 745 SRSDG 1004 SRSDG 1008 SRSDG 1019 SRSDG 1023 SRSDG 1033 SRSDG 1034 |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | DG reservoir drain time out |  | If the DG reservoir drain is not complete in 120 secs. |  | DG reservoir drain time out | If the DG reservoir drain is not complete in 120 secs. | 241 | ALARM_ID_DG_RESERVOIR_DRAIN_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG reservoir drain time out | DG | Cleaning: Resv Drain Timeout
cleaning | PRS 873 |  |  | SRSDG 850 SRSDG 573 SRSDG 574 SRSDG 579 SRSDG 580 SRSDG 492 SRSDG 493 SRSDG 498 SRSDG 499 SRSDG 508 SRSDG 509 SRSDG 519 SRSDG 690 SRSDG 691 SRSDG 703 SRSDG 711 SRSDG 1020 SRSDG 1024 |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | DG reservoir fill time out |  | If the DG reservoir fill is not complete in 120 secs. |  | DG reservoir fill time out | If the DG reservoir fill is not complete in 120 secs. | 242 | ALARM_ID_DG_RESERVOIR_FILL_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG reservoir fill time out | DG | Cleaning: Resv Fill Timeout
cleaning |  |  |  | SRSDG 853 |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | DG reservoir leak time out |  | If the reservoir level is drifted by 100 mL for more than 30 seconds. |  | DG reservoir leak time out | If the reservoir level is drifted by 100 mL for more than 30 seconds. | 243 | ALARM_ID_DG_RESERVOIR_LEAK_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG reservoir leak time out | DG | Cleaning: Fluid Loss Detected
cleaning |  |  |  | SRSDG 852 |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | DG temperature sensors difference out of range |  | If any of the temperature sensors are out of the average of the sensors by more than 3 C. |  | DG temperature sensors difference out of range | If any of the temperature sensors are out of the average of the sensors by more than 3 C. | 244 | ALARM_ID_DG_CLEANING_MODE_TEMP_SENSORS_DIFF_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode temperature sensors difference out of range | DG | Cleaning: Temp Sensor Check
cleaning |  |  |  | SRSDG 854 |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | DG heat disinfect target temperature time out (could not reach to temperature) |  | If heating up to minimum temperature for heat disinfect failed after 4 hours. |  | DG heat disinfect target temperature time out (could not reach to temperature) | If heating up to minimum temperature for heat disinfect failed after 4 hours. | 245 | ALARM_ID_DG_HEAT_DISINFECT_TARGET_TEMP_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG heat disinfect target temperature time out (could not reach to temperature) | DG | Cleaning: Heat Disinfect Temp TO
cleaning | [5 empty] | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | DG chemical disinfect inlet conducitivty out of range |  | If CPo and CD2 conductivity values are above 100 uS/cm. |  | DG chemical disinfect inlet conducitivty out of range | If CPo and CD2 conductivity values are above 100 uS/cm. | 246 | ALARM_ID_DG_CLEANING_MODE_COND_SENSORS_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode conductivity sensors out of range | DG | Cleaning: RO Water Cond Error
cleaning | [12 empty] | 247 | ALARM_ID_AVAILABLE_7 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Alarm Id available
cleaning | [5 empty] | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | DG chemical disinfect out of target temperature |  | If TPo > 50 C after 1 second. |  | DG chemical disinfect out of target temperature | If TPo > 50 C after 1 second. | 248 | ALARM_ID_DG_CHEM_DISINFECT_TARGET_TEMP_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG chemical disinfect could not reach to target temperature out of range | DG | Cleaning: Chem Disinfect Proc
cleaning |  |  |  | SRSDG 992 |  | Unable to sense proper chemical disinfectant.\n- Cleaning mode process has terminated.\n- Check that proper chemical disinfectant is connected to the \nbicarb port.\n- Check chemical disinfectant bottle is not empty. | DG chemical disinfect out of target conductivity |  | If CD2 < 150 uS/cm or CD2 > 650 uS/cm after 10 seconds. |  | DG chemical disinfect out of target conductivity | If CD2 < 150 uS/cm or CD2 > 650 uS/cm after 10 seconds. | 249 | ALARM_ID_DG_CHEM_DISINFECT_TARGET_COND_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG chemical disinfect target conductivity out of range | DG | Cleaning: Chem Disinf PAA Cond
cleaning |  |  |  | SRSDG 1046 |  | Connect chemical disinfectant bottle to bicarb port. | Insert the acid to start chemical disinfect |  | Start of chemical disinfect |  | Insert the acid to start chemical disinfect | Start of chemical disinfect | 250 | ALARM_ID_DG_CHEM_DISINFECT_INSERT_ACID |  | ALARM_PRIORITY_LOW | 750 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG chemical disinfect insert acid and remove the concentrate cap | DG | Cleaning: Chem Disinf Connect PAA
x | PRS 398 | SA 196 | SRSHD 1058 |  |  | Call service to schedule a repair. | HD invalid system record |  | If the calculated CRC of the HD system record image does not match the expected CRC. |  | HD invalid system record | If the calculated CRC of the HD system record image does not match the expected CRC. | 251 | ALARM_ID_HD_INVALID_SYSTEM_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD invalid system record | HD | HD POST: Sys Rec NVRAM CRC
x | PRS 398 | SA 196 | SRSHD 1062 |  |  | Call service to schedule a repair. | HD invalid service record |  | If the calculated CRC of the HD service record image does not match the expected CRC. |  | HD invalid service record | If the calculated CRC of the HD service record image does not match the expected CRC. | 252 | ALARM_ID_HD_INVALID_SERVICE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD invalid service record | HD | HD POST: Serv Rec NVRAM CRC
x | PRS 863 | SA 283 |  | SRSDG 609 |  | Call service to schedule a repair. | DG invalid system record |  | If the calculated CRC of the DG system record image does not match the expected CRC. |  | DG invalid system record | If the calculated CRC of the DG system record image does not match the expected CRC. | 253 | ALARM_ID_DG_INVALID_SYSTEM_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid system record | DG | DG POST: Sys Rec NVRAM CRC
x | PRS 863 | SA 283 |  | SRSDG 610 |  | Call service to schedule a repair. | DG invalid service record |  | If the calculated CRC of the DG service record image does not match the expected CRC. |  | DG invalid service record | If the calculated CRC of the DG service record image does not match the expected CRC. | 254 | ALARM_ID_DG_INVALID_SERVICE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid service record | DG | DG POST: Serv Rec NVRAM CRC
x | PRS 398 | SA 206 | SRSHD 1358 |  |  | Call service to schedule a repair. | HD and UI software builds are not compatible |  | If the compatibility revision field of the UI is not equal to SW_COMPATIBILITY_REV. |  | HD and UI software builds are not compatible | If the compatibility revision field of the UI is not equal to SW_COMPATIBILITY_REV. | 255 | ALARM_ID_HD_UI_COMPATIBILITY_ERROR |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD and UI software builds are not compatible | HD | HD POST: UI Compatibility
x | PRS 398 | SA 205 |  |  | SRSUI 497 | Chemical disinfect flush in progress\n- Please wait for the process to complete\n- Treatment will not be able to be started until process \ncompletes.\n- Process may be terminated by pressing the HD stop button. | DG flush started |  | If DG flush successfully started from the UI. |  | DG flush started | If DG flush successfully started from the UI. | 256 | ALARM_ID_HD_DISINFECT_CHEM_FLUSH | OK | ALARM_PRIORITY_LOW | 843 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD processor is in Disinfect Chemical Flush mode | HD | Cleaning: Chem Flush In Progress
[13 empty] | 257 | ALARM_ID_AVAILABLE_8 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Alarm ID available for use
[13 empty] | 258 | ALARM_ID_AVAILABLE_9 |  | ALARM_PRIORITY_LOW | 999 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | Alarm Id available
x | PRS 494 | SA 220 SA 215 SA 216 SA 235 SA 238 | SRSHD 1390 SRSHD 1388 SRSHD 1387 SRSHD 1497 SRSHD 1499 SRSHD 1682 |  |  | Call service to schedule a repair. | HD temperatures out of range |  | If thermistor reports temperature > 80.0 C or < 0.0 C for a certain period of time. |  | HD temperatures out of range | If thermistor reports temperature > 80.0 C or < 0.0 C for a certain period of time. | 259 | ALARM_ID_HD_TEMPERATURES_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD temperatures out of range | HD | HD Fault: Temp Sensors Error
x | PRS 398 | SA 200 | SRSHD 977 |  | SRSUI 1067 | Call service to schedule a repair. | UI POST Application Integrity (Sha256Sum) failure |  | If the UI POST reports a shasum faillure. |  | UI POST Application Integrity (Sha256Sum) failure | If the UI POST reports a shasum faillure. | 260 | ALARM_ID_HD_UI_POST_FAILURE_SHASUM | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI POST Application Integrity (Sha256Sum) failure | UI HD | HD POST: UI Image
x | PRS 398 | SA 201 | SRSHD 1655 |  | SRSUI 1068 | Call service to schedule a repair. | UI POST CANBus failure |  | If the UI POST reports a can bus failure. |  | UI POST CANBus failure | If the UI POST reports a can bus failure. | 261 | ALARM_ID_HD_UI_POST_FAILURE_CANBUS | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI POST CANBus failure | UI HD | HD POST: UI CAN Driver
x | PRS 398 | SA 201 | SRSHD 1656 |  | SRSUI 1069 | Call service to schedule a repair. | UI POST Display failure |  | If the UI POST reports a display failure. |  | UI POST Display failure | If the UI POST reports a display failure. | 262 | ALARM_ID_HD_UI_POST_FAILURE_DISPLAY | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI POST Display failure | UI HD | HD POST: UI Disp Driver
x | PRS 398 | SA 201 | SRSHD 1657 |  | SRSUI 1070 | Call service to schedule a repair. | UI POST Touch failure |  | If the UI POST reports a touch failure. |  | UI POST Touch failure | If the UI POST reports a touch failure. | 263 | ALARM_ID_HD_UI_POST_FAILURE_TOUCH | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI POST Touch failure | UI HD | HD POST: UI Touch Driver
x | PRS 398 | SA 201 | SRSHD 1658 |  | SRSUI 1071 | SD-Card Failure \n- Power cycle the system.\n- Call service to schedule a repair if problem persists. | UI POST SD-Card failure |  | If the UI POST reports SD card failure. |  | UI POST SD-Card failure | If the UI POST reports SD card failure. | 264 | ALARM_ID_HD_UI_POST_FAILURE_SDCARD | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI POST SD-Card failure | UI HD | HD POST: UI SD-Card
x | PRS 398 | SA 201 SA 202 SA 203 | SRSHD 1659 |  | SRSUI 1072 | Call service to schedule a repair. | UI POST RTC failure |  | If the UI POST reports RTC failure. |  | UI POST RTC failure | If the UI POST reports RTC failure. | 265 | ALARM_ID_HD_UI_POST_FAILURE_RTC | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI POST RTC failure | UI HD | HD POST: UI RTC
x | PRS 398 | SA 201 | SRSHD 1660 |  | SRSUI 1073 | Call service to schedule a repair. | UI POST WiFi failure |  | If the UI POST reports WiFi failure. |  | UI POST WiFi failure | If the UI POST reports WiFi failure. | 266 | ALARM_ID_HD_UI_POST_FAILURE_WIFI | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI POST WiFi failure | UI HD | HD POST: UI WiFi
x | PRS 398 | SA 201 | SRSHD 1661 |  | SRSUI 1074 | Call service to schedule a repair. | UI POST Bluetooth failure |  | If the UI POST reports bluetooth failure. |  | UI POST Bluetooth failure | If the UI POST reports bluetooth failure. | 267 | ALARM_ID_HD_UI_POST_FAILURE_BLUETOOTH | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI POST Bluetooth failure | UI HD | HD POST: UI Bluetooth
x | PRS 398 | SA 201 | SRSHD 1662 |  | SRSUI 1075 | Call service to schedule a repair. | UI POST Ethernet failure |  | If the UI POST reports ethernet failure. |  | UI POST Ethernet failure | If the UI POST reports ethernet failure. | 268 | ALARM_ID_HD_UI_POST_FAILURE_ETHERNET | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI POST Ethernet failure | UI HD | HD POST: UI Ethernet
x | PRS 398 | SA 201 | SRSHD 1663 |  | SRSUI 1076 | Call service to schedule a repair. | UI POST Sound failure |  | If the UI POST reports sound failure. |  | UI POST Sound failure | If the UI POST reports sound failure. | 269 | ALARM_ID_HD_UI_POST_FAILURE_SOUND | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI POST Sound failure | UI HD | HD POST: UI Sound
x | PRS 398 | SA 459 | SRSHD 1481 |  |  | Call service to schedule a repair. | HD POST Safety Shutdown failure |  | If the 24 volts power source is greater than 5.0 volts during safety shutdown self test state. If the 24 volts power source is less than 22.6 volts during safety shutdown self test recover state. |  | HD POST Safety Shutdown failure | If the 24 volts power source is greater than 5.0 volts during safety shutdown self test state. If the 24 volts power source is less than 22.6 volts during safety shutdown self test recover state. | 270 | ALARM_ID_HD_SAFETY_SHUTDOWN_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD POST Safety Shutdown failure | HD | HD POST: Proc Safestate
x | PRS 863 | SA 292 |  | SRSDG 822 |  | Call service to schedule a repair. | DG POST Safety Shutdown failure |  | If the 24 volts power source is greater than 5.0 volts during safety shutdown self test state. If the 24 volts power source is less than 22.6 volts during safety shutdown self test recover state. |  | DG POST Safety Shutdown failure | If the 24 volts power source is greater than 5.0 volts during safety shutdown self test state. If the 24 volts power source is less than 22.6 volts during safety shutdown self test recover state. | 271 | ALARM_ID_DG_SAFETY_SHUTDOWN_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG POST Safety Shutdown failure | DG | DG POST: Proc Safestate
x | PRS 872 | SA 221 | SRSHD 1394 |  |  | Fan failure.\n- Check that the fans inlets are not blocked or obstructed. | HD Fan RPM out of range |  | If the fan RPM is out of Min/Max range. |  | HD Fan RPM out of range | If the fan RPM is out of Min/Max range. | 272 | ALARM_ID_HD_FAN_RPM_OUT_OF_RANGE |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD Fan RPM out of range | HD | HD Alarm (Low): Fan Speed
[13 empty] | 273 | ALARM_ID_AVAILABLE_10 |  | ALARM_PRIORITY_LOW | 999 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | Available for use
Blaine - Remove SRS (Done) Sean - Remove from code |  |  | Removed | [9 empty] | 274 | ALARM_ID_AVAILABLE_11 |  | ALARM_PRIORITY_LOW | 999 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | Available for use
x | PRS 494 | SA 237 | SRSHD 1706 |  |  | Call service to schedule a repair. | HD arterial pressure sensor is reading out of range |  | If the arterial pressure > 2000.0 mmHg or < -1500.0 mmHg. |  | HD arterial pressure sensor is reading out of range | If the arterial pressure > 2000.0 mmHg or < -1500.0 mmHg. | 275 | ALARM_ID_HD_ARTERIAL_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD arterial pressure sensor is reading out of range | HD | HD Fault: Arterial Press Range
x | PRS 494 | SA 234 | SRSHD 1707 |  |  | Call service to schedule a repair. | HD venous pressure sensor is reading out of range |  | If the venous pressure > 2000.0 mmHg or < -1500.0 mmHg. |  | HD venous pressure sensor is reading out of range | If the venous pressure > 2000.0 mmHg or < -1500.0 mmHg. | 276 | ALARM_ID_HD_VENOUS_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD venous pressure sensor is reading out of range | HD | HD Fault: Venous Press Range
Sean to remove from code check of 32767 check (align to SRS) | PRS 494 | SA 231 | SRSHD 1708 |  |  | Call service to schedule a repair. | HD BP occlusion sensor is reading out of range |  | If the blood pump occlusion count is < 2000. |  | HD BP occlusion sensor is reading out of range | If the blood pump occlusion count is < 2000. | 277 | ALARM_ID_HD_BP_OCCLUSION_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD BP occlusion sensor is reading out of range | HD | HD Fault: Occlusion Sens Range
Dara - update code to 1900g |  |  | SRSHD 793 |  |  | Call service to schedule a repair. | HD active reservoir weight out of range. |  | If the active reservoir weight is > 1900.0 g. |  | HD active reservoir weight out of range. | If the active reservoir weight is > 1900.0 g. | 278 | ALARM_ID_HD_ACTIVE_RESERVOIR_WEIGHT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD active reservoir weight out of range. | HD | DG Fault: Reservoir Overfill
|  |  |  | SRSDG 752 |  | Call service to schedule a repair. | DG dialysate drain time out |  | If the DG dialysate drain exceeded 120 seconds. |  | DG dialysate drain time out | If the DG dialysate drain exceeded 120 seconds. | 279 | ALARM_ID_DG_DIALYSATE_DRAIN_TIME_OUT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG dialysate drain time out | DG | DG Fault: Reservoir Drain
| PRS 863 | SA 340 SA 344 |  | SRSDG 888 |  | Call service to schedule a repair. | DG drain pump invalid direction |  | If the drain pump's direction is not 1 when draining. |  | DG drain pump invalid direction | If the drain pump's direction is not 1 when draining. | 280 | ALARM_ID_DG_DRAIN_PUMP_DIRECTION_INVALID |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain pump direction invalid | DG | DG Fault: DRP Direction Fault
| PRS 827 |  |  | SRSDG 836 |  | Acid concentrate low.\n- Replace acid concentrate bottle. | DG Acid concentration bottle low volume alarm |  | If the acid concentration bottle volume > 3240 (3340 - 100). |  | DG Acid concentration bottle low volume alarm | If the acid concentration bottle volume > 3240 (3340 - 100). | 281 | ALARM_ID_DG_ACID_BOTTLE_LOW_VOLUME |  | ALARM_PRIORITY_MEDIUM | 412 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | Acid concentration bottle low volume alarm | DG | Dial. Gen: Acid Conc. Low
| PRS 828 |  |  | SRSDG 837 |  | Bicarbonate concentrate low.\n- Replace bicarbonate concentrate bottle. | DG Bicarbonate concentration bottle low volume alarm |  | If the bicarbonate concentration bottle volume > 3680 (3780 - 100). |  | DG Bicarbonate concentration bottle low volume alarm | If the bicarbonate concentration bottle volume > 3680 (3780 - 100). | 282 | ALARM_ID_DG_BICARB_BOTTLE_LOW_VOLUME |  | ALARM_PRIORITY_MEDIUM | 413 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | Bicarbonate concentration bottle low volume alarm | DG | Dial. Gen: Bicarb Conc. Low
x | PRS 863 | SA 309 SA 310 SA 314 SA 315 |  | SRSDG 615 |  | Call service to schedule a repair. | DG load cells weight out of range |  | If the weight of the DG load cells > 4500.0 g or < 0.0 g. |  | DG load cells weight out of range | If the weight of the DG load cells > 4500.0 g or < 0.0 g. | 283 | ALARM_ID_DG_LOAD_CELL_WEIGHT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells weight out of range | DG | DG Fault: LC Range
x | PRS 863 | SA 317 SA 318 |  | SRSDG 383 |  | Call service to schedule a repair. | DG load cells primary/back up drift out of range |  | If redundant load cells deviate by more than 60g after tare for 5 seconds. |  | DG load cells primary/back up drift out of range | If redundant load cells deviate by more than 60g after tare for 5 seconds. | 284 | ALARM_ID_DG_LOAD_CELL_PRIMARY_BACKUP_DRIFT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells primary/back up drift out of range | DG | DG Fault: Redundant LC Check
[13 empty] | 285 | ALARM_ID_AVAILABLE_12 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | Available for use
[6 empty] | Concentrate cap removed from DG during cleaning.\n- Cleaning Process Terminated. | DG cap in wrong state |  | If the concentrate cap is not in the right position in the cleaning modes. |  | DG cap in wrong state | If the concentrate cap is not in the right position in the cleaning modes. | 286 | ALARM_ID_DG_CONCENTRATE_CAP_NOT_IN_PROPER_POSITION |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG concentrate cap not in proper position | DG | Cleaning: Conc Cap Removed
x | PRS 494 | SA 222 | SRSHD 1403 |  |  | Call service to schedule a repair. | HD processor clock speed checks against FPGA clock failure |  | If during initialization mode, the difference between the processor clock speed and the FPGA clock speed > 1. |  | HD processor clock speed checks against FPGA clock failure | If during initialization mode, the difference between the processor clock speed and the FPGA clock speed > 1. | 287 | ALARM_ID_HD_FPGA_CLOCK_SPEED_CHECK_FAILURE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD processor clock speed checks against FPGA clock failure | HD | HD Fault: Proc-FPGA Clock
|  | SA 491 SA 492 | SRSHD 562 |  |  | Call service to schedule a repair. | HD load cell inconsistency |  | If redundant load cells deviate from start to end of a reservoir use during treatment by more than 10g. |  | HD load cell inconsistency | If redundant load cells deviate from start to end of a reservoir use during treatment by more than 10g. | 288 | ALARM_ID_HD_LOAD_CELL_PRIMARY_BACKUP_DRIFT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD load cells primary/back up drift out of range | HD | DG Fault: LC Drift Check
Cleaning |  |  |  | SRSDG 928 |  | Dialysate cap removed from DG during cleaning.\n- Cleaning Process Terminated. | DG cap in wrong state |  | If the dialysate cap is not in the right position in the cleaning modes. |  | DG cap in wrong state | If the dialysate cap is not in the right position in the cleaning modes. | 289 | ALARM_ID_DG_DIALYSATE_CAP_NOT_IN_PROPER_POSITION |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG dialysate cap not closed | DG | Cleaning: Dialysate Cap Removed
x | PRS 494 | SA 233 | SRSHD 1584 |  |  | Call service to schedule a repair. | HD venous sensor FPGA fault |  | If the FPGA does not report fresh data for a certain period of time |  | HD venous sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 290 | ALARM_ID_HD_VENOUS_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD venous pressure sensor FPGA fault | HD | HD Fault: Ven. Pres. Com
Cleaning | [5 empty] | System flush in progress\n- Please wait for the process to complete.\n- Treatment will not be able to be started until process \ncompletes.\n- Process may be terminated by pressing the HD stop button. | DG flush started |  | If DG flush successfully started from the UI. |  | DG flush started | If DG flush successfully started from the UI. | 291 | ALARM_ID_HD_DISINFECT_FLUSH |  | ALARM_PRIORITY_LOW | 840 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD processor is in Disinfect Flush mode | HD | Cleaning: Flush In Progress
Cleaning | [5 empty] | Heat disinfect in progress\n- Please wait for the process to complete.\n- Treatment will not be able to be started until process \ncompletes.\n- Process may be terminated by pressing the HD stop button. | DG heat disinfect started |  | If DG heat disinfect successfully started from the UI. |  | DG heat disinfect started | If DG heat disinfect successfully started from the UI. | 292 | ALARM_ID_HD_DISINFECT_HEAT |  | ALARM_PRIORITY_LOW | 841 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD processor is in Disinfect Heat mode | HD | Cleaning: Heat Disinf In Progress
Cleaning | [5 empty] | Chemical disinfect flush in progress\n- Please wait for the process to complete\n- Treatment will not be able to be started until process \ncompletes.\n- Process may be terminated by pressing the HD stop button. | DG flush started |  | If DG flush successfully started from the UI. |  | DG flush started | If DG flush successfully started from the UI. | 293 | ALARM_ID_HD_DISINFECT_CHEM |  | ALARM_PRIORITY_LOW | 842 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD processor is in Disinfect Chemical mode | HD | Cleaning: Chem Flush In Progress
remove | [12 empty] | 294 | ALARM_ID_AVAILABLE_13 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Available for use
x | PRS 398 | SA 470 | SRSHD 1664 |  | SRSUI 1238 | Call service to schedule a repair. | UI cloud sync POST failure |  | If the UI POST reports cloud sync failure. |  | UI cloud sync POST failure | If the UI POST reports cloud sync failure. | 295 | ALARM_ID_HD_UI_POST_FAILURE_CLOUDSYNC | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI POST CloudSync failure | UI HD | HD POST: UI CloudSync
x | PRS 863 | SA 330 SA 331 |  | SRSDG 951 |  | Call service to schedule a repair. | DG TDi sensor FPGA fault |  | If the FPGA does not report fresh data for a certain period of time |  | DG TDi sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 296 | ALARM_ID_DG_TDI_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG TDi sensors FPGA fault | DG | DG Fault: TDi RTD Comm
x | PRS 863 | SA 326 SA 327 |  | SRSDG 950 |  | Call service to schedule a repair. | DG TRo sensor FPGA fault |  | If the FPGA does not report fresh data for a certain period of time |  | DG TRo sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 297 | ALARM_ID_DG_TRO_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG TRo sensors FPGA fault | DG | DG Fault: TRo RTD Comm
x | PRS 863 | SA 349 |  | SRSDG 948 |  | Call service to schedule a repair. | DG barometric sensor FPGA fault |  | If the FPGA does not report fresh data for a certain period of time |  | DG barometric sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 298 | ALARM_ID_DG_BARO_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG baro sensor FPGA fault | DG | DG Fault: Baro Pres Comm
x | PRS 863 | SA 483 |  | SRSDG 952 |  | Call service to schedule a repair. | DG invalid serial number |  | If the DG serial number is empty and does not have any characters. |  | DG invalid serial number | If the DG serial number is empty and does not have any characters. | 299 | ALARM_ID_DG_INVALID_SERIAL_NUMBER |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid serial number | DG | DG POST: Serial Number Check
x | PRS 398 | SA 473 | SRSHD 1582 |  |  | Call service to schedule a repair. | HD invalid serial number |  | If the HD serial number is empty and does not have any characters. |  | HD invalid serial number | If the HD serial number is empty and does not have any characters. | 300 | ALARM_ID_HD_INVALID_SERIAL_NUMBER |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD invalid serial number | HD | HD POST: Serial Number Check
remove redundant with SRSDG 866 | [12 empty] | 301 | ALARM_ID_AVAILABLE_14 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Alarm ID available
x | PRS 863 | SA 485 |  | SRSDG 126 |  | Call service to schedule a repair. | DG_CPU RAM error |  | If the RAM status indicates any of the following: Double bit error, Address decode failure, Address decode logic element failure, Read address parity failure, Write address parity failure. |  | DG_CPU RAM error | If the RAM status indicates any of the following: Double bit error, Address decode failure, Address decode logic element failure, Read address parity failure, Write address parity failure. | 302 | ALARM_ID_DG_CPU_RAM_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG processor RAM error | DG | DG Fault: Processor RAM error
x | PRS 863 | SA 345 |  | SRSDG 958 |  | Call service to schedule a repair. | DG drain pump direction FPGA fault |  | If the FPGA indicates an error for a certain period of time. |  | DG drain pump direction FPGA fault | If the FPGA indicates an error for a certain period of time. | 303 | ALARM_ID_DG_DRAIN_PUMP_DIRECTION_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain pump direction FPGA fault | DG | DG Fault: DRP Hall Commutation
x | PRS 398 | SA 197 | SRSHD 1587 |  |  | Call service to schedule a repair. | HD invalid usage reocord |  | If the CRC of the HD usage record does not match the calculated CRC. |  | HD invalid usage reocord | If the CRC of the HD usage record does not match the calculated CRC. | 304 | ALARM_ID_HD_INVALID_USAGE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD invalid usage record | HD | HD POST: Usage Rec NVRAM CRC
x | PRS 494 | SA 484 | SRSHD 1577 |  |  | Call service to schedule a repair. | HD_CPU RAM error |  | If the RAM status indicates any of the following: Double bit error, Address decode failure, Address decode logic element failure, Read address parity failure, Write address parity failure. |  | HD_CPU RAM error | If the RAM status indicates any of the following: Double bit error, Address decode failure, Address decode logic element failure, Read address parity failure, Write address parity failure. | 305 | ALARM_ID_HD_CPU_RAM_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD processor RAM error | HD | HD Fault: Processor RAM error
[6 empty] | Call service to schedule a repair. |  |  | CD1 and CD2 conductivity measurement during acid check is not within 5% of each other |  |  | CD1 and CD2 conductivity measurement during acid check is not within 5% of each other | 306 | ALARM_ID_DG_CD1_CD2_CHECK_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CD1 CD2 check fault | DG | DG Fault: CD1 CD2 Check
x | PRS 396 |  | SRSHD 1594 |  |  | HD power has been lost. \n- Treatment must be terminated.\n-Check all electrical connections.\n- The display and CPU will run off of the battery until \npower returns.\n-If power is returned prior to blood stopped countdown elapses, \nrinseback will be allowed. | HD AC power lost alarm |  | The HD Software shall trigger an alarm if an AC power loss is detected while the HD system is in pre-treatment, treatment, or post-treatment mode. |  | HD AC power lost alarm | The HD Software shall trigger an alarm if an AC power loss is detected while the HD system is in pre-treatment, treatment, or post-treatment mode. | 307 | ALARM_ID_HD_AC_POWER_LOST_IN_TREATMENT |  | ALARM_PRIORITY_HIGH | 100 | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | HD AC power lost second alarm | HD | Blood Stop: Loss of AC Power
x | PRS 863 | SA 348 |  | SRSDG 959 |  | Call service to schedule a repair. | DG CPi conductivity sensor transmits invalid characters |  | If the CPi conductivity sensor transmits invalid characters for a certain period of time. |  | DG CPi conductivity sensor transmits invalid characters | If the CPi conductivity sensor transmits invalid characters for a certain period of time. | 308 | ALARM_ID_DG_CPI_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CPi conductivity sensor invalid character received | DG | DG Fault: CPi Comm
x | PRS 863 | SA 348 |  | SRSDG 960 |  | Call service to schedule a repair. | DG CPo conductivity sensor transmits invalid characters |  | If the CPo conductivity sensor transmits invalid characters for a certain period of time. |  | DG CPo conductivity sensor transmits invalid characters | If the CPo conductivity sensor transmits invalid characters for a certain period of time. | 309 | ALARM_ID_DG_CPO_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CPo conductivity sensor invalid character received | DG | DG Fault: CPo Comm
x | PRS 863 | SA 347 |  | SRSDG 962 |  | Call service to schedule a repair. | DG CD1 conductivity sensor transmits invalid characters |  | If the CD1 conductivity sensor transmits invalid characters for a certain period of time. |  | DG CD1 conductivity sensor transmits invalid characters | If the CD1 conductivity sensor transmits invalid characters for a certain period of time. | 310 | ALARM_ID_DG_CD1_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CD1 conductivity sensor invalid character received | DG | DG Fault: CD1 Comm
x | PRS 863 | SA 347 |  | SRSDG 961 |  | Call service to schedule a repair. | DG CD2 conductivity sensor transmits invalid characters |  | If the CD2 conductivity sensor transmits invalid characters for a certain period of time. |  | DG CD2 conductivity sensor transmits invalid characters | If the CD2 conductivity sensor transmits invalid characters for a certain period of time. | 311 | ALARM_ID_DG_CD2_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CD2 conductivity sensor invalid character received | DG | DG Fault: CD2 Comm
Sean - review code and requirements use raw FMD flow (add DG to publish if not doing so already) long persistance |  |  | SRSHD 1600 |  |  | Dialysate in flow rate is low\n- Check dialysate line connection to the DG. | The dialysate inlet flow is too low. |  | If the dialysate inlet flow < 10 mL/min (raw) while in treatment mode or pre-treatment mode after prime and DPi is running. |  | The dialysate inlet flow is too low. | If the dialysate inlet flow < 10 mL/min (raw) while in treatment mode or pre-treatment mode after prime and DPi is running. | 312 | ALARM_ID_HD_DIAL_IN_FLOW_CHECK_FAILURE |  | ALARM_PRIORITY_MEDIUM | 400 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD Dialysate In flow rate to too low. | HD | Dial. Gen: Dialysate In Flow
x | PRS 852 |  |  | SRSDG 1064 SRSDG 1052 SRSDG 968 |  | Cleaning Mode Inlet water temperature too low.\n- Cleaning mode process has terminated.\n- Check inlet water before restarting cleaning mode. | DG cleaning mode inlet water temperature too low |  | If inlet water temperature is < 24 C. |  | DG cleaning mode inlet water temperature too low | If inlet water temperature is < 24 C. | 313 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_TEMP_TOO_LOW |  | ALARM_PRIORITY_MEDIUM | 602 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water temperature too low. | DG | Cleaning: Inlet Water Temp Low
x | PRS 855 |  |  | SRSDG 1063 SRSDG 1051 SRSDG 967 |  | Cleaning Mode Inlet water pressure too high.\n- Cleaning mode process has terminated.\n- Check inlet water before restarting cleaning mode. | DG cleaning mode inlet water pressure too high |  | If inlet water temperature is > 80 psig. |  | DG cleaning mode inlet water pressure too high | If inlet water temperature is > 80 psig. | 314 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_PRESSURE_TOO_HIGH |  | ALARM_PRIORITY_MEDIUM | 605 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water pressure too high. | DG | Cleaning: Inlet Water Press High
x | PRS 856 |  |  | SRSDG 1068 SRSDG 1056 SRSDG 966 |  | Cleaning Mode Inlet water pressure too low.\n- Cleaning mode process has terminated.\n- Check inlet water before restarting cleaning mode. | DG cleaning mode inlet water pressure too low |  | If inlet water pressure is < 14 psig. |  | DG cleaning mode inlet water pressure too low | If inlet water pressure is < 14 psig. | 315 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_PRESSURE_TOO_LOW |  | ALARM_PRIORITY_MEDIUM | 606 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water pressure too low. | DG | Cleaning: Inlet Water Press Low
x | PRS 863 | SA 495 |  | SRSDG 983 |  | Call service to schedule a repair. | Acid concentrate pump did not park successfully |  | If the park took more than 1 second. |  | Acid concentrate pump did not park successfully | If the park took more than 1 second. | 316 | ALARM_ID_DG_ACID_CONCENTRATE_PUMP_PARK_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG acid concentrate pump park fault. | DG | DG Fault: Acid CP Pump Park
x | PRS 863 | SA 496 |  | SRSDG 984 |  | Call service to schedule a repair. | Bicarb concentrate pump did not park successfully |  | If the park took more than 1 second. |  | Bicarb concentrate pump did not park successfully | If the park took more than 1 second. | 317 | ALARM_ID_DG_BICARB_CONCENTRATE_PUMP_PARK_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG bicarb concentrate pump part fault. | DG | DG Fault: Bicarb CP Pump Park
Blaine to check with michael |  |  | SRSHD 1609 |  |  | Unable to reduce fluid level.\n- Air trap may be overfilled.\n- Pressure transducer protector may be contaminated.\n- Treatment must be terminated without option for rinseback.\n- Call service to schedule a repair. | HD Air Pump operation timeout |  | If air pump has been turned on for greater than 10s. |  | HD Air Pump operation timeout | If air pump has been turned on for greater than 10s. | 318 | ALARM_ID_HD_AIR_PUMP_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | HD Air Pump timeout alarm | HD | HD Fault: Air Pump Timeout
Cleaning |  |  |  | SRSDG 1040 |  | Collect chemical residuals water sample from drain line. | Collect sample to check the chemical disinfectant acid |  | Chemical disinfect flush sample alarm is prompted. |  | Collect sample to check the chemical disinfectant acid | Chemical disinfect flush sample alarm is prompted. | 319 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_FLUSH_SAMPLE |  | ALARM_PRIORITY_LOW | 752 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | DG chemical disinfect flush flush sample. | DG | Cleaning: Chem Residual Sampling
Cleaning |  |  |  | SRSDG 1043 |  | During Chemical Disinfection Flush\n-user residual sampling timed out. | The user did not proceed with collecting chemical disinfect flush sample for more than 30 minutes |  | Chemical disinfect flush sample was not collected form more than 30 minutes. |  | The user did not proceed with collecting chemical disinfect flush sample for more than 30 minutes | Chemical disinfect flush sample was not collected form more than 30 minutes. | 320 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_SAMPLE_TIMEOUT |  | ALARM_PRIORITY_LOW | 752 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG chemical disinfect flush sample timeout. | DG | Cleaning: Chem Residual Timeout
x | PRS 863 | SA 356 |  | SRSDG 1069 |  | Call service to schedule a repair. | Outlet reactor not healthy |  | If the outlet reactor are not healthy. |  | Outlet reactor not healthy | If the outlet reactor are not healthy. | 321 | ALARM_ID_DG_OUTLET_UV_REACTOR_NOT_HEALTHY |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG outlet UV reactor not healthy | DG | DG Fault: UVo Error
[6 empty] | Call service to schedule a repair. | UI POST Invalid Year |  | If the UI POST reports a invalid year. |  | UI POST Invalid Year | If the UI POST reports a invalid year. | 322 | ALARM_ID_HD_UI_POST_FAILURE_INVALID_YEAR |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | HD UI POST Year check failure | UI HD | HD POST Alarm: UI Invalid Year

## Property Definitions
Property Name | Definition |  | Column | Description
AlarmID | Unique alarm identification number |  | PRS ID | PRS requirement ID associated with the alarm.
Alarm Name | Name of the alarm |  | HD SRS ID | HD SRS requirement ID (if applicable) associated with the alarm.
Priority | Priority of the alarm (High, Medium, Low) |  | DG SRS ID | DG SRS requirement ID (if applicable) associated with the alarm.
Rank | Within the same priority, higher rank alarms have higher priority |  | Device | Sub-system that detects the alarm.
Fault | The alarm activation causes an HD fault |  | Description | Detailed description of the alarm.
DGFault | The alarm activation causes an DG fault |  | Displayed Title | Alarm title displayed prominantly at top of alarm dialog when the alarm is top priority.
Stop | The alarm activation causes a controlled stop |  | Displayed Instructions | Instructions displayed in alarm dialog when the alarm is top priority.
Clear Immediate | Alarm condition is cleared immediately after alarm activation |  | Trigger Conditions | Condition(s) required to trigger the alarm.
Non Recoverable | Alarm cannot be cleared (unrecoverable) |  | AlarmID | Unique alarm ID #.
No Resume | Alarm activation prevents treatment resume |  | Alarm Name | Enumeration name (code reference) for the alarm.
No Rinseback | Alarm activation prevents rinseback |  | Priority | Priority of the alarm (high, medium or low per 60601-1-8.
No End Treatment | Alarm activation prevents ending treatment |  | Rank | Rank of the alarm (used to determine which alarm is top priority) (low rank=high priority).
No New Treatment | Alarm activation prevents new treatment |  | Fault | Does HD f/w consider the alarm to be a system fault.
No Dialysate Recirc | Alarm activation prevents dialysate re-circulation |  | DGFault | Does DG f/w consider the alarm to be a system fault (will not prevent HD from some activities).
Clear Top Only | User acknowledgment clears only displayed alarm |  | Stop | Does the HD f/w need to pause treatment (or perhaps other activities) when alarm is active.
Escalate In | Alarm will escalate if not cleared within this time (zero indicates no escalation) |  | Clear Immediate | Should the f/w immediately clear the alarm condition after trigger (not clear the alarm itself).
Escalate To | ID of the alarm that this alarm will be escalated to |  | Non Recoverable | Is the alarm recoverable (can user clear the alarm).
|  |  | No Resume | Should treatment resume be prevented when the alarm is active.
|  |  | No Rinseback | Should rinseback be prevented when the alarm is active.
|  |  | No End Treatment | Should ending treatment be prevented when the alarm is active.
Category | Priority |  | No Blood Recirc | Should blood recirculation be prevented when the alarm is active (treatment is paused). Dialyzer is bypassed regardless.
HD POST | High |  | No Dialysate Recirc | Should dialysate recirculation be prevented when the alarm is active (treatment is paused). Dialyzer is bypassed regardless.
HD Machine Alarm | High |  | Clear Top Only | If user acknowledges alarm, should all active alarms be cleared or just the alarm top (displayed) alarm.
DG POST DG Machine Alarm | High |  | TreatmentLog | Should alarm be included in treatment log.
Blood Stopped - No Rinseback Allowed | HIgh |  | Escalate In | Duration (in seconds) until alarm escalates to a higher priority alarm (0 indicates no escalation).
Blood Stopped - Rinseback Allowed | High |  | Escalate To | Higher priority alarm to escalate to when escalation period expires (None indicates no escalation).
End Pre-Tx | Medium |  | Doxygen Title | Used for design documentation purposes.
Concentrate Warning | Low
Concentrate Alarm | Medium
Saline Warning | Low
Saline Alarm | Medium
Dialysate Alarms (Temp, conductivity, recirc?) | Medium
Other Procedural (setup, pre-tx, etc) | Medium/Low
Disinfection Alarms | Medium
HD Machine Alert | Low
DG Machine Alert | Low
Blood Pressure Alarms | N/A - not supported
Priority | Type | Mode