## Alarms
Action | PRS ID2 | SA | HD SRS ID | DG SRS ID | UI SRS ID | Device | Displayed Title | Char count | Displayed Instructions | Char Count2 | Description | Trigger Conditions | AlarmID | Alarm Name | Alarm ID Check | Priority | Rank | Fault | DGFault | Stop | Clear Immediate | Non Recoverable | No Resume | No Rinseback | No End Treatment | No Blood Recirc | No Dialysate Recirc | Clear Top Only | TreatmentLog | Doxygen Title
N/A | N/A |  | N/A | N/A |  | FALSE | Indication for when no alarms is a possible situation |  | N/A |  | ` | N/A | 0 | ALARM_ID_NO_ALARM |  | ALARM_PRIORITY_NONE | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 398 | SA 213 | SRSHD 652 |  |  | FALSE | HD stuck button POST failure. Stop or Off button detected to be pressed for at least 1 second during test shortly after power up |  | HD |  | HD POST: Stuck Button | Stop or On button pressed during startup\n-Release button and power cycle\n-Call service to schedule a repair if problem persists | 1 | ALARM_ID_HD_STUCK_BUTTON_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 207 | SRSHD 995 SRSHD 1357 |  |  | FALSE | HD FPGA POST failure |  | HD |  | HD POST: FPGA Compatibility | Call service to schedule a repair. | 2 | ALARM_ID_HD_FPGA_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 287 |  | SRSDG 602 SRSDG 814 SRSDG 1061 |  | FALSE | DG FPGA POST failure |  | DG |  | DG POST: FPGA Compatibility | Call service to schedule a repair. | 3 | ALARM_ID_DG_FPGA_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
Sean - check if DG and HD POST test is the same in code. Consider updating trigger conditions | PRS 398 | SA 208 SA 209 SA 210 SA 211 | SRSHD 650 |  |  | FALSE | HD Watchdog POST failure |  | HD |  | HD POST: WD Safestate | Call service to schedule a repair. | 4 | ALARM_ID_HD_WATCHDOG_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
Sean - check if DG and HD POST test is the same in code. Consider updating trigger conditions | PRS 863 | SA 288 SA 290 SA 289 |  | SRSDG 53 |  | FALSE | DG Watchdog POST failure |  | DG |  | DG POST: WD Safestate | Call service to schedule a repair. | 5 | ALARM_ID_DG_WATCHDOG_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 398 | SA 204 | SRSHD 1484 |  |  | FALSE | HD UI communication POST failure. UI failed to communicate within a reasonable time after power up |  | HD |  | HD POST: Proc-UI Comm | Call service to schedule a repair. | 6 | ALARM_ID_HD_UI_COMM_POST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 872 | SA 457 | SRSHD 1595 |  |  | FALSE | HD RTC battery low |  | HD |  | HD Alarm (Low): RTC Bat Low | Call service to schedule a repair. | 7 | ALARM_ID_HD_RTC_BATTERY_LOW |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 398 | SA 471 | SRSHD 632 |  |  | FALSE | HD accelerometer failed POST |  | HD |  | HD POST: Acc Vector | Call service to schedule a repair. | 8 | ALARM_ID_HD_ACCELEROMETER_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 472 |  | SRSDG 323 |  | FALSE | DG accelerometer failed POST |  | DG |  | DG POST: Acc Vector | Call service to schedule a repair. | 9 | ALARM_ID_DG_ACCELEROMETER_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 398 | SA 480 | SRSHD 1478 |  |  | FALSE | HD blood leak sensor setting embedded mode failure. |  | HD |  | HD POST: BLD Config Emb Mode | Call service to schedule a repair. | 10 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_EMBEDDED_MODE_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
Dialysate Temp | [5 empty] | FALSE | HD dialysate temperature below target temperature or below safety temp |  | HD |  | Dial. Gen: Dialysate Temp Low | Dialysate temperature is low\n- Treatment is paused and will be ready to resume when \ndialysate is within target range. | 11 | ALARM_ID_HD_DIALYSATE_TEMP_BELOW_TARGET_TEMP |  | ALARM_PRIORITY_MEDIUM | 402 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE
x | PRS 398 | SA 195 | SRSHD 658 |  |  | FALSE | HD firmware image integrity POST test failed |  | HD |  | HD POST: HD Image | Call service to schedule a repair. | 12 | ALARM_ID_HD_INTEGRITY_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 282 |  | SRSDG 813 |  | FALSE | DG firmware image integrity POST test failed |  | DG |  | DG POST: DG Image | Call service to schedule a repair. | 13 | ALARM_ID_DG_INTEGRITY_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 284 |  | SRSDG 957 |  | FALSE | DG invalid usage record CRC |  | DG |  | DG POST: Usage Rec NVRAM CRC | Call service to schedule a repair. | 14 | ALARM_ID_DG_INVALID_USAGE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
Dara - remove Alarm ID 15 (Duplicate with Alarm ID 304 - SRS 1587) | [5 empty] | FALSE | HD dialysate temperature above high safety. |  | HD |  | Dial. Gen: Dialysate Temp High | Dialysate temperature is high\n-Treatment is paused and will be ready to resume when \ndialysate is within target range. | 15 | ALARM_ID_HD_DIALYSATE_TEMP_ABOVE_SAFETY_TEMP |  | ALARM_PRIORITY_MEDIUM | 401 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 398 | SA 212 | SRSHD 656 |  |  | FALSE | HD alarm audio failed POST |  | HD |  | HD POST: Primary Alarm Current | Call service to schedule a repair. | 16 | ALARM_ID_HD_ALARM_AUDIO_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
Sean to review and update |  |  | SRSHD 1677 |  |  | FALSE | HD UI POST failed |  | HD |  | HD POST: UI POST Not Complete | Call service to schedule a repair. | 17 | ALARM_ID_HD_UI_POST_TIMEOUT |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
[6 empty] | FALSE | HD didn't get ACK on message to DG that required acknowledgment |  | HD |  | Loss of DG Comm (DG ACK) | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | 18 | ALARM_ID_HD_CAN_MESSAGE_NOT_ACKED_BY_DG |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
Dialysate Temp |  |  | SRSHD 901 |  |  | FALSE | HD dialysate temperature above target temperature. |  | HD |  | Dial. Gen: Dialysate Temp High | Dialysate temperature is high\n- Treatment is paused and will be ready to resume when \ndialysate is within target range. | 19 | ALARM_ID_HD_DIALYSATE_TEMP_ABOVE_TARGET_TEMP |  | ALARM_PRIORITY_MEDIUM | 401 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE
x | PRS 863 | SA 369 |  | SRSDG 16 |  | FALSE | DG conductivity sensors invalid calibration record |  | DG |  | DG POST: Conductivity Cal | Call service to schedule a repair. | 20 | ALARM_ID_DG_COND_SENSORS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 369 |  | SRSDG 936 |  | FALSE | DG drain line volume invalid calibration record |  | DG |  | DG POST: Drain Line Vol Cal | Call service to schedule a repair. | 21 | ALARM_ID_DG_DRAIN_LINE_VOLUME_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 369 |  | SRSDG 935 |  | FALSE | DG reservoirs invalid calibration record |  | DG |  | DG POST: Reservoir Vol Cal | Call service to schedule a repair. | 22 | ALARM_ID_DG_RESERVOIRS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 369 |  | SRSDG 934 |  | FALSE | DG acid concentrate calibration record |  | DG |  | DG POST: Acid Conc Cal | Call service to schedule a repair. | 23 | ALARM_ID_DG_ACID_CONCENTRATE_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 369 |  | SRSDG 933 |  | FALSE | DG bicarb concentrate calibration record |  | DG |  | DG POST: Bicarb Conc Cal | Call service to schedule a repair. | 24 | ALARM_ID_DG_BICARB_CONCENTRATE_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 369 |  | SRSDG 817 |  | FALSE | DG accelerometer invalid calibration record |  | DG |  | DG POST: Acc Cal | Call service to schedule a repair. | 25 | ALARM_ID_DG_ACCELEROMETERS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 398 | SA 368 | SRSHD 1576 |  |  | FALSE | HD accelerometer invalid calibration record |  | HD |  | HD POST: Acc Cal | Call service to schedule a repair. | 26 | ALARM_ID_HD_ACCELEROMETERS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
Move to wet self-test - procedural |  |  | SRSHD 1428 |  |  | FALSE | HD blood leak sensor zero and self test sequence failed |  | HD |  | Process: Pre-Tx Wet Self-Test | Self-test detected has a problem with the blood leak sensor\n- Check cartridge installation on to machine\n- If problem persists, terminate pre-treatment and restart \nwith a new cartridge. | 27 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_ZERO_SEQUENCE_FAILED |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 863 | SA 319 SA 320 |  | SRSDG 864 |  | FALSE | DG two wire sensors FPGA fault |  | DG |  | DG Fault: Conductivity RTD Comm | Call service to schedule a repair. | 28 | ALARM_ID_DG_TWO_WIRE_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 398 | SA 368 | SRSHD 1482 |  |  | FALSE | HD heparin force sensor invalid calibration record |  | HD |  | HD POST: SP DAC Cal | Call service to schedule a repair. | 29 | ALARM_ID_HD_HEPARIN_FORCE_SENSOR_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 416 | SRSHD 1704 |  |  | FALSE | HD Software fault. Software found itself in an unexpected state |  | HD |  | HD Fault: SW Invalid State | Call service to schedule a repair. | 30 | ALARM_ID_HD_SOFTWARE_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 241 SA 242 | SRSHD 515 |  |  | FALSE | HD blood pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running |  | HD |  | HD Fault: BP Current | Call service to schedule a repair. | 31 | ALARM_ID_HD_BLOOD_PUMP_MC_CURRENT_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 239 | SRSHD 1087 |  |  | FALSE | HD blood pump failed motor off check. Measured speed while commanded off |  | HD |  | HD Fault: BP Speed | Call service to schedule a repair. | 32 | ALARM_ID_HD_BLOOD_PUMP_OFF_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 240 SA 244 | SRSHD 1079 |  |  | FALSE | HD blood pump failed motor direction check. Measured vs. commanded |  | HD |  | HD Fault: BP Direction | Call service to schedule a repair. | 33 | ALARM_ID_HD_BLOOD_PUMP_MC_DIRECTION_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 246 | SRSHD 1081 |  |  | FALSE | HD blood pump failed rotor speed check. Mismatch with rotor and motor speeds |  | HD |  | HD Fault: BP Rotor | Blood pump rotor failure.\n- Check the pump rotor for secure connection. | 34 | ALARM_ID_HD_BLOOD_PUMP_ROTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 250 SA 251 | SRSHD 542 |  |  | FALSE | HD dialysis inlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running |  | HD |  | HD Fault: DPi Current | Call service to schedule a repair. | 35 | ALARM_ID_HD_DIAL_IN_PUMP_MC_CURRENT_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 248 | SRSHD 1088 |  |  | FALSE | HD dialysis inlet pump failed motor off check. Measured speed while commanded off |  | HD |  | HD Fault: DPi Speed | Call service to schedule a repair. | 36 | ALARM_ID_HD_DIAL_IN_PUMP_OFF_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 249 SA 253 | SRSHD 1050 |  |  | FALSE | HD dialysis inlet pump failed motor direction check. Measured vs. commanded |  | HD |  | HD Fault: DPi Direction | Call service to schedule a repair. | 37 | ALARM_ID_HD_DIAL_IN_PUMP_MC_DIRECTION_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 255 | SRSHD 1082 |  |  | FALSE | HD dialysis inlet pump failed rotor speed check. Mismatch with rotor and motor speeds |  | HD |  | HD Fault: DPi Rotor | Dialysate in pump rotor failure.\n- Check the pump rotor for secure connection. | 38 | ALARM_ID_HD_DIAL_IN_PUMP_ROTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 258 SA 259 | SRSHD 554 |  |  | FALSE | HD dialysis outlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running |  | HD |  | HD Fault: DPo Current | Call service to schedule a repair. | 39 | ALARM_ID_HD_DIAL_OUT_PUMP_MC_CURRENT_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 256 | SRSHD 1089 |  |  | FALSE | HD dialysis outlet pump failed motor off check. Measured speed while commanded off |  | HD |  | HD Fault: DPo Speed | Call service to schedule a repair. | 40 | ALARM_ID_HD_DIAL_OUT_PUMP_OFF_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 257 SA 261 | SRSHD 1051 |  |  | FALSE | HD dialysis outlet pump failed motor direction check. Measured vs. commanded |  | HD |  | HD Fault: DPo Direction | Call service to schedule a repair. | 41 | ALARM_ID_HD_DIAL_OUT_PUMP_MC_DIRECTION_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 263 | SRSHD 1083 |  |  | FALSE | HD dialysis outlet pump failed rotor speed check. Mismatch with rotor and motor speeds |  | HD |  | HD Fault: DPo Rotor | Dialysate out pump rotor failure.\n- Check the pump rotor for secure connection. | 42 | ALARM_ID_HD_DIAL_OUT_PUMP_ROTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 357 |  | SRSDG 954 |  | FALSE | DG main primary heater FPGA fault |  | DG |  | DG Fault: Iso ADC Comm | Call service to schedule a repair. | 43 | ALARM_ID_DG_MAIN_PRIMARY_HEATER_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 494 | SA 226 | SRSHD 1001 |  |  | FALSE | HD UI communication timeout |  | HD |  | HD Fault: UI Comm | Call service to schedule a repair. | 44 | ALARM_ID_HD_UI_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 226 SA 477 | SRSHD 967 |  |  | FALSE | HD too many bad communications CRC |  | HD |  | HD Fault: CAN CRC Com | Call service to schedule a repair. | 45 | ALARM_ID_HD_COMM_TOO_MANY_BAD_CRCS |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 226 SA 477 | SRSHD 964 |  |  | FALSE | HD didn't get ACK on message to UI that required acknowledgment |  | HD |  | HD Fault: CAN ACK UI Com | Call service to schedule a repair. | 46 | ALARM_ID_HD_CAN_MESSAGE_NOT_ACKED_BY_UI |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 476 | SRSHD 907 |  |  | FALSE | Available for use | [5 empty] | 47 | ALARM_ID_AVAILABLE_16 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 494 | SA 452 | SRSHD 1029 |  |  | FALSE | HD ultrafiltration volume accuracy error during treatment |  | HD |  | HD Fault: UF Rate Tare Err | Call service to schedule a repair. | 48 | ALARM_ID_HD_UF_VOLUME_ACCURACY_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 224 | SRSHD 1356 SRSHD 1490 |  |  | FALSE | HD FPGA communication down for too long |  | HD |  | HD Fault: FPGA Com | Call service to schedule a repair. | 49 | ALARM_ID_HD_FPGA_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 481 |  | SRSDG 26 |  | FALSE | DG FPGA not accepting commanded valve states |  | DG |  | DG Fault: Valve Control | Call service to schedule a repair. | 50 | ALARM_ID_DG_VALVE_CONTROL_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 494 | SA 239 SA 243 | SRSHD 1098 |  |  | FALSE | HD blood pump failed motor speed check. Measured vs. commanded |  | HD |  | HD Fault: BP Speed | Call service to schedule a repair. | 51 | ALARM_ID_HD_BLOOD_PUMP_MOTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 252 | SRSHD 1099 |  |  | FALSE | HD dialysate inlet pump failed motor speed check. Measured vs. commanded |  | HD |  | HD Fault: DPi Speed | Call service to schedule a repair. | 52 | ALARM_ID_HD_DIAL_IN_PUMP_MOTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 260 SA 256 | SRSHD 1198 SRSHD 1431 SRSHD 1670 |  |  | FALSE | HD dialysate outlet pump failed motor speed check. Measured vs. commanded |  | HD |  | HD Fault: DPo Speed | Call service to schedule a repair. | 53 | ALARM_ID_HD_DIAL_OUT_PUMP_MOTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 501 | SRSHD 1705 |  |  | FALSE | HD critical data integrity check failed |  | HD |  | HD Fault: Critical Data Err | Call service to schedule a repair. | 54 | ALARM_ID_HD_CRITICAL_DATA_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
| PRS 863 | SA 500 |  | SRSDG 921 |  | FALSE | DG critical data integrity check failed |  | DG |  | DG Fault: Critical Data Err | Call service to schedule a repair. | 55 | ALARM_ID_DG_CRITICAL_DATA_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 494 | SA 232 | SRSHD 1526 |  |  | FALSE | HD accelerometer error (no readings or FPGA reports error) |  | HD |  | HD Fault: Acc Com | Call service to schedule a repair. | 56 | ALARM_ID_HD_ACCELEROMETER_FAILURE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 346 |  | SRSDG 869 SRSDG 870 |  | FALSE | DG accelerometer error (no readings or FPGA reports error) |  | DG |  | DG Fault: Acc Com | Call service to schedule a repair. | 57 | ALARM_ID_DG_ACCELEROMETER_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 494 | SA 384 SA 385 SA 386 SA 387 | SRSHD 1340 |  |  | FALSE | HD valve homing failed |  | HD |  | HD Fault: Valve Home | Call service to schedule a repair. | 58 | ALARM_ID_HD_VALVE_HOMING_FAILED |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 264 SA 266 SA 268 SA 270 | SRSHD 1338 |  |  | FALSE | HD valve transition time out |  | HD |  | HD Fault: Valve Transition Err | Call service to schedule a repair. | 59 | ALARM_ID_HD_VALVE_TRANSITION_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 502 SA 505 SA 503 SA 504 | SRSHD 1515 SRSHD 1516 SRSHD 1632 |  |  | FALSE | HD valve not functional |  | HD |  | HD Fault: Valve Control | Call service to schedule a repair. | 60 | ALARM_ID_HD_VALVE_NOT_FUNCTIONAL |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 265 SA 267 SA 269 SA 271 | SRSHD 1339 |  |  | FALSE | HD valve current out of range |  | HD |  | HD Fault: Valve Current | Call service to schedule a repair. | 61 | ALARM_ID_HD_VALVE_CURRENT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 264 SA 266 SA 268 SA 270 | SRSHD 1336 |  |  | FALSE | HD valve position out of target |  | HD |  | HD Fault: Valve Position | Call service to schedule a repair. | 62 | ALARM_ID_HD_VALVE_POSITION_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 453 | SRSHD 1586 |  |  | FALSE | HD blood leak FPGA fault |  | HD |  | HD Fault: BLD Com | Call service to schedule a repair. | 63 | ALARM_ID_HD_BLOOD_LEAK_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 350 |  | SRSDG 874 |  | FALSE | DG barometric pressure sensor out of range |  | DG |  | DG Fault: Baro Range | Call service to schedule a repair. | 64 | ALARM_ID_DG_BARO_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 400 |  | SRSHD 1531 |  |  | FALSE | HD DG requests DG command with invalid parameter fault |  | HD |  | Loss of DG Comm (Reject) | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | 65 | ALARM_ID_HD_DG_COMMAND_INVALID_PARAMETER_FAULT |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 398 | SA 480 | SRSHD 1427 SRSHD 1429 SRSHD 1631 |  |  | FALSE | HD blood leak sensor set point set failure |  | HD |  | HD POST: BLD Config Set Point | Call service to schedule a repair. | 66 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_SET_POINT_SET_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
No cartridge self-test (procedural) |  |  | SRSHD 1511 |  |  | FALSE | HD blood pump occlusion self-test failure alarm |  | HD |  | Process: Pre-Tx Remove Cartridge | Cartridge has been detected.\n- Open door\n- Remove cartridge\n- Close door\n- Resume | 67 | ALARM_ID_HD_BP_OCCLUSION_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 700 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
agree to change rank and re-assess similar to temperature alarm Blaine |  |  | SRSHD 1426 |  |  | FALSE | HD active reservoir recirculation out of range |  | HD |  | Dial. Gen: Dialysate not ready | Dialysate is not yet ready.\n- Treatment is paused and will be ready to resume when \ndialysate is ready for use. | 68 | ALARM_ID_HD_ACTIVE_RESERVOIR_RECIRCULATION_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 421 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 398 | SA 368 | SRSHD 532 |  |  | FALSE | HD blood leak sensor invalid calibration record |  | HD |  | HD POST: BLD Cal | Call service to schedule a repair. | 69 | ALARM_ID_HD_BLOOD_LEAK_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
This should be in-treatment alarm? Not machine alarm? See alarm ID 275 for machine alarm (valid range check) agree - make this in-treatment alarms not faults Blaine to update - not a fault Sean to update the range to what is in the SRS |  |  | SRSHD 737 |  |  | FALSE | HD arterial pressure self-test failure alarm |  | HD |  | Process: Pre-Tx Remove Cartridge | Cartridge has been detected.\n- Open door\n- Remove cartridge\n- Close door\n- Resume | 70 | ALARM_ID_HD_ARTERIAL_PRESSURE_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 700 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
This should be in-treatment alarm? Not machine alarm? See alarm ID 275 for machine alarm (valid range check) agree - make this in-treatment alarms not faults Blaine to update - not a fault Sean to update the range to what is in the SRS |  |  | SRSHD 738 |  |  | FALSE | HD venous pressure self-test failure alarm |  | HD |  | Process: Pre-Tx Remove Cartridge | Cartridge has been detected.\n- Open door\n- Remove cartridge\n- Close door\n- Resume | 71 | ALARM_ID_HD_VENOUS_PRESSURE_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 700 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 400 |  | SRSHD 1126 |  |  | FALSE | HD No load cell data message receive at least once every 2 seconds |  | HD |  | Loss of DG Comm (LC) | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | 72 | ALARM_ID_HD_NEW_LOAD_CELL_DATA_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 400 |  | SRSHD 1573 |  |  | FALSE | HD No dialysate temperature data message receive at least once every 2 seconds |  | HD |  | Loss of DG Comm (TDI/TRO) | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | 73 | ALARM_ID_HD_NEW_DIALYSATE_TEMP_DATA_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 364 |  | SRSDG 600 |  | FALSE | DG inlet UV reactor is on with no flow |  | DG |  | DG Fault: UVi On No Flow | Call service to schedule a repair. | 74 | ALARM_ID_DG_INLET_UV_REACTOR_ON_WITH_NO_FLOW |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 494 | SA 469 | SRSHD 1488 |  |  | FALSE | HD syringe pump self-test failure alarm |  | HD |  | HD Fault: Syringe DAC | Call service to schedule a repair. | 75 | ALARM_ID_HD_SYRINGE_PUMP_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
Sean to update code | PRS 494 | SA 217 SA 218 SA 219 | SRSHD 635 SRSHD 1678 SRSHD 1681 |  |  | FALSE | HD monitored voltage is out of range |  | HD |  | HD Fault: Voltage Monitor | Call service to schedule a repair. | 76 | ALARM_ID_HD_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
Sean to update code | PRS 863 | SA 295 SA 296 SA 297 |  | SRSDG 46 SRSDG 866 SRSDG 1088 |  | FALSE | DG monitored voltage is out of range |  | DG |  | DG Fault: Voltage Monitor | Call service to schedule a repair. | 77 | ALARM_ID_DG_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 494 | SA 275 | SRSHD 1324 |  |  | FALSE | HD syringe pump direction (from encoder) error |  | HD |  | HD Fault: Syringe Pump Enc Dir | Call service to schedule a repair. | 78 | ALARM_ID_HD_SYRINGE_PUMP_ENCODER_DIRECTION_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 489 | SRSHD 1507 |  |  | FALSE | HD syringe pump direction (from controller) error |  | HD |  | HD Fault: Syringe Pump Drv Dir | Call service to schedule a repair. | 79 | ALARM_ID_HD_SYRINGE_PUMP_CONTROLLER_DIRECTION_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 272 | SRSHD 1317 |  |  | FALSE | HD syringe pump fault reported by FPGA |  | HD |  | HD Fault: Syringe Pump Driver | Call service to schedule a repair. | 80 | ALARM_ID_HD_SYRINGE_PUMP_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 488 | SRSHD 1330 |  |  | FALSE | HD syringe pump over travel error |  | HD |  | HD Fault: Syringe Pump OT | Call service to schedule a repair. | 81 | ALARM_ID_HD_SYRINGE_PUMP_OVER_TRAVEL_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 277 | SRSHD 1506 |  |  | FALSE | HD syringe pump DAC write failure |  | HD |  | HD Fault: Syringe Pump DAC Err | Call service to schedule a repair. | 82 | ALARM_ID_HD_SYRINGE_PUMP_DAC_WRITE_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 415 | SRSHD 1328 |  |  | FALSE | HD syringe pump is running while the BP is off |  | HD |  | HD Fault: Syringe Pump On Err | Call service to schedule a repair. | 83 | ALARM_ID_HD_SYRINGE_PUMP_RUNNING_WHILE_BP_OFF_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 286 |  | SRSDG 815 |  | FALSE | DG set RTC year is invalid |  | DG |  | DG POST: HD RTC Time | Call service to schedule a repair. | 84 | ALARM_ID_DG_SET_RTC_YEAR_INVALID |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
revisit - PRS has this as low priority - why did SRS make this high priority? Probably needs to be high priority as blood pump needs to stop - update PRS to make high priority | [5 empty] | FALSE | HD pump track latch opened alarm |  | HD |  | Blood Stop: Cartridge Lever Open | Cartridge Lock Lever is not closed.\n- Open front door and push lever all the way down.\n- Close door and press RESUME button to continue. | 85 | ALARM_ID_HD_PUMP_TRACK_LATCH_OPENED |  | ALARM_PRIORITY_HIGH | 209 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 199 | SRSHD 1479 |  |  | FALSE | HD set RTC year is invalid |  | HD |  | HD POST: HD RTC Time | Call service to schedule a repair. | 86 | ALARM_ID_HD_SET_RTC_YEAR_INVALID |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 369 |  | SRSDG 929 |  | FALSE | DG heating invalid calibration record |  | DG |  | DG POST: Heat Cal | Call service to schedule a repair. | 87 | ALARM_ID_DG_HEATING_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
[6 empty] | FALSE | DG concentrate pumps hall sensor out of range |  | DG |  | DG Fault: CP1 CP2 Speed Range | Call service to schedule a repair. | 88 | ALARM_ID_DG_CONC_PUMP_HALL_SENSOR_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 365 |  | SRSDG 599 |  | FALSE | DG outlet UV reactor on with no flow |  | DG |  | DG Fault: UVo On No Flow | Call service to schedule a repair. | 89 | ALARM_ID_DG_OUTLET_UV_REACTOR_ON_WITH_NO_FLOW |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 307 SA 308 |  | SRSDG 604 |  | FALSE | DG load cells A1/B1 FPGA fault |  | DG |  | DG Fault: LC A1/B1 Comm | Call service to schedule a repair. | 90 | ALARM_ID_DG_LOAD_CELL_A1_B1_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 400 |  | SRSHD 1408 |  |  | FALSE | HD No dialysate flow data receive in the last 3 seconds |  | HD |  | Loss of DG Comm (FMD) | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | 91 | ALARM_ID_HD_DIALYSATE_FLOW_DATA_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
Remove (redundant with SRSDG 913) | [12 empty] | 92
[6 empty] | FALSE | Available for use | [5 empty] | 92 | ALARM_ID_AVAILABLE_1 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 380 |  | SRSDG 430 |  | FALSE | DG outlet primary conductivity out of range |  | DG |  | DG Fault: RO Permeate Cond High | Call service to schedule a repair. | 93 | ALARM_ID_DG_OUTLET_PRIMARY_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 303 SA 304 SA 305 SA 306 |  | SRSDG 629 |  | FALSE | DG pressure out of range |  | DG |  | DG Fault: Pressure Sensor Range | Call service to schedule a repair. | 94 | ALARM_ID_DG_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 486 |  | SRSDG 907 |  | FALSE | DG watchdog expired |  | DG |  | DG Fault: WD Error | Call service to schedule a repair. | 95 | ALARM_ID_DG_WATCHDOG_EXPIRED |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 406 |  |  | SRSDG 410 |  | TRUE | DG inlet water temperature in high range |  | DG |  | Dial. Gen: Inlet Water Temp High | Inlet water temperature is high.\n- Treatment is paused\n- Check the incoming water supply. | 96 | ALARM_ID_DG_INLET_WATER_TEMPERATURE_IN_HIGH_RANGE |  | ALARM_PRIORITY_MEDIUM | 432 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
Check that software behavior and PRS behavior are aligned | PRS 543 |  |  | SRSDG 397 |  | TRUE | DG fill conductivity out of range |  | DG |  | Dial. Gen: Dialysate Cond. | Dialysate conductivity is out of range.\n- Check acid and bicarb concentrate is correct and not empty. | 97 | ALARM_ID_DG_FILL_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 420 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE
x | PRS 872 | SA 227 SA 228 | SRSHD 1395 |  |  | FALSE | HD battery communication fault |  | HD |  | HD Alarm (Low): Battery Com | Call service to schedule a repair. | 98 | ALARM_ID_HD_BATTERY_COMM_FAULT |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 494 | SA 274 | SRSHD 1683 |  |  | FALSE | HD syringe pump stall alarm |  | HD |  | HD Fault: Syringe Pump Stall | Call service to schedule a repair. | 99 | ALARM_ID_HD_SYRINGE_PUMP_STALL |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 416 | SRSHD 1679 |  |  | FALSE | HD no cartridge self-test timeout |  | HD |  | HD Fault: Timeout No Cart ST | Call service to schedule a repair. | 100 | ALARM_ID_HD_NO_CART_SELF_TEST_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
[6 empty] | FALSE | UI POST HD communication failure |  | UI |  | HD POST: UI-Proc Comm | Call service to schedule a repair. | 101 | ALARM_ID_UI_POST_HD_COMM |  | ALARM_PRIORITY_HIGH | 0 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
cleaning | PRS 863 | SA 482 |  | SRSDG 856 SRSDG 930 |  | FALSE | DG heat disinfect temperature gradient out of range |  | DG |  | Cleaning: Tpo THD Temp Delta | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | 102 | ALARM_ID_DG_HEAT_DISINFECT_TEMP_GRAD_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 196 | SRSHD 1059 |  |  | FALSE | HD invalid calibration CRC |  | HD |  | HD POST: Cal NVRAM CRC | Call service to schedule a repair. | 103 | ALARM_ID_HD_NVDATAMGMT_CAL_GROUP_RECORD_CRC_INVALID |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 478 | SRSHD 1115 |  |  | FALSE | HD air trap level sensors reporting illegal combination of air/fluid |  | HD |  | HD Fault: Air Trap Sensor | Call service to schedule a repair. | 104 | ALARM_ID_HD_AIR_TRAP_ILLEGAL_LEVELS |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 283 |  | SRSDG 608 |  | FALSE | DG invalid calibration CRC |  | DG |  | DG POST: Cal NVRAM CRC | Call service to schedule a repair. | 105 | ALARM_ID_DG_NVDATAMGMT_CAL_GROUP_RECORD_CRC_INVALID |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 369 |  | SRSDG 820 SRSDG 821 SRSDG 1087 |  | FALSE | DG dialysate flow sensor invalid calibration record |  | DG |  | DG POST: Flow Sensor Cal | Call service to schedule a repair. | 106 | ALARM_ID_DG_FLOW_SENSORS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
No SRS ID - need to add then assess Move to DG machine alarm? (allow properties table to be aligned) Sean - to revisit code to assess | [5 empty] | FALSE | HD reports DG restarted fault |  | HD |  | DG Fault: DG Restarted | Call service to schedule a repair. | 107 | ALARM_ID_HD_DG_RESTARTED_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 494 | SA 276 | SRSHD 1318 |  |  | FALSE | HD syringe pump ADC FPGA fault |  | HD |  | HD Fault: Syringe Pump ADC Com | Call service to schedule a repair. | 108 | ALARM_ID_HD_SYRINGE_PUMP_FPGA_ADC_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 490 | SRSHD 580 |  |  | FALSE | HD syringe pump volume check error |  | HD |  | HD Fault: Syringe Pump Vol | Call service to schedule a repair. | 109 | ALARM_ID_HD_SYRINGE_PUMP_VOLUME_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 274 | SRSHD 579 |  |  | FALSE | HD syringe pump speed check error |  | HD |  | HD Fault: Syringe Pump Speed | Call service to schedule a repair. | 110 | ALARM_ID_HD_SYRINGE_PUMP_SPEED_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 274 | SRSHD 1333 |  |  | FALSE | HD syringe pump not stopped in off state error |  | HD |  | HD Fault: Syringe Pump Speed | Call service to schedule a repair. | 111 | ALARM_ID_HD_SYRINGE_PUMP_NOT_STOPPED_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 312 SA 313 |  | SRSDG 953 |  | FALSE | DG load cells A1/B1 FPGA fault |  | DG |  | DG Fault: LC A2/B2 Comm | Call service to schedule a repair. | 112 | ALARM_ID_DG_LOAD_CELL_A2_B2_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 341 |  | SRSDG 890 |  | FALSE | DG Drain pump current out of range |  | DG |  | DG Fault: DRP Current | Call service to schedule a repair. | 113 | ALARM_ID_DG_DRAIN_PUMP_CURRENT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
move to wet self-test? | PRS 494 | SA 479 | SRSHD 1532 |  |  | FALSE | HD venous air bubble detector self-test failure |  | HD |  | Process: Pre-Tx Wet Self-Test | Self-test has detected a problem with the air bubble sensor\n- Check cartridge installation on to machine\n- If problem persists, terminate pre-treatment and \nrestart with a new cartridge. | 114 | ALARM_ID_HD_VENOUS_BUBBLE_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
Sean - update code to remove temperatures from code that are no longer needed (see updated SRSDG 224) | PRS 863 | SA 321 SA 322 SA 323 SA 324 SA 328 SA 332 SA 336 SA 298 SA 460 |  | SRSDG 224 |  | FALSE | DG temperature sensor out of range |  | DG |  | DG Fault: Temp Sensor Range | Call service to schedule a repair. | 115 | ALARM_ID_DG_TEMPERATURE_SENSOR_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
[6 empty] | FALSE | HD UI SD card failure |  | UI HD |  | HD Alarm (Low): SD-Card Error | SD-Card Failure \n- Power cycle the system.\n- Call service to schedule a repair if problem persists. | 116 | ALARM_ID_HD_UI_SDCARD_FAILURE |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
|  |  |  | SRSDG 909 |  | FALSE | DG CAN message not acked |  | DG |  | Loss of HD Comm (HD ACK) | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | 117 | ALARM_ID_DG_CAN_MESSAGE_NOT_ACKED |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 458 |  | SRSDG 43 |  | FALSE | DG RTC config error |  | DG |  | DG Fault: RTC Config | Call service to schedule a repair. | 118 | ALARM_ID_DG_RTC_CONFIG_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 874 | SA 499 |  | SRSDG 908 |  | FALSE | DG RTC battery low error |  | DG |  | DG Alarm (Low): RTC Bat Low | Call service to schedule a repair. | 119 | ALARM_ID_DG_RTC_BATTERY_LOW |  | ALARM_PRIORITY_LOW | 851 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
|  |  | SRSHD 777 |  |  | FALSE | HD pre-treatment mode wet self-test prime check failure |  | HD |  | Process: Pre-Tx Wet Self-Test | Self-test has detected a bubble in the venous bubble detector\n- The cartridge should be primed\n- Check cartridge installation on to machine\n- If problem persists, terminate pre-treatment and \nrestart with a new cartridge. | 120 | ALARM_ID_HD_PRE_TREATMENT_WET_PRIME_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 863 | SA 363 |  | SRSDG 611 |  | FALSE | DG main primary heater voltage out of range |  | DG |  | DG Fault: Prim Heater Lg Control | Call service to schedule a repair. | 121 | ALARM_ID_DG_MAIN_PRIMARY_HEATER_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 362 |  | SRSDG 612 |  | FALSE | DG small primary heater voltage out of range |  | DG |  | DG Fault: Prim Heater Sm Control | Call service to schedule a repair. | 122 | ALARM_ID_DG_SMALL_PRIMARY_HEATER_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 361 |  | SRSDG 613 |  | FALSE | DG trimmer heater voltage out of range |  | DG |  | DG Fault: Trimmer Heater Control | Call service to schedule a repair. | 123 | ALARM_ID_DG_TRIMMER_HEATER_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 390 |  | SRSHD 954 |  |  | FALSE | HD end of treatment alarm (high priority - user still not acting to end treatment) |  | HD |  | Blood Stop: Blood Clot Risk | Blood pump has been off for too long.\n- Treatment must be terminated. Blood clot risk.\n- NO RINSEBACK ALLOWED. | 124 | ALARM_ID_HD_END_OF_TREATMENT_HIGH |  | ALARM_PRIORITY_HIGH | 11 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 390 |  | SRSHD 948 |  |  | FALSE | HD blood sitting too long after treatment stopped by user alarm (>5 min) |  | HD |  | Blood Stop: Blood Clot Risk | Blood pump has been off for too long.\n- Treatment must be terminated. Blood clot risk.\n- NO RINSEBACK ALLOWED. | 125 | ALARM_ID_HD_TREATMENT_STOPPED_NO_RINSEBACK |  | ALARM_PRIORITY_HIGH | 11 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 387 |  | SRSHD 925 |  |  | TRUE | HD blood leak detected alarm |  | HD |  | Blood Stop: Dialyzer Blood Leak | A blood leak has been detected. Treatment must end. | 126 | ALARM_ID_HD_BLOOD_LEAK_DETECTED |  | ALARM_PRIORITY_HIGH | 150 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 911 |  |  | TRUE | HD venous pressure too low during treatment |  | HD |  | Blood Stop: Venous Pres. Low | Venous pressure is low.\n- Check to be sure your blue venous connections are secure.\n- Check that your venous access is connected and secured \nproperly. | 127 | ALARM_ID_HD_VENOUS_PRESSURE_LOW |  | ALARM_PRIORITY_HIGH | 200 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE
x | PRS 494 | SA 506 | SRSHD 919 |  |  | TRUE | HD venous air bubble detected alarm |  | HD |  | Blood Stop: Venous Air Bubble | Venous air bubble detected indicating failure of air trap \ncontrol system.\n- Treatment must be terminated without option for rinseback.\n- Check access for dislodgement. | 128 | ALARM_ID_HD_VENOUS_BUBBLE_DETECTED |  | ALARM_PRIORITY_HIGH | 12 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE
x | PRS 494 | SA 507 | SRSHD 921 |  |  | FALSE | Available for use | [5 empty] | 129 | ALARM_ID_AVAILABLE_17 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 913 |  |  | TRUE | HD venous pressure too high during treatment |  | HD |  | Blood Stop: Venous Pres. High | Venous pressure is high.\n- Check for access or patient discomfort.\n- Resolve access issue if applicable.\n- Check that there are no kinks or clots in the tubing.\n- Unkink the lines if applicable.\n- End treatment and do not return blood if clots are present. | 130 | ALARM_ID_HD_VENOUS_PRESSURE_HIGH |  | ALARM_PRIORITY_HIGH | 201 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 915 |  |  | TRUE | HD arterial pressure too low during treatment |  | HD |  | Blood Stop: Arterial Pres. Low | Arterial pressure is low.\n- Check that the access is properly connected and secured.\n- Resolve access issue if applicable.\n- Check to be sure the red arterial connections are secure.\n- Check that there are no kinks or clots in the tubing.\n- Unkink lines if applicable.\n- End treatment and do not return blood if clots are present. | 131 | ALARM_ID_HD_ARTERIAL_PRESSURE_LOW |  | ALARM_PRIORITY_HIGH | 202 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 917 |  |  | TRUE | HD arterial pressure too high during treatment |  | HD |  | Blood Stop: Arterial Pres. High | Arterial pressure is high.\n- Check that the access is properly connected and secured.\n- Resolve access issue if applicable.\n- Check to be sure your red arterial connections are secure.\n- Check that there are no kinks or clots in the tubing.\n- Unkink lines if applicable.\n- End treatment and do not return blood if clots are present. | 132 | ALARM_ID_HD_ARTERIAL_PRESSURE_HIGH |  | ALARM_PRIORITY_HIGH | 203 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE
x | PRS 863 | SA 419 |  | SRSDG 328 |  | TRUE | DG fluid leak detected alarm |  | DG |  | DG Fault: Fluid Leak in DG | Fluid leak detected in Dialysate Generation Delivery Device.\n-Treatment must be terminated.\n- Shut off water source and disconnect water and \nelectrical connections from system.\n- Call service to schedule a repair. | 133 | ALARM_ID_DG_FLUID_LEAK_DETECTED |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 409 |  | SRSHD 927 |  |  | TRUE | HD fluid leak detected alarm |  | HD |  | Blood Stop: Fluid Leak in HD | Fluid leak detected in Hemodialysis Delivery device.\n- If in treatment, BLOOD STOPPED.\n-Identify and correct the cause of the spill.\n- Clean up the spill and dry leak sensor.\n- If in treatment, resume treatment. | 134 | ALARM_ID_HD_FLUID_LEAK_DETECTED |  | ALARM_PRIORITY_HIGH | 208 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 411 |  | SRSHD 938 |  |  | FALSE | HD experienced large acceleration |  | HD |  | Blood Stop: HD Shock Detected | Excessive movement detected with Hemodialysis Delivery \ndevice.\n- Confirm the device is stationary and level before continuing. | 135 | ALARM_ID_HD_SHOCK |  | ALARM_PRIORITY_HIGH | 210 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 686 |  |  | SRSDG 327 |  | FALSE | DG experienced large acceleration |  | DG |  | Dial. Gen: Shock-Vibration | Excessive movement detected with Dialysis Generation \ndevice.\n- Treatment is paused.\n- Confirm the device is stationary and level before continuing. | 136 | ALARM_ID_DG_SHOCK |  | ALARM_PRIORITY_MEDIUM | 440 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 412 |  | SRSHD 936 |  |  | FALSE | HD tilt exceeds maximum |  | HD |  | Blood Stop: HD Tilted | Excessive tilt detected with Hemodialysis Delivery device.\n- Confirm the device is stationary and level before continuing. | 137 | ALARM_ID_HD_EXCESSIVE_TILT |  | ALARM_PRIORITY_HIGH | 211 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 687 |  |  | SRSDG 326 |  | FALSE | DG tilt exceeds maximum |  | DG |  | Dial. Gen: Tilt | Excessive tilt detected with Dialysis Generation device.\n- Treatment is paused.\n-Confirm the device is stationary and level before continuing. | 138 | ALARM_ID_DG_EXCESSIVE_TILT |  | ALARM_PRIORITY_MEDIUM | 441 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 823 |  | SRSHD 1593 |  |  | FALSE | HD AC power lost alarm |  | HD |  | Loss of AC Power | Power has been lost. Check all electrical connections.\n- The display and CPU will run off of the battery until \npower returns. | 139 | ALARM_ID_HD_AC_POWER_LOST |  | ALARM_PRIORITY_LOW | 852 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 400 |  | SRSHD 1124 |  |  | FALSE | HD DG communication timeout |  | HD |  | Loss of DG Comm (Timeout) | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | 140 | ALARM_ID_HD_DG_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
|  |  | SRSHD 1169 |  |  | TRUE | HD air trap fill timeout during treatment |  | HD |  | Blood Stop: Air Trap Fill | Failure to fill the Air Trap.\n- Treatment must be terminated without option for rinseback.\n- Check access for dislodgement. | 141 | ALARM_ID_HD_AIR_TRAP_FILL_DURING_TREATMENT |  | ALARM_PRIORITY_HIGH | 13 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 496 |  | SRSHD 1361 SRSHD 1362 |  |  | TRUE | HD blood pump occlusion detected |  | HD |  | Blood Stop: Blood Pump Occlusion | Occlusion detected after the blood pump.\n- Check for clots or kinks around the dialyzer and in the \nvenous line.\n- Resolve access issue if applicable.\n- End treatment and do not return blood if clots are present.\n- Check for access or patient discomfort.\n- Unkink the lines if applicable. | 142 | ALARM_ID_HD_OCCLUSION_BLOOD_PUMP |  | ALARM_PRIORITY_HIGH | 206 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE
x | PRS 863 | SA 338 |  | SRSDG 912 |  | FALSE | DG dialysate temperature sensors out of range |  | DG |  | DG Fault: TRo TDi Delta | Call service to schedule a repair. | 143 | ALARM_ID_DG_DIALYSATE_TEMPERATURE_SENSORS_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 853 |  |  | SRSDG 1066 SRSDG 1054 SRSDG 970 |  | FALSE | DG cleaning mode inlet water conductivity out of high range. |  | DG |  | Cleaning: Inlet Water Cond High | Cleaning Mode Inlet water conductivity too high.\n- Cleaning mode process has terminated \n- Check inlet water before restarting cleaning mode. | 144 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_COND_TOO_HIGH |  | ALARM_PRIORITY_MEDIUM | 603 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 394 |  |  | SRSDG 846 |  | FALSE | DG concentrate conductivity after adding acid out of range alarm |  | DG |  | Dial. Gen: Acid Conc. Cond. | Acid concentrate conductivity during prime is out of range.\n- Check correct acid concentrate is connected and full.\n- Resume to reprime concentrate. | 145 | ALARM_ID_DG_ACID_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 410 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE
x | PRS 863 | SA 285 |  | SRSDG 312 |  | FALSE | DG RTC or timer accuracy failure. |  | DG |  | DG POST: DG RTC Timer | Call service to schedule a repair. | 146 | ALARM_ID_DG_RTC_OR_TIMER_ACCURACY_FAILURE |  | ALARM_PRIORITY_MEDIUM | 421 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 543 |  |  | SRSDG 398 |  | FALSE | DG wait for the DG to produce dialysate |  | DG |  | Dial. Gen: Dialysate not ready | Dialysate is not yet ready.\n- Treatment is paused and will be ready to resume when \ndialysate is ready for use. | 147 | ALARM_ID_DG_CREATING_DIALYSATE_PLEASE_WAIT |  | ALARM_PRIORITY_MEDIUM | 421 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 557 |  |  | SRSDG 1065 SRSDG 1053 SRSDG 969 |  | FALSE | DG Cleaning mode inlet water temperature too high |  | DG |  | Cleaning: Inlet Water Temp High | Cleaning Mode Inlet water temperature too high.\n- Cleaning mode process has terminated.\n- Check inlet water before restarting cleaning mode. | 148 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_TEMP_TOO_HIGH |  | ALARM_PRIORITY_MEDIUM | 601 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 475 | SRSHD 646 |  |  | FALSE | HD watchdog expired |  | HD |  | HD Fault: WD Error | Call service to schedule a repair. | 149 | ALARM_ID_HD_WATCHDOG_EXPIRED |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 404 |  |  | SRSDG 391 |  | TRUE | DG inlet water conductivity in high range |  | DG |  | Dial. Gen: Inlet Water Cond High | Inlet water conductivity is high.\n- Treatment is paused.\n- Check the incoming water supply. | 150 | ALARM_ID_DG_INLET_WATER_CONDUCTIVITY_IN_HIGH_RANGE |  | ALARM_PRIORITY_MEDIUM | 434 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 403 |  |  | SRSDG 826 |  | TRUE | DG inlet water conductivity in low range |  | DG |  | Dial. Gen: Inlet Water Cond Low | Inlet water conductivity is low.\n- Treatment is paused.\n- Check the incoming water supply. | 151 | ALARM_ID_DG_INLET_WATER_CONDUCTIVITY_IN_LOW_RANGE |  | ALARM_PRIORITY_MEDIUM | 435 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 401 |  |  | SRSDG 417 |  | TRUE | DG inlet water pressure in low range |  | DG |  | Dial. Gen: Inlet Water Pres Low | Inlet water pressure is low.\n- Treatment is paused.\n- Check the incoming water supply. | 152 | ALARM_ID_DG_INLET_WATER_PRESSURE_IN_LOW_RANGE |  | ALARM_PRIORITY_MEDIUM | 431 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 529 |  | SRSHD 1122 |  |  | FALSE | HD prime completed high priority alarm |  | HD |  | Process: Pre-Tx Terminated | Pre-Treatment timeout.\n- Patient was not connected to device\n- Pre-Treatment has been terminated\n- Proceed to end treatment to remove cartridge from system. | 153 | ALARM_ID_HD_PRIME_COMPLETED_HIGH |  | ALARM_PRIORITY_HIGH | 292 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 854 |  |  | SRSDG 1067 SRSDG 1055 SRSDG 971 |  | FALSE | DG cleaning mode inlet water conductivity out of low range. |  | DG |  | Cleaning: Inlet Water Cond Low | Cleaning Mode Inlet water conductivity too low.\n- Cleaning mode process has terminated \n- Check inlet water before restarting cleaning mode. | 154 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_COND_TOO_LOW |  | ALARM_PRIORITY_MEDIUM | 604 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 400 |  | SRSHD 1574 |  |  | FALSE | HD No reservoirs data message receive at least once every 2 seconds |  | HD |  | Loss of DG Comm (Reservoir) | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | 155 | ALARM_ID_HD_NEW_RESERVOIRS_DATA_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 400 |  | SRSHD 1575 |  |  | FALSE | HD No DG operation mode message receive at least once every 2 seconds |  | HD |  | Loss of DG Comm (Op Mode) | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. | 156 | ALARM_ID_HD_DG_NEW_OPERATION_MODE_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
cleaning |  |  |  | SRSDG 988 |  | FALSE | DG chemical disinfect prime acid line timeout. |  | DG |  | Cleaning: Chem Disinf PAA Prime | Unable to sense proper chemical disinfectant.\n- Cleaning mode process has terminated.\n- Check that proper chemical disinfectant is connected to the \nbicarb port.\n- Check chemical disinfectant bottle is not empty. | 157 | ALARM_ID_DG_CHEM_DISINFECT_PRIME_ACID_LINE_TIME_OUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 405 |  |  | SRSDG 408 |  | TRUE | DG inlet water temperature is in the low range |  | DG |  | Dial. Gen: Inlet Water Temp Low | Inlet water temperature is low.\n- Treatment is paused.\n- Check the incoming water supply. | 158 | ALARM_ID_DG_INLET_WATER_TEMPERATURE_IN_LOW_RANGE |  | ALARM_PRIORITY_MEDIUM | 433 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
cleaning |  |  |  | SRSDG 1044 |  | FALSE | DG chemical disinfect flush remove acid and close the concentrate cap. |  | DG |  | Cleaning: Chem Disinf Connect Cap | Remove chemical disinfectant connection and close cap. | 159 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_REMOVE_ACID |  | ALARM_PRIORITY_LOW | 751 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE
x | PRS 494 | SA 239 | SRSHD 972 |  |  | FALSE | HD blood pump rotor speed too high |  | HD |  | HD Fault: BP Rotor Max Speed | Call service to schedule a repair. | 160 | ALARM_ID_HD_BLOOD_PUMP_ROTOR_SPEED_TOO_HIGH |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 497 |  | SRSDG 122 |  | FALSE | DG comm too many bad CRCs |  | DG |  | DG Fault: HD-DG Comm | Call service to schedule a repair. | 161 | ALARM_ID_DG_COMM_TOO_MANY_BAD_CRCS |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 300 |  | SRSDG 867 |  | FALSE | DG FPGA clock speed check failure. |  | DG |  | DG Fault: Proc-FPGA Clock | Call service to schedule a repair. | 162 | ALARM_ID_DG_FPGA_CLOCK_SPEED_CHECK_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
revisit - determine if this is blood stopped or dialysate may need to add to PRS |  |  | SRSHD 1260 |  |  | FALSE | HD sees primary load cell for reservoir 1 change too much too fast |  | HD |  | Dial. Gen: UF Rate Check | Ultrafiltration control detected and incorrect rate.\n- This can be caused by excessive movement of system\n(shock/vibration)\n- Check system is stable and level\n- Resume and If problem persists, call service to schedule a\nrepair. | 163 | ALARM_ID_HD_LOAD_CELL_ACCELERATION_RES_1_ALARM |  | ALARM_PRIORITY_MEDIUM | 442 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
revisit - determine if this is blood stopped or dialysate may need to add to PRS |  |  | SRSHD 1601 |  |  | FALSE | HD sees primary load cell for reservoir 2 change too much too fast |  | HD |  | Dial. Gen: UF Rate Check | Ultrafiltration control detected and incorrect rate.\n- This can be caused by excessive movement of system \n(shock/vibration).\n- Check system is stable and level\n- Resume and If problem persists, call service to schedule a \nrepair. | 164 | ALARM_ID_HD_LOAD_CELL_ACCELERATION_RES_2_ALARM |  | ALARM_PRIORITY_MEDIUM | 442 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
sean remove from code 15min limint (SRS 1556 has been updated) |  |  | SRSHD 1228 SRSHD 1399 SRSHD 1556 SRSHD 1709 |  |  | FALSE | HD in treatment rinseback operation timeout |  | HD |  | Blood Stop: Blood Clot Risk | Blood pump has been off for too long.\n- Treament must be terminated. Blood clot risk.\n- NO RINSEBACK ALLOWED. | 165 | ALARM_ID_HD_TREATMENT_RINSEBACK_TIMEOUT_ALARM |  | ALARM_PRIORITY_HIGH | 11 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
[6 empty] | FALSE | Available for use | [5 empty] | 166 | ALARM_ID_AVAILABLE_3 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE
needs to be high priority due to blood stopped | [5 empty] | FALSE | HD cartridge door opened alarm |  | HD |  | Blood Stop: Door Open | Device door is open.\n - Close door and press RESUME button to continue. | 167 | ALARM_ID_HD_CARTRIDGE_DOOR_OPENED |  | ALARM_PRIORITY_HIGH | 209 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
[6 empty] | FALSE | Available for use | [5 empty] | 168 | ALARM_ID__AVAILABLE_70 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE
x | PRS 863 | SA 493 |  | SRSDG 643 |  | FALSE | DG Dialysate flow rate out of maximum range |  | DG |  | DG Fault: Dialysate Flow Sens Rg | Call service to schedule a repair. | 169 | ALARM_ID_DG_DIALYSATE_FLOW_RATE_OUT_OF_MAX_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
Currently in CBIT section - should be somewhere else? | [5 empty] | FALSE | HD syringe empty alarm |  | HD |  | Heparin Syringe Empty | Heparin syringe empty.\n- Treatment will continue\n- If additional heparin is required to complete treatment, \nadminister manually. | 170 | ALARM_ID_HD_SYRINGE_PUMP_SYRINGE_EMPTY |  | ALARM_PRIORITY_LOW | 730 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE
|  |  | SRSHD 1525 |  |  | FALSE | HD syringe pump occlusion alarm |  | HD |  | Process: Pre-Tx Hep. Pump Occ. | Heparin pump occlusion detected.\n- Check heparin line is unclamped.\n- Check there are no kinks in the heparin line. | 171 | ALARM_ID_HD_SYRINGE_PUMP_OCCLUSION |  | ALARM_PRIORITY_LOW | 716 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
|  |  | SRSHD 1383 |  |  | FALSE | HD syringe pump not enough Heparin alarm |  | HD |  | Process: Pre-Tx Not Enough Hep. | Not enough heparin is available for treatment.\n- Check if syringe pump is touching syringe plunger.\n- If syringe pump is touching syringe plunger, end treatment.\n- If syringe pump is not touching syringe plunger, reload \nsyringe with enough heparin. | 172 | ALARM_ID_HD_SYRINGE_PUMP_NOT_ENOUGH_HEPARIN_ALARM |  | ALARM_PRIORITY_LOW | 715 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 494 | SA 474 | SRSHD 659 |  |  | FALSE | HD RTC configuration error |  | HD |  | HD Fault: RTC Config | Call service to schedule a repair. | 173 | ALARM_ID_HD_RTC_CONFIG_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 198 | SRSHD 630 |  |  | FALSE | HD RTC or timer accuracy failure |  | HD |  | HD POST: HD RTC Timer | Call service to schedule a repair. | 174 | ALARM_ID_HD_RTC_OR_TIMER_ACCURACY_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 245 SA 254 SA 262 | SRSHD 1502 SRSHD 1503 SRSHD 1568 |  |  | FALSE | HD pump direction status error |  | HD |  | HD Fault: Pump Commutation Err | Call service to schedule a repair. | 175 | ALARM_ID_HD_PUMP_DIRECTION_STATUS_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
Dialysate Temp |  |  | SRSHD 895 |  |  | FALSE | HD dialysate temperature above high safety range | [5 empty] | 176 | ALARM_ID_AVAILABLE_4 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 863 | SA 417 |  | SRSDG 919 |  | FALSE | DG software fault. Software found itself in an unexpected state |  | DG |  | DG Fault: Software Invalid State | Call service to schedule a repair. | 177 | ALARM_ID_DG_SOFTWARE_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 494 | SA 225 |  |  | SRSUI 1066 | FALSE | HD communication timeout |  | UI |  | HD Fault: UI-Proc Comm | Call service to schedule a repair. | 178 | ALARM_ID_HD_COMM_TIMEOUT | OK | ALARM_PRIORITY_HIGH | 0 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
|  |  |  | SRSDG 605 |  | FALSE | DG FPGA communication down for too long |  | DG |  | DG Fault: FPGA Com | Call service to schedule a repair. | 179 | ALARM_ID_DG_FPGA_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 494 |  | SRSDG 883 |  | FALSE | DG RO flow out of maximum range |  | DG |  | DG Fault: RO Flow Sensor Range | Call service to schedule a repair. | 180 | ALARM_ID_DG_RO_FLOW_RATE_OUT_OF_MAX_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
|  |  |  | SRSDG 849 SRSDG 838 |  | FALSE | DG load cells weight out of range for tare |  | DG |  | DG Fault: Tare Out of Range | Call service to schedule a repair. | 181 | ALARM_ID_DG_LOAD_CELLS_TARE_WEIGHT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 369 |  | SRSDG 162 |  | FALSE | DG load cells invalid calibration |  | DG |  | DG POST: LC Cal | Call service to schedule a repair. | 182 | ALARM_ID_DG_LOAD_CELLS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
cleaning |  |  |  | SRSDG 851 |  | FALSE | DG invalid load cell value |  | DG |  | Cleaning: Invalid Resv Transfer | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | 183 | ALARM_ID_DG_INVALID_LOAD_CELL_VALUE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 355 SA 356 |  | SRSDG 559 |  | FALSE | DG inlet UV reactor not healthy |  | DG |  | DG Fault: UVi Error | Call service to schedule a repair. | 184 | ALARM_ID_DG_INLET_UV_REACTOR_NOT_HEALTHY |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 874 | SA 299 |  | SRSDG 628 |  | FALSE | DG fan RPM out of range |  | DG |  | DG Alarm (Low): Fan Speed | Fan failure.\n- Check that the fans inlets are not blocked or obstructed. | 185 | ALARM_ID_DG_FAN_RPM_OUT_OF_RANGE |  | ALARM_PRIORITY_LOW | 851 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 863 | SA 351 SA 353 |  | SRSDG 891 |  | FALSE | DG concentrate pump fault |  | DG |  | DG Fault: CP1/CP2 Fault | Call service to schedule a repair. | 186 | ALARM_ID_DG_CONCENTRATE_PUMP_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 352 |  | SRSDG 905 |  | FALSE | DG concentrate pump CP1 speed control error |  | DG |  | DG Fault: CP1 Speed Fault | Call service to schedule a repair. | 187 | ALARM_ID_DG_CP1_SPEED_CONTROL_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 354 |  | SRSDG 906 |  | FALSE | DG concentrate pump CP2 speed control error |  | DG |  | DG Fault: CP2 Speed Fault | Call service to schedule a repair. | 188 | ALARM_ID_DG_CP2_SPEED_CONTROL_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 339 SA 343 |  | SRSDG 448 |  | FALSE | DG drain pump RPM out of range |  | DG |  | DG Fault: DRP Speed Fault | Call service to schedule a repair. | 189 | ALARM_ID_DG_DRAIN_PUMP_RPM_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
[6 empty] | FALSE | DG drain pump off fault |  | DG |  | DG Fault: DRP Speed Not Off | Call service to schedule a repair. | 190 | ALARM_ID_DG_DRAIN_PUMP_OFF_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 498 |  | SRSDG 885 |  | FALSE | DG flow rate out of upper range |  | DG |  | DG Fault: RO Flow High | Call service to schedule a repair. | 191 | ALARM_ID_DG_FLOW_RATE_OUT_OF_UPPER_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 498 |  | SRSDG 945 |  | FALSE | DG flow rate out of lower range |  | DG |  | DG Fault: RO Flow Low | Call service to schedule a repair. | 192 | ALARM_ID_DG_FLOW_RATE_OUT_OF_LOWER_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
Dialysate Temp |  |  | SRSHD 897 |  |  | FALSE | HD dialysate temperature out of low safety range. | [5 empty] | 193 | ALARM_ID_AVAILABLE_5 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 863 | SA 360 |  | SRSDG 881 |  | FALSE | DG RO pump duty cycle out of range |  | DG |  | DG Fault: RO Driver Feedback | Call service to schedule a repair. | 194 | ALARM_ID_DG_RO_PUMP_DUTY_CYCLE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 487 |  | SRSDG 925 |  | FALSE | DG RO pump pressure out of range |  | DG |  | DG Fault: RO Pump High Pressure | Call service to schedule a repair. | 195 | ALARM_ID_DG_RO_PUMP_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 348 |  | SRSDG 955 |  | FALSE | DG CPi/CPo sensors FPGA fault |  | DG |  | DG Fault: CPi/CPo Comm | Call service to schedule a repair. | 196 | ALARM_ID_DG_CPI_CPO_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 347 |  | SRSDG 956 |  | FALSE | DG CD1/CD2 sensors FPGA fault |  | DG |  | DG Fault: CD1/CD2 Comm | Call service to schedule a repair. | 197 | ALARM_ID_DG_CD1_CD2_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 367 |  | SRSDG 872 |  | FALSE | DG RO flow too low while primary heater is on |  | DG |  | DG Fault: HP On With No Flow | Call service to schedule a repair. | 198 | ALARM_ID_DG_RO_FLOW_TOO_LOW_WHILE_PRIMARY_HEATER_IS_ON |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 366 |  | SRSDG 871 |  | FALSE | DG Dialysate flow too low while trimmer heater is on |  | DG |  | DG Fault: HR On With No Flow | Call service to schedule a repair. | 199 | ALARM_ID_DG_DIALYSATE_FLOW_TOO_LOW_WHILE_TRIMMER_HEATER_IS_ON |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 293 SA 358 SA 359 |  | SRSDG 178 |  | FALSE | DG thermistors/sensors temperature out of range |  | DG |  | DG Fault: Thermistor Range | Call service to schedule a repair. | 200 | ALARM_ID_DG_THERMISTORS_TEMPERATURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
|  |  | SRSHD 746 SRSHD 747 |  |  | FALSE | HD pre-treatment mode wet self-test lc vs FMD failure |  | HD |  | Process: Pre-Tx Wet Self-Test | Self-test has detected a problem in the dialysate flow \nmeasurement or there is a leak in the dialysate in/out lines.\n- The cartridge should be primed\n- Check cartridge installation on to machine\n- If problem persists, terminate pre-treatment and restart \nwith a new cartridge. | 201 | ALARM_ID_HD_PRE_TREATMENT_WET_FLOW_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
[6 empty] | FALSE | HD pre-treatment mode dry pressure normal self-test failure |  | HD |  | Process: Pre-Tx Dry Self-Test | Self-test has detected blood circuit pressure cannot be \nrelieved.\n- Check PBo connection to HD for kinks\n -Check PBo transducer protector for occlusion\n- If problem persists, install a new cartridge and \nPBo transducer protector. | 202 | ALARM_ID_HD_PRE_TREATMENT_DRY_PRESSURE_NORMAL_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
[6 empty] | FALSE | Available for use | [5 empty] | 203 | ALARM_ID_AVAILABLE_6 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 373 |  | SRSDG 429 |  | FALSE | DG RO rejection ratio out of range |  | DG |  | DG Fault: RO Rejection | Call service to schedule a repair. | 204 | ALARM_ID_DG_RO_REJECTION_RATIO_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
|  |  |  | SRSDG 913 |  | FALSE | DG conductivity sensor fault |  | DG |  | DG Fault: Cond Sensors Error | Call service to schedule a repair. | 205 | ALARM_ID_DG_CONDUCTIVITY_SENSOR_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
|  |  |  | SRSDG 437 |  | FALSE | DG dialysate fill runs out of time |  | DG |  | DG Fault: Reservoir Fill Timeout | Call service to schedule a repair. | 206 | ALARM_ID_DG_DIALYSATE_FILL_OUT_OF_TIME |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
|  |  |  | SRSDG 240 |  | FALSE | DG flow meter check failure alarm |  | DG |  | DG Fault: FMP Flow Rate Error | Call service to schedule a repair. | 207 | ALARM_ID_DG_FLOW_METER_CHECK_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
blaine to revisit redundant with venous pressure alarm limit? | [5 empty] | TRUE | HD venous pressure occlusion alarm |  | HD |  | Blood Stop: Venous Pres. High | Venous pressure is high.\n- Check for access or patient discomfort.\n- Resolve access issue if applicable.\n- Check that there are no kinks or clots in the tubing.\n- Unkink the lines if applicable.\n- End treatment and do not return blood if clots are present. | 208 | ALARM_ID_HD_VENOUS_LINE_OCCLUSION |  | ALARM_PRIORITY_HIGH | 201 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE
|  |  |  | SRSDG 980 |  | FALSE | DG drain circulation line timeout |  | DG |  | Cleaning: Flush Circ Error | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | 209 | ALARM_ID_DG_DRAIN_CIRCULATION_LINE_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 872 | SA 229 | SRSHD 1397 |  |  | FALSE | HD battery pack detected an error |  | HD |  | HD Alarm (Low): Battery Fault | Call service to schedule a repair. | 210 | ALARM_ID_HD_BATTERY_PACK_ERROR_DETECTED |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
need to check with sean how this will escalate with other alarm active that will affect rank | PRS 389 |  | SRSHD 946 |  |  | FALSE | HD blood sitting too long warning (>4 min) |  | HD |  | Blood Stopped: Warning | Blood pump is stopped.\n- Resume treatment before countdown timer expires. | 211 | ALARM_ID_HD_BLOOD_SITTING_WARNING |  | ALARM_PRIORITY_HIGH | 290 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
need to check with sean how this will escalate with other alarm active that will affect rank | PRS 389 |  | SRSHD 952 |  |  | FALSE | HD end of treatment alarm (user not acting to end treatment) |  | HD |  | Blood Stopped: Warning | Blood pump is stopped.\n- Resume treatment before countdown timer expires. | 212 | ALARM_ID_HD_END_OF_TREATMENT_ALARM |  | ALARM_PRIORITY_HIGH | 290 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
sean to advise on properties | PRS 530 |  | SRSHD 1120 |  |  | FALSE | HD prime completed medium priority alarm |  | HD |  | Process: Pre-Tx Prime Complete | Prime of the cartridge is complete.\n- Proceed to connect to patient. | 213 | ALARM_ID_HD_PRIME_COMPLETED_MEDIUM |  | ALARM_PRIORITY_MEDIUM | 590 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE
|  |  | SRSHD 909 |  |  | FALSE | Available for use | [5 empty] | 214 | ALARM_ID_AVAILABLE_30 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
|  |  | SRSHD 826 |  |  | FALSE | Available for use | [5 empty] | 215 | ALARM_ID_AVAILABLE_31 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 393 |  | SRSHD 1405 |  |  | FALSE | HD end treatment sub-mode timeout alarm |  | HD |  | Process: End of Dialysis Tx | Dialysis Treatment has completed\n- Proceed to end treatment and rinseback. | 216 | ALARM_ID_HD_END_TREATMENT_TIMEOUT_ALARM |  | ALARM_PRIORITY_HIGH | 291 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
|  |  | SRSHD 1454 |  |  | FALSE | Available for use | [5 empty] | 217 | ALARM_ID_AVAILABLE_32 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
|  |  | SRSHD 1294 |  |  | FALSE | HD syringe detected alarm |  | HD |  | Process: Remove Syringe | Syringe not removed from machine.\n- Remove heparin syringe. | 218 | ALARM_ID_HD_SYRINGE_DETECTED |  | ALARM_PRIORITY_LOW | 745 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
Currently in CBIT section - should be somewhere else? | [5 empty] | FALSE | HD syringe pump syringe removed alarm |  | HD |  | Blood Stop: Syringe Not Detected | Heparin is needed for treatment but syringe has been removed \nfrom syringe pump.\n- Install Heparin syringe. | 219 | ALARM_ID_HD_SYRINGE_PUMP_SYRINGE_REMOVED |  | ALARM_PRIORITY_HIGH | 207 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE
x | PRS 863 | SA 334 SA 335 |  | SRSDG 947 |  | FALSE | DG THd sensors FPGA fault |  | DG |  | DG Fault: THD RTD Comm | Call service to schedule a repair. | 220 | ALARM_ID_DG_THD_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
|  |  | SRSHD 1306 SRSHD 1307 SRSHD 1308 |  |  | FALSE | HD Saline bag is empty |  | HD |  | Blood Stop: Saline Bag Empty | Saline bag appears empty.\n- Check saline bag and line and replace if saline bag is \nempty. | 221 | ALARM_ID_HD_EMPTY_SALINE_BAG |  | ALARM_PRIORITY_HIGH | 204 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE
x | PRS 494 | SA 230 | SRSHD 1585 |  |  | FALSE | HD occlusion sensor FPGA fault |  | HD |  | HD Fault: BP Occ Com | Call service to schedule a repair. | 222 | ALARM_ID_HD_OCCLUSION_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 236 | SRSHD 1583 |  |  | FALSE | HD arterial sensor FPGA fault |  | HD |  | HD Fault: Art. Pres. Com | Call service to schedule a repair. | 223 | ALARM_ID_HD_ARTERIAL_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 388 |  | SRSHD 944 |  |  | FALSE | HD treatment stopped by user action - pressed stop button |  | HD |  | Blood Stopped: Treatment Pause | Treatment paused by user. | 224 | ALARM_ID_HD_TREATMENT_STOPPED_BY_USER |  | ALARM_PRIORITY_MEDIUM | 300 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
revisit priority - should be low | PRS 391 |  | SRSHD 810 |  |  | FALSE | HD end of treatment warning |  | HD |  | Process: End of Treatment | Dialysis Treatment has completed\n- Proceed to end treatment and rinseback. | 225 | ALARM_ID_HD_END_OF_TREATMENT_WARNING |  | ALARM_PRIORITY_LOW | 739 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
sean to advise on properties | PRS 531 |  | SRSHD 1118 |  |  | FALSE | HD prime completed low priority alarm |  | HD |  | Process: Pre-Tx Priming Complete | Prime of the cartridge is complete.\n- Proceed to connect to patient. | 226 | ALARM_ID_HD_PRIME_COMPLETED_LOW_PRIORITY |  | ALARM_PRIORITY_LOW | 729 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE
[6 empty] | FALSE | Available for use | [5 empty] | 227 | ALARM_ID_AVAILABLE_15 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
|  |  | SRSHD 1571 |  |  | FALSE | HD time out on prime saline purge air state. |  | HD |  | Process: Pre-Tx Prime Timeout | Blood line prime with saline timeout.\n- Check blood and dialysate lines for clamps, kinks or leaks. | 228 | ALARM_ID_HD_PRIME_SALINE_PURGE_AIR_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
|  |  | SRSHD 1189 |  |  | FALSE | HD prime dialysate dialyzer time out alarm |  | HD |  | Process: Pre-Tx Prime Timeout | Dialyzer prime timeout.\n- Check blood and dialysate lines for clamps, kinks or leaks. | 229 | ALARM_ID_HD_PRIME_DIALYSATE_DIALYZER_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
|  |  | SRSHD 1190 |  |  | FALSE | HD prime dialysate bypass time out alarm |  | HD |  | Process: Pre-Tx Prime Timeout | Dialyzer prime timeout.\n- Check blood and dialysate lines for clamps, kinks or leaks. | 230 | ALARM_ID_HD_PRIME_DIALYSATE_BYPASS_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
|  |  | SRSHD 771 SRSHD 1599 SRSHD 1579 SRSHD 1580 SRSHD 772 |  |  | FALSE | HD pre-treatment mode dry pressure self-test failure |  | HD |  | Process: Pre-Tx Dry Self-Test | Self-test has detected a leak in the circuit\n- Check blood line connections\n- Check dialysate line connections\n- Check PBo and PBa transducer protectors for occlusion.\n- Check clamps.\n- If problem persists, install a new cartridge. | 231 | ALARM_ID_HD_PRE_TREATMENT_DRY_PRESSURE_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
|  |  | SRSHD 1534 SRSHD 1535 |  |  | FALSE | HD pre-treatment mode wet self-test LC vs. LC failure |  | HD |  | Process: Pre-Tx Wet Self-Test | Self-test has detected a problem in the load cell \nmeasurement or there is a leak in the dialysate in/out lines\n- The cartridge should be primed\n- Check cartridge installation onto machine\n- If problem persists, terminate pre-treatment and restart \nwith a new cartridge. | 232 | ALARM_ID_HD_PRE_TREATMENT_WET_LC_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 863 | SA 461 |  | SRSDG 964 SRSDG 949 SRSDG 1089 |  | FALSE | DG barometric sensor coefficients bad CRC |  | DG |  | DG Fault: Baro Cal CRC | Call service to schedule a repair. | 233 | ALARM_ID_DG_BAROMETRIC_SENSOR_COEFFS_BAD_CRC |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 841 |  |  | SRSDG 1037 |  | TRUE | DG inlet water pressure in high range. |  | DG |  | Dial. Gen: Inlet Water Pres High | Inlet water pressure is high.\n- Treatment is paused\n- Check the incoming water supply. | 234 | ALARM_ID_DG_INLET_WATER_PRESSURE_IN_HIGH_RANGE |  | ALARM_PRIORITY_MEDIUM | 430 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
[6 empty] | FALSE | HD in treatment stopped sub-mode after rinseback completed (no escalation) |  | HD |  | Process: Rinseback Complete | Rinseback has been completed. | 235 | ALARM_ID_HD_TREATMENT_STOPPED_AFTER_RINSEBACK |  | ALARM_PRIORITY_LOW | 738 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE
|  |  | SRSHD 1205 |  |  | FALSE | HD needs new cartridge to be installed |  | HD |  | Process: Pre-Tx Dry Self-Test | Self-test has detected fluid in the cartridge.\n- Cartridges are single-use only\n- Install new cartridge to resume. | 236 | ALARM_ID_HD_INSTALL_NEW_CARTRIDGE |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
|  |  | SRSHD 1188 |  |  | FALSE | HD prime saline dialyzer time out alarm |  | HD |  | Process: Pre-Tx Prime Timeout | Dialyzer prime timeout.\n- Check blood and dialysate lines for clamps, kinks or leaks. | 237 | ALARM_ID_HD_PRIME_SALINE_DIALYZER_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
|  |  | SRSHD 1359 |  |  | FALSE | HD no cartridge loaded or installed improperly alarm |  | HD |  | Process: Pre-Tx Dry Self-Test | Self-test has detected cartridge is not installed.\n- Check that a cartridge is properly installed. | 238 | ALARM_ID_HD_NO_CARTRIDGE_LOADED |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
|  |  | SRSHD 1542 |  |  | FALSE | HD fail to remove cartridge alarm |  | HD |  | Process: Post-Tx Remove Cartridge | Cartridge not unloaded, or unloaded improperly.\n- Check that the cartridge has been removed. | 239 | ALARM_ID_HD_CARTRIDGE_REMOVAL_FAILURE |  | ALARM_PRIORITY_LOW | 746 | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 542 |  |  | SRSDG 482 |  | TRUE | DG bicarb conductivity out of range during bicarb pump check alarm |  | DG |  | Dial. Gen: Bicarb Conc. Cond. | Bicarbonate conductivity during prime is out of range.\n- Check correct concentrate is connected and full.\n- Resume to reprime concentrate. | 240 | ALARM_ID_DG_BICARB_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 411 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE
x | PRS 873 |  |  | SRSDG 848 SRSDG 90 SRSDG 561 SRSDG 577 SRSDG 590 SRSDG 591 SRSDG 336 SRSDG 879 SRSDG 340 SRSDG 496 SRSDG 501 SRSDG 505 SRSDG 532 SRSDG 535 SRSDG 556 SRSDG 557 SRSDG 1048 SRSDG 1049 SRSDG 651 SRSDG 655 SRSDG 744 SRSDG 745 SRSDG 1004 SRSDG 1008 SRSDG 1019 SRSDG 1023 SRSDG 1033 SRSDG 1034 |  | FALSE | DG reservoir drain time out |  | DG |  | Cleaning: Resv Drain Timeout | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | 241 | ALARM_ID_DG_RESERVOIR_DRAIN_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
cleaning | PRS 873 |  |  | SRSDG 850 SRSDG 573 SRSDG 574 SRSDG 579 SRSDG 580 SRSDG 492 SRSDG 493 SRSDG 498 SRSDG 499 SRSDG 508 SRSDG 509 SRSDG 519 SRSDG 690 SRSDG 691 SRSDG 703 SRSDG 711 SRSDG 1020 SRSDG 1024 |  | FALSE | DG reservoir fill time out |  | DG |  | Cleaning: Resv Fill Timeout | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | 242 | ALARM_ID_DG_RESERVOIR_FILL_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
cleaning |  |  |  | SRSDG 853 |  | FALSE | DG reservoir leak time out |  | DG |  | Cleaning: Fluid Loss Detected | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | 243 | ALARM_ID_DG_RESERVOIR_LEAK_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
cleaning |  |  |  | SRSDG 852 |  | FALSE | DG cleaning mode temperature sensors difference out of range |  | DG |  | Cleaning: Temp Sensor Check | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | 244 | ALARM_ID_DG_CLEANING_MODE_TEMP_SENSORS_DIFF_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
cleaning |  |  |  | SRSDG 854 |  | FALSE | DG heat disinfect target temperature time out (could not reach to temperature) |  | DG |  | Cleaning: Heat Disinfect Temp TO | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | 245 | ALARM_ID_DG_HEAT_DISINFECT_TARGET_TEMP_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
cleaning | [5 empty] | FALSE | DG cleaning mode conductivity sensors out of range |  | DG |  | Cleaning: RO Water Cond Error | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | 246 | ALARM_ID_DG_CLEANING_MODE_COND_SENSORS_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
cleaning | [5 empty] | FALSE | Alarm Id available | [5 empty] | 247 | ALARM_ID_AVAILABLE_7 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
cleaning | [5 empty] | FALSE | DG chemical disinfect could not reach to target temperature out of range |  | DG |  | Cleaning: Chem Disinfect Proc | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. | 248 | ALARM_ID_DG_CHEM_DISINFECT_TARGET_TEMP_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
cleaning |  |  |  | SRSDG 992 |  | FALSE | DG chemical disinfect target conductivity out of range |  | DG |  | Cleaning: Chem Disinf PAA Cond | Unable to sense proper chemical disinfectant.\n- Cleaning mode process has terminated.\n- Check that proper chemical disinfectant is connected to the \nbicarb port.\n- Check chemical disinfectant bottle is not empty. | 249 | ALARM_ID_DG_CHEM_DISINFECT_TARGET_COND_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
cleaning |  |  |  | SRSDG 1046 |  | FALSE | DG chemical disinfect insert acid and remove the concentrate cap |  | DG |  | Cleaning: Chem Disinf Connect PAA | Connect chemical disinfectant bottle to bicarb port. | 250 | ALARM_ID_DG_CHEM_DISINFECT_INSERT_ACID |  | ALARM_PRIORITY_LOW | 750 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE
x | PRS 398 | SA 196 | SRSHD 1058 |  |  | FALSE | HD invalid system record |  | HD |  | HD POST: Sys Rec NVRAM CRC | Call service to schedule a repair. | 251 | ALARM_ID_HD_INVALID_SYSTEM_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 196 | SRSHD 1062 |  |  | FALSE | HD invalid service record |  | HD |  | HD POST: Serv Rec NVRAM CRC | Call service to schedule a repair. | 252 | ALARM_ID_HD_INVALID_SERVICE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 283 |  | SRSDG 609 |  | FALSE | DG invalid system record |  | DG |  | DG POST: Sys Rec NVRAM CRC | Call service to schedule a repair. | 253 | ALARM_ID_DG_INVALID_SYSTEM_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 283 |  | SRSDG 610 |  | FALSE | DG invalid service record |  | DG |  | DG POST: Serv Rec NVRAM CRC | Call service to schedule a repair. | 254 | ALARM_ID_DG_INVALID_SERVICE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 398 | SA 206 | SRSHD 1358 |  |  | FALSE | HD and UI software builds are not compatible |  | HD |  | HD POST: UI Compatibility | Call service to schedule a repair. | 255 | ALARM_ID_HD_UI_COMPATIBILITY_ERROR |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 205 |  |  | SRSUI 497 | FALSE | HD processor is in Disinfect Chemical Flush mode |  | HD |  | Cleaning: Chem Flush In Progress | Chemical disinfect flush in progress\n- Please wait for the process to complete\n- Treatment will not be able to be started until process \ncompletes.\n- Process may be terminated by pressing the HD stop button. | 256 | ALARM_ID_HD_DISINFECT_CHEM_FLUSH | OK | ALARM_PRIORITY_LOW | 843 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
[6 empty] | FALSE | Alarm ID available for use | [5 empty] | 257 | ALARM_ID_AVAILABLE_8 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
[6 empty] | FALSE | Alarm Id available | [5 empty] | 258 | ALARM_ID_AVAILABLE_9 |  | ALARM_PRIORITY_LOW | 999 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 220 SA 215 SA 216 SA 235 SA 238 | SRSHD 1390 SRSHD 1388 SRSHD 1387 SRSHD 1497 SRSHD 1499 SRSHD 1682 |  |  | FALSE | HD temperatures out of range |  | HD |  | HD Fault: Temp Sensors Error | Call service to schedule a repair. | 259 | ALARM_ID_HD_TEMPERATURES_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 200 | SRSHD 977 |  | SRSUI 1067 | FALSE | HD UI POST Application Integrity (Sha256Sum) failure |  | UI HD |  | HD POST: UI Image | Call service to schedule a repair. | 260 | ALARM_ID_HD_UI_POST_FAILURE_SHASUM | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 201 | SRSHD 1655 |  | SRSUI 1068 | FALSE | HD UI POST CANBus failure |  | UI HD |  | HD POST: UI CAN Driver | Call service to schedule a repair. | 261 | ALARM_ID_HD_UI_POST_FAILURE_CANBUS | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 201 | SRSHD 1656 |  | SRSUI 1069 | FALSE | HD UI POST Display failure |  | UI HD |  | HD POST: UI Disp Driver | Call service to schedule a repair. | 262 | ALARM_ID_HD_UI_POST_FAILURE_DISPLAY | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 201 | SRSHD 1657 |  | SRSUI 1070 | FALSE | HD UI POST Touch failure |  | UI HD |  | HD POST: UI Touch Driver | Call service to schedule a repair. | 263 | ALARM_ID_HD_UI_POST_FAILURE_TOUCH | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 201 | SRSHD 1658 |  | SRSUI 1071 | FALSE | HD UI POST SD-Card failure |  | UI HD |  | HD POST: UI SD-Card | SD-Card Failure \n- Power cycle the system.\n- Call service to schedule a repair if problem persists. | 264 | ALARM_ID_HD_UI_POST_FAILURE_SDCARD | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 201 SA 202 SA 203 | SRSHD 1659 |  | SRSUI 1072 | FALSE | HD UI POST RTC failure |  | UI HD |  | HD POST: UI RTC | Call service to schedule a repair. | 265 | ALARM_ID_HD_UI_POST_FAILURE_RTC | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 201 | SRSHD 1660 |  | SRSUI 1073 | FALSE | HD UI POST WiFi failure |  | UI HD |  | HD POST: UI WiFi | Call service to schedule a repair. | 266 | ALARM_ID_HD_UI_POST_FAILURE_WIFI | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 201 | SRSHD 1661 |  | SRSUI 1074 | FALSE | HD UI POST Bluetooth failure |  | UI HD |  | HD POST: UI Bluetooth | Call service to schedule a repair. | 267 | ALARM_ID_HD_UI_POST_FAILURE_BLUETOOTH | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 201 | SRSHD 1662 |  | SRSUI 1075 | FALSE | HD UI POST Ethernet failure |  | UI HD |  | HD POST: UI Ethernet | Call service to schedule a repair. | 268 | ALARM_ID_HD_UI_POST_FAILURE_ETHERNET | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 201 | SRSHD 1663 |  | SRSUI 1076 | FALSE | HD UI POST Sound failure |  | UI HD |  | HD POST: UI Sound | Call service to schedule a repair. | 269 | ALARM_ID_HD_UI_POST_FAILURE_SOUND | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 398 | SA 459 | SRSHD 1481 |  |  | FALSE | HD POST Safety Shutdown failure |  | HD |  | HD POST: Proc Safestate | Call service to schedule a repair. | 270 | ALARM_ID_HD_SAFETY_SHUTDOWN_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 292 |  | SRSDG 822 |  | FALSE | DG POST Safety Shutdown failure |  | DG |  | DG POST: Proc Safestate | Call service to schedule a repair. | 271 | ALARM_ID_DG_SAFETY_SHUTDOWN_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 872 | SA 221 | SRSHD 1394 |  |  | FALSE | HD Fan RPM out of range |  | HD |  | HD Alarm (Low): Fan Speed | Fan failure.\n- Check that the fans inlets are not blocked or obstructed. | 272 | ALARM_ID_HD_FAN_RPM_OUT_OF_RANGE |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
[6 empty] | FALSE | Available for use | [5 empty] | 273 | ALARM_ID_AVAILABLE_10 |  | ALARM_PRIORITY_LOW | 999 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
Blaine - Remove SRS (Done) Sean - Remove from code |  |  | Removed |  |  | FALSE | Available for use | [5 empty] | 274 | ALARM_ID_AVAILABLE_11 |  | ALARM_PRIORITY_LOW | 999 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 237 | SRSHD 1706 |  |  | FALSE | HD arterial pressure sensor is reading out of range |  | HD |  | HD Fault: Arterial Press Range | Call service to schedule a repair. | 275 | ALARM_ID_HD_ARTERIAL_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 234 | SRSHD 1707 |  |  | FALSE | HD venous pressure sensor is reading out of range |  | HD |  | HD Fault: Venous Press Range | Call service to schedule a repair. | 276 | ALARM_ID_HD_VENOUS_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
Sean to remove from code check of 32767 check (align to SRS) | PRS 494 | SA 231 | SRSHD 1708 |  |  | FALSE | HD BP occlusion sensor is reading out of range |  | HD |  | HD Fault: Occlusion Sens Range | Call service to schedule a repair. | 277 | ALARM_ID_HD_BP_OCCLUSION_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
Dara - update code to 1900g |  |  | SRSHD 793 |  |  | FALSE | HD active reservoir weight out of range. |  | HD |  | DG Fault: Reservoir Overfill | Call service to schedule a repair. | 278 | ALARM_ID_HD_ACTIVE_RESERVOIR_WEIGHT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
|  |  |  | SRSDG 752 |  | FALSE | DG dialysate drain time out |  | DG |  | DG Fault: Reservoir Drain | Call service to schedule a repair. | 279 | ALARM_ID_DG_DIALYSATE_DRAIN_TIME_OUT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
| PRS 863 | SA 340 SA 344 |  | SRSDG 888 |  | FALSE | DG drain pump direction invalid |  | DG |  | DG Fault: DRP Direction Fault | Call service to schedule a repair. | 280 | ALARM_ID_DG_DRAIN_PUMP_DIRECTION_INVALID |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
| PRS 827 |  |  | SRSDG 836 |  | TRUE | Acid concentration bottle low volume alarm |  | DG |  | Dial. Gen: Acid Conc. Low | Acid concentrate low.\n- Replace acid concentrate bottle. | 281 | ALARM_ID_DG_ACID_BOTTLE_LOW_VOLUME |  | ALARM_PRIORITY_MEDIUM | 412 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
| PRS 828 |  |  | SRSDG 837 |  | TRUE | Bicarbonate concentration bottle low volume alarm |  | DG |  | Dial. Gen: Bicarb Conc. Low | Bicarbonate concentrate low.\n- Replace bicarbonate concentrate bottle. | 282 | ALARM_ID_DG_BICARB_BOTTLE_LOW_VOLUME |  | ALARM_PRIORITY_MEDIUM | 413 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
x | PRS 863 | SA 309 SA 310 SA 314 SA 315 |  | SRSDG 615 |  | FALSE | DG load cells weight out of range |  | DG |  | DG Fault: LC Range | Call service to schedule a repair. | 283 | ALARM_ID_DG_LOAD_CELL_WEIGHT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 317 SA 318 |  | SRSDG 383 |  | FALSE | DG load cells primary/back up drift out of range |  | DG |  | DG Fault: Redundant LC Check | Call service to schedule a repair. | 284 | ALARM_ID_DG_LOAD_CELL_PRIMARY_BACKUP_DRIFT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
[6 empty] | FALSE | Available for use | [5 empty] | 285 | ALARM_ID_AVAILABLE_12 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE
[6 empty] | FALSE | DG concentrate cap not in proper position |  | DG |  | Cleaning: Conc Cap Removed | Concentrate cap removed from DG during cleaning.\n- Cleaning Process Terminated. | 286 | ALARM_ID_DG_CONCENTRATE_CAP_NOT_IN_PROPER_POSITION |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 222 | SRSHD 1403 |  |  | FALSE | HD processor clock speed checks against FPGA clock failure |  | HD |  | HD Fault: Proc-FPGA Clock | Call service to schedule a repair. | 287 | ALARM_ID_HD_FPGA_CLOCK_SPEED_CHECK_FAILURE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
|  | SA 491 SA 492 | SRSHD 562 |  |  | FALSE | HD load cells primary/back up drift out of range |  | HD |  | DG Fault: LC Drift Check | Call service to schedule a repair. | 288 | ALARM_ID_HD_LOAD_CELL_PRIMARY_BACKUP_DRIFT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
Cleaning |  |  |  | SRSDG 928 |  | FALSE | DG dialysate cap not closed |  | DG |  | Cleaning: Dialysate Cap Removed | Dialysate cap removed from DG during cleaning.\n- Cleaning Process Terminated. | 289 | ALARM_ID_DG_DIALYSATE_CAP_NOT_IN_PROPER_POSITION |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 233 | SRSHD 1584 |  |  | FALSE | HD venous pressure sensor FPGA fault |  | HD |  | HD Fault: Ven. Pres. Com | Call service to schedule a repair. | 290 | ALARM_ID_HD_VENOUS_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
Cleaning | [5 empty] | FALSE | HD processor is in Disinfect Flush mode |  | HD |  | Cleaning: Flush In Progress | System flush in progress\n- Please wait for the process to complete.\n- Treatment will not be able to be started until process \ncompletes.\n- Process may be terminated by pressing the HD stop button. | 291 | ALARM_ID_HD_DISINFECT_FLUSH |  | ALARM_PRIORITY_LOW | 840 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
Cleaning | [5 empty] | FALSE | HD processor is in Disinfect Heat mode |  | HD |  | Cleaning: Heat Disinf In Progress | Heat disinfect in progress\n- Please wait for the process to complete.\n- Treatment will not be able to be started until process \ncompletes.\n- Process may be terminated by pressing the HD stop button. | 292 | ALARM_ID_HD_DISINFECT_HEAT |  | ALARM_PRIORITY_LOW | 841 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
Cleaning | [5 empty] | FALSE | HD processor is in Disinfect Chemical mode |  | HD |  | Cleaning: Chem Flush In Progress | Chemical disinfect flush in progress\n- Please wait for the process to complete\n- Treatment will not be able to be started until process \ncompletes.\n- Process may be terminated by pressing the HD stop button. | 293 | ALARM_ID_HD_DISINFECT_CHEM |  | ALARM_PRIORITY_LOW | 842 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE
remove | [5 empty] | FALSE | Available for use | [5 empty] | 294 | ALARM_ID_AVAILABLE_13 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 398 | SA 470 | SRSHD 1664 |  | SRSUI 1238 | FALSE | HD UI POST CloudSync failure |  | UI HD |  | HD POST: UI CloudSync | Call service to schedule a repair. | 295 | ALARM_ID_HD_UI_POST_FAILURE_CLOUDSYNC | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 330 SA 331 |  | SRSDG 951 |  | FALSE | DG TDi sensors FPGA fault |  | DG |  | DG Fault: TDi RTD Comm | Call service to schedule a repair. | 296 | ALARM_ID_DG_TDI_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 326 SA 327 |  | SRSDG 950 |  | FALSE | DG TRo sensors FPGA fault |  | DG |  | DG Fault: TRo RTD Comm | Call service to schedule a repair. | 297 | ALARM_ID_DG_TRO_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 349 |  | SRSDG 948 |  | FALSE | DG baro sensor FPGA fault |  | DG |  | DG Fault: Baro Pres Comm | Call service to schedule a repair. | 298 | ALARM_ID_DG_BARO_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 483 |  | SRSDG 952 |  | FALSE | DG invalid serial number |  | DG |  | DG POST: Serial Number Check | Call service to schedule a repair. | 299 | ALARM_ID_DG_INVALID_SERIAL_NUMBER |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 398 | SA 473 | SRSHD 1582 |  |  | FALSE | HD invalid serial number |  | HD |  | HD POST: Serial Number Check | Call service to schedule a repair. | 300 | ALARM_ID_HD_INVALID_SERIAL_NUMBER |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
remove redundant with SRSDG 866 | [5 empty] | FALSE | Alarm ID available | [5 empty] | 301 | ALARM_ID_AVAILABLE_14 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 485 |  | SRSDG 126 |  | FALSE | DG processor RAM error |  | DG |  | DG Fault: Processor RAM error | Call service to schedule a repair. | 302 | ALARM_ID_DG_CPU_RAM_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 345 |  | SRSDG 958 |  | FALSE | DG drain pump direction FPGA fault |  | DG |  | DG Fault: DRP Hall Commutation | Call service to schedule a repair. | 303 | ALARM_ID_DG_DRAIN_PUMP_DIRECTION_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 398 | SA 197 | SRSHD 1587 |  |  | FALSE | HD invalid usage record |  | HD |  | HD POST: Usage Rec NVRAM CRC | Call service to schedule a repair. | 304 | ALARM_ID_HD_INVALID_USAGE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 494 | SA 484 | SRSHD 1577 |  |  | FALSE | HD processor RAM error |  | HD |  | HD Fault: Processor RAM error | Call service to schedule a repair. | 305 | ALARM_ID_HD_CPU_RAM_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
[6 empty] | FALSE | DG CD1 CD2 check fault |  | DG |  | DG Fault: CD1 CD2 Check | Call service to schedule a repair. | 306 | ALARM_ID_DG_CD1_CD2_CHECK_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 396 |  | SRSHD 1594 |  |  | TRUE | HD AC power lost second alarm |  | HD |  | Blood Stop: Loss of AC Power | HD power has been lost. \n- Treatment must be terminated.\n-Check all electrical connections.\n- The display and CPU will run off of the battery until \npower returns.\n-If power is returned prior to blood stopped countdown elapses, \nrinseback will be allowed. | 307 | ALARM_ID_HD_AC_POWER_LOST_IN_TREATMENT |  | ALARM_PRIORITY_HIGH | 100 | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 348 |  | SRSDG 959 |  | FALSE | DG CPi conductivity sensor invalid character received |  | DG |  | DG Fault: CPi Comm | Call service to schedule a repair. | 308 | ALARM_ID_DG_CPI_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 348 |  | SRSDG 960 |  | FALSE | DG CPo conductivity sensor invalid character received |  | DG |  | DG Fault: CPo Comm | Call service to schedule a repair. | 309 | ALARM_ID_DG_CPO_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 347 |  | SRSDG 962 |  | FALSE | DG CD1 conductivity sensor invalid character received |  | DG |  | DG Fault: CD1 Comm | Call service to schedule a repair. | 310 | ALARM_ID_DG_CD1_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 347 |  | SRSDG 961 |  | FALSE | DG CD2 conductivity sensor invalid character received |  | DG |  | DG Fault: CD2 Comm | Call service to schedule a repair. | 311 | ALARM_ID_DG_CD2_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
Sean - review code and requirements use raw FMD flow (add DG to publish if not doing so already) long persistance |  |  | SRSHD 1600 |  |  | FALSE | HD Dialysate In flow rate to too low. |  | HD |  | Dial. Gen: Dialysate In Flow | Dialysate in flow rate is low\n- Check dialysate line connection to the DG. | 312 | ALARM_ID_HD_DIAL_IN_FLOW_CHECK_FAILURE |  | ALARM_PRIORITY_MEDIUM | 400 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 852 |  |  | SRSDG 1064 SRSDG 1052 SRSDG 968 |  | FALSE | DG cleaning mode inlet water temperature too low. |  | DG |  | Cleaning: Inlet Water Temp Low | Cleaning Mode Inlet water temperature too low.\n- Cleaning mode process has terminated.\n- Check inlet water before restarting cleaning mode. | 313 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_TEMP_TOO_LOW |  | ALARM_PRIORITY_MEDIUM | 602 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 855 |  |  | SRSDG 1063 SRSDG 1051 SRSDG 967 |  | FALSE | DG cleaning mode inlet water pressure too high. |  | DG |  | Cleaning: Inlet Water Press High | Cleaning Mode Inlet water pressure too high.\n- Cleaning mode process has terminated.\n- Check inlet water before restarting cleaning mode. | 314 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_PRESSURE_TOO_HIGH |  | ALARM_PRIORITY_MEDIUM | 605 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 856 |  |  | SRSDG 1068 SRSDG 1056 SRSDG 966 |  | FALSE | DG cleaning mode inlet water pressure too low. |  | DG |  | Cleaning: Inlet Water Press Low | Cleaning Mode Inlet water pressure too low.\n- Cleaning mode process has terminated.\n- Check inlet water before restarting cleaning mode. | 315 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_PRESSURE_TOO_LOW |  | ALARM_PRIORITY_MEDIUM | 606 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
x | PRS 863 | SA 495 |  | SRSDG 983 |  | FALSE | DG acid concentrate pump park fault. |  | DG |  | DG Fault: Acid CP Pump Park | Call service to schedule a repair. | 316 | ALARM_ID_DG_ACID_CONCENTRATE_PUMP_PARK_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
x | PRS 863 | SA 496 |  | SRSDG 984 |  | FALSE | DG bicarb concentrate pump part fault. |  | DG |  | DG Fault: Bicarb CP Pump Park | Call service to schedule a repair. | 317 | ALARM_ID_DG_BICARB_CONCENTRATE_PUMP_PARK_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
Blaine to check with michael |  |  | SRSHD 1609 |  |  | TRUE | HD Air Pump timeout alarm |  | HD |  | HD Fault: Air Pump Timeout | Unable to reduce fluid level.\n- Air trap may be overfilled.\n- Pressure transducer protector may be contaminated.\n- Treatment must be terminated without option for rinseback.\n- Call service to schedule a repair. | 318 | ALARM_ID_HD_AIR_PUMP_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
Cleaning |  |  |  | SRSDG 1040 |  | FALSE | DG chemical disinfect flush flush sample. |  | DG |  | Cleaning: Chem Residual Sampling | Collect chemical residuals water sample from drain line. | 319 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_FLUSH_SAMPLE |  | ALARM_PRIORITY_LOW | 752 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE
Cleaning |  |  |  | SRSDG 1043 |  | FALSE | DG chemical disinfect flush sample timeout. |  | DG |  | Cleaning: Chem Residual Timeout | During Chemical Disinfection Flush\n-user residual sampling timed out. | 320 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_SAMPLE_TIMEOUT |  | ALARM_PRIORITY_LOW | 752 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE
x | PRS 863 | SA 356 |  | SRSDG 1069 |  | FALSE | DG outlet UV reactor not healthy |  | DG |  | DG Fault: UVo Error | Call service to schedule a repair. | 321 | ALARM_ID_DG_OUTLET_UV_REACTOR_NOT_HEALTHY |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE
[6 empty] | FALSE | HD UI POST Year check failure |  | UI HD |  | HD POST Alarm: UI Invalid Year | Call service to schedule a repair. | 322 | ALARM_ID_HD_UI_POST_FAILURE_INVALID_YEAR |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE

## Property Definitions
Property Name | Definition |  | Column | Description
AlarmID | Unique alarm identification number |  | PRS ID | PRS requirement ID associated with the alarm.
Alarm Name | Name of the alarm |  | HD SRS ID | HD SRS requirement ID (if applicable) associated with the alarm.
Priority | Priority of the alarm (High, Medium, Low) |  | DG SRS ID | DG SRS requirement ID (if applicable) associated with the alarm.
Rank | Within the same priority, higher rank alarms have higher priority |  | Device | Sub-system that detects the alarm.
Fault | The alarm activation causes an HD fault |  | Description | Detailed description of the alarm.
DGFault | The alarm activation causes an DG fault |  | Displayed Title | Alarm title displayed prominantly at top of alarm dialog when the alarm is top priority.
Stop | The alarm activation causes a controlled stop |  | Displayed Instructions | Instructions displayed in alarm dialog when the alarm is top priority.
Clear Immediate | Alarm condition is cleared immediately after alarm activation |  | Trigger Conditions | Condition(s) required to trigger the alarm.
Non Recoverable | Alarm cannot be cleared (unrecoverable) |  | AlarmID | Unique alarm ID #.
No Resume | Alarm activation prevents treatment resume |  | Alarm Name | Enumeration name (code reference) for the alarm.
No Rinseback | Alarm activation prevents rinseback |  | Priority | Priority of the alarm (high, medium or low per 60601-1-8.
No End Treatment | Alarm activation prevents ending treatment |  | Rank | Rank of the alarm (used to determine which alarm is top priority) (low rank=high priority).
No New Treatment | Alarm activation prevents new treatment |  | Fault | Does HD f/w consider the alarm to be a system fault.
No Dialysate Recirc | Alarm activation prevents dialysate re-circulation |  | DGFault | Does DG f/w consider the alarm to be a system fault (will not prevent HD from some activities).
Clear Top Only | User acknowledgment clears only displayed alarm |  | Stop | Does the HD f/w need to pause treatment (or perhaps other activities) when alarm is active.
Escalate In | Alarm will escalate if not cleared within this time (zero indicates no escalation) |  | Clear Immediate | Should the f/w immediately clear the alarm condition after trigger (not clear the alarm itself).
Escalate To | ID of the alarm that this alarm will be escalated to |  | Non Recoverable | Is the alarm recoverable (can user clear the alarm).
|  |  | No Resume | Should treatment resume be prevented when the alarm is active.
|  |  | No Rinseback | Should rinseback be prevented when the alarm is active.
|  |  | No End Treatment | Should ending treatment be prevented when the alarm is active.
Category | Priority |  | No Blood Recirc | Should blood recirculation be prevented when the alarm is active (treatment is paused). Dialyzer is bypassed regardless.
HD POST | High |  | No Dialysate Recirc | Should dialysate recirculation be prevented when the alarm is active (treatment is paused). Dialyzer is bypassed regardless.
HD Machine Alarm | High |  | Clear Top Only | If user acknowledges alarm, should all active alarms be cleared or just the alarm top (displayed) alarm.
DG POST DG Machine Alarm | High |  | TreatmentLog | Should alarm be included in treatment log.
Blood Stopped - No Rinseback Allowed | HIgh |  | Escalate In | Duration (in seconds) until alarm escalates to a higher priority alarm (0 indicates no escalation).
Blood Stopped - Rinseback Allowed | High |  | Escalate To | Higher priority alarm to escalate to when escalation period expires (None indicates no escalation).
End Pre-Tx | Medium |  | Doxygen Title | Used for design documentation purposes.
Concentrate Warning | Low
Concentrate Alarm | Medium
Saline Warning | Low
Saline Alarm | Medium
Dialysate Alarms (Temp, conductivity, recirc?) | Medium
Other Procedural (setup, pre-tx, etc) | Medium/Low
Disinfection Alarms | Medium
HD Machine Alert | Low
DG Machine Alert | Low
Blood Pressure Alarms | N/A - not supported
Priority | Type | Mode