## Alarms
PRS ID | HD SRS ID | DG SRS ID | Device | Description | Displayed Title | Displayed Instructions | Trigger Conditions | AlarmID | Alarm Name | Priority | Rank | Fault | DGFault | Stop | Clear Immediate | Non Recoverable | No Resume | No Rinseback | No End Treatment | No Blood Recirc | No Dialysate Recirc | Clear Top Only | TreatmentLog | Escalate In | Escalate To | Doxygen Title
N/A | N/A | N/A | N/A | Indication for when no alarms is a possible situation | N/A | N/A | N/A | 0 | ALARM_ID_NO_ALARM | ALARM_PRIORITY_NONE | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Indication for when no alarms is a possible situation | Indication for when no alarms is a possible situation
|  |  | HD | Stuck button POST failure. Stop or Off button detected to be pressed for at least 1 second during test shortly after power up | Self Test Error | Call service to schedule a repair. | Stop or Off button are pressed for more than 1 sec during POST. | 1 | ALARM_ID_STUCK_BUTTON_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Stuck button POST failure. Stop or Off button detected to be pressed for at least 1 second during test shortly after power up | Stuck button POST failure. Stop or Off button detected to be pressed for at least 1 second during test shortly after power up
|  |  | HD | HD FPGA POST failure | Self Test Error | Call service to schedule a repair. | Minimum HD FPGA major revision is not compatible with this version of HD firmware or the FPGA Id is not correct. | 2 | ALARM_ID_HD_FPGA_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD FPGA POST failure | HD FPGA POST failure
|  |  | DG | DG FPGA POST failure | Self Test Error | Call service to schedule a repair. | Minimum DG FPGA major and minor revision are not equal to the FPGA major and minor revision in FPGA header | 3 | ALARM_ID_DG_FPGA_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG FPGA POST failure | DG FPGA POST failure
|  |  | HD | HD Watchdog POST failure | Self Test Error | Call service to schedule a repair. | If voltage reading is less than 22.6Volts when the Watchdog recovered If PIN_SIGNAL is not LOW If backup alarm audio is on when Watchdog expired | 4 | ALARM_ID_HD_WATCHDOG_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD Watchdog POST failure | HD Watchdog POST failure
|  |  | DG | DG Watchdog POST failure | Self Test Error | Call service to schedule a repair. | If voltage reading > 5.0 Volts when the Watchdog recovered If PIN_SIGNAL is not LOW If voltage reading is less than 22.6Volts when the Watchdog recovered | 5 | ALARM_ID_DG_WATCHDOG_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG Watchdog POST failure | DG Watchdog POST failure
|  |  | HD | UI communication POST failure. UI failed to communicate within a reasonable time after power up | Self Test Error | Call service to schedule a repair. | If the UI failed to checkin via the CAN bus within 40 seconds | 6 | ALARM_ID_UI_COMM_POST_FAILED | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | UI communication POST failure. UI failed to communicate within a reasonable time after power up | UI communication POST failure. UI failed to communicate within a reasonable time after power up
|  |  | ??? | Real-time clock configuration error | Machine Error | Call service to schedule a repair. | If the RTC is not functional as indicated by the bits in the first 3 control registers. | 7 | ALARM_ID_HD_RTC_BATTERY_LOW | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | HD RTC battery low | Alarm ID available for use
|  |  | HD | HD accelerometer failed POST | Self Test Error | Call service to schedule a repair. | If the calculated vector length of the accelerometer x, y, z is not the nominal length. | 8 | ALARM_ID_HD_ACCELEROMETER_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD accelerometer failed POST | HD accelerometer failed POST
|  |  | DG | DG accelerometer failed POST | Self Test Error | Call service to schedule a repair. | If the calculated vector length of the accelerometer x, y, z is not the nominal length. | 9 | ALARM_ID_DG_ACCELEROMETER_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG accelerometer failed POST | DG accelerometer failed POST
|  |  | ??? | RTC or MCU timer inaccurate | Machine Error | Call service to schedule a repair. | If the time that has been elapsed is not within the specified time tolerance. For RTC tolerance is 75. For FPGA tolerance is 12. | 10 | ALARM_ID_HD_BLOOD_LEAK_FAULT | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Alarm ID available for use | Alarm ID available for use
|  |  | DG | DG heaters failed POST |  |  | This alarm is not used in the DG code | 11 | ALARM_ID_HD_BLOOD_LEAK_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Alarm ID available for use | DG heaters failed POST
|  |  | HD | HD firmware image integrity POST test failed | Self Test Error | Call service to schedule a repair. | If the calculated CRC of the HD image does not match the expected CRC. | 12 | ALARM_ID_HD_INTEGRITY_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD firmware image integrity POST test failed | HD firmware image integrity POST test failed
|  |  | DG | DG firmware image integrity POST test failed | Self Test Error | Call service to schedule a repair. | If the calculated CRC of the DG firmware image does not matched the expected CRC | 13 | ALARM_ID_DG_INTEGRITY_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG firmware image integrity POST test failed | DG firmware image integrity POST test failed
|  |  | HD | HD blood flow invalid calibration |  |  | This alarm ID is not used in the HD code. | 14 | ALARM_ID_HD_BLOOD_FLOW_INVALID_CALIBRATION | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood flow invalid calibration | HD blood flow invalid calibration
|  |  | HD | HD dialysate flow invalid calibration |  |  | This alarm ID is not used in the HD code. | 15 | ALARM_ID_HD_DIALYSATE_FLOW_INVALID_CALIBRATION | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD dialysate flow invalid calibration | HD dialysate flow invalid calibration
|  |  | HD | HD alarm audio failed POST | Self Test Error | Call service to schedule a repair. | If alarm audio current is NOT sufficiently high for at more than 1 second indicating alarm tone is being output | 16 | ALARM_ID_HD_ALARM_AUDIO_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD alarm audio failed POST | HD alarm audio failed POST
|  |  | HD | HD UI POST failed | Self Test Error | Call service to schedule a repair. | If User Interface has not sent POST results within 2 seconds. | 17 | ALARM_ID_HD_UI_POST_FAILED | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD UI POST failed | HD UI POST failed
|  |  | DG | DG load cells invalid calibration record | Calibration Error | Call service to schedule a repair. | If the non-volatile data is valid and if not raise the alarm | 18 | ALARM_ID_DG_PRESSURE_SENSORS_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG load cells invalid calibration record | DG load cells invalid calibration record
|  |  | DG | DG flow sensors invalid calibration record | Calibration Error | Call service to schedule a repair. | If the non-volatile data is valid and if not raise the alarm | 19 | ALARM_ID_DG_FLOW_SENSORS_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG flow sensors invalid calibration record | DG flow sensors invalid calibration record
|  |  | DG | DG conductivity sensors invalid calibration record | Calibration Error | Call service to schedule a repair. | If the calibration data that was received from NVDataMgmt is NOT legitimate | 20 | ALARM_ID_DG_COND_SENSORS_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG conductivity sensors invalid calibration record | DG conductivity sensors invalid calibration record
|  |  | DG | DG drain line volume invalid calibration record | Calibration Error | Call service to schedule a repair. | If the calibration data that was received from NVDataMgmt is NOT legitimate | 21 | ALARM_ID_DG_DRAIN_LINE_VOLUME_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG drain line volume invalid calibration record | DG drain line volume invalid calibration record
|  |  | DG | DG reservoirs invalid calibration record | Calibration Error | Call service to schedule a repair. | If the calibration data that was received from NVDataMgmt is NOT legitimate | 22 | ALARM_ID_DG_RESERVOIRS_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG reservoirs invalid calibration record | DG reservoirs invalid calibration record
|  |  | DG | DG acid concentrate calibration record | Calibration Error | Call service to schedule a repair. | If the non-volatile data is valid and if not raise the alarm | 23 | ALARM_ID_DG_ACID_CONCENTRATE_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG acid concentrate calibration record | DG acid concentrate calibration record
|  |  | DG | DG bicarb concentrate calibration record | Calibration Error | Call service to schedule a repair. | If the non-volatile data is valid and if not raise the alarm | 24 | ALARM_ID_DG_BICARB_CONCENTRATE_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG bicarb concentrate calibration record | DG bicarb concentrate calibration record
|  |  | DG | DG accelerometer invalid calibration record | Calibration Error | Call service to schedule a repair. | If the calibration data that was received from NVDataMgmt is NOT legitimate | 25 | ALARM_ID_DG_ACCELEROMETERS_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG accelerometer invalid calibration record | DG accelerometer invalid calibration record
|  |  | HD | HD accelerometer invalid calibration record | Calibration Error | Call service to schedule a repair. | If the accel calibration data that was received from NVDataMgmt is NOT legitimate. | 26 | ALARM_ID_HD_ACCELEROMETERS_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD accelerometer invalid calibration record | HD accelerometer invalid calibration record
|  |  | HD | HD blood flow invalid calibration record |  |  | This alarm ID is not used in the HD code. | 27 | ALARM_ID_HD_BLOOD_FLOW_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood flow invalid calibration record | HD blood flow invalid calibration record
|  |  | HD | HD dialyzer flow invalid calibration record |  |  | This alarm ID is not used in the HD code. | 28 | ALARM_ID_HD_DIALYSATE_FLOW_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD dialyzer flow invalid calibration record | HD dialyzer flow invalid calibration record
|  |  | HD | HD heparin force sensor invalid calibration record | Calibration Error | Call service to schedule a repair. | If the heparin force sensor calibration data that was received from NVDataMgmt is NOT legitimate. | 29 | ALARM_ID_HD_HEPARIN_FORCE_SENSOR_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD heparin force sensor invalid calibration record | HD heparin force sensor invalid calibration record
|  |  | HD | HD Software fault. Software found itself in an unexpected state | Machine Error | Call service to schedule a repair. | If the specific software faults listed in alarm management SW faults has been detected | 30 | ALARM_ID_HD_SOFTWARE_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD Software fault. Software found itself in an unexpected state | HD Software fault. Software found itself in an unexpected state
|  |  | HD | Blood pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | Blood Pump Error | Call service to schedule a repair. | If the measured MC current is not greater than BP_MAX_CURR_WHEN_STOPPED_MA or is not greater than BP_MAX_CURR_WHEN_RUNNING_MA. | 31 | ALARM_ID_BLOOD_PUMP_MC_CURRENT_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | Blood pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running
|  |  | HD | Blood pump failed motor off check. Measured speed while commanded off | Blood Pump Error | Call service to schedule a repair. | If the measured motor speed is greater than BP_MAX_MOTOR_SPEED_WHILE_OFF_RPM | 32 | ALARM_ID_BLOOD_PUMP_OFF_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood pump failed motor off check. Measured speed while commanded off | Blood pump failed motor off check. Measured speed while commanded off
|  |  | HD | Blood pump failed motor direction check. Measured vs. commanded | Blood Pump Error | Call service to schedule a repair. | If the set direction versus hall sensors direction versus sign of motor speed DOES NOT agree. | 33 | ALARM_ID_BLOOD_PUMP_MC_DIRECTION_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood pump failed motor direction check. Measured vs. commanded | Blood pump failed motor direction check. Measured vs. commanded
|  |  | HD | Blood pump failed rotor speed check. Mismatch with rotor and motor speeds | Blood Pump Error | Call service to schedule a repair. | If the delta of the measured motor speed and the commanded motor speed is not within tolerance. | 34 | ALARM_ID_BLOOD_PUMP_ROTOR_SPEED_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood pump failed rotor speed check. Mismatch with rotor and motor speeds | Blood pump failed rotor speed check. Mismatch with rotor and motor speeds
|  |  | HD | Dialysis inlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | Dialysate Pump Error | Call service to schedule a repair. | If the Motor Controller measured DPi pump current > 150.0 mA for more than 250.0 msec. | 35 | ALARM_ID_DIAL_IN_PUMP_MC_CURRENT_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis inlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | Dialysis inlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running
|  |  | HD | Dialysis inlet pump failed motor off check. Measured speed while commanded off | Dialysate Pump Error | Call service to schedule a repair. | If during command off state, the DPi pump motor speed is measured > 100.0 RMP. | 36 | ALARM_ID_DIAL_IN_PUMP_OFF_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis inlet pump failed motor off check. Measured speed while commanded off | Dialysis inlet pump failed motor off check. Measured speed while commanded off
|  |  | HD | Dialysis inlet pump failed motor direction check. Measured vs. commanded | Dialysate Pump Error | Call service to schedule a repair. | If the set direction versus hall sensors direction DO NOT agree for more than 250.0 msec. | 37 | ALARM_ID_DIAL_IN_PUMP_MC_DIRECTION_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis inlet pump failed motor direction check. Measured vs. commanded | Dialysis inlet pump failed motor direction check. Measured vs. commanded
|  |  | HD | Dialysis inlet pump failed rotor speed check. Mismatch with rotor and motor speeds | Dialysate Pump Error | Call service to schedule a repair. | If the measured motor speed versus the measured motor speed while controlling the target > 5.0 for more than 12 seconds. | 38 | ALARM_ID_DIAL_IN_PUMP_ROTOR_SPEED_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis inlet pump failed rotor speed check. Mismatch with rotor and motor speeds | Dialysis inlet pump failed rotor speed check. Mismatch with rotor and motor speeds
|  |  | HD | Dialysis outlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | Dialysate Pump Error | Call service to schedule a repair. | If the DPo pump current > 150.0 mA while stopped or Dpo current > 2000.0 mA for more than 2 seconds while running. | 39 | ALARM_ID_DIAL_OUT_PUMP_MC_CURRENT_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis outlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | Dialysis outlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running
|  |  | HD | Dialysis outlet pump failed motor off check. Measured speed while commanded off | Dialysate Pump Error | Call service to schedule a repair. | If the Dpo pump current > 100.0 RPM while commanded off for more than 5000 msec. | 40 | ALARM_ID_DIAL_OUT_PUMP_OFF_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis outlet pump failed motor off check. Measured speed while commanded off | Dialysis outlet pump failed motor off check. Measured speed while commanded off
|  |  | HD | Dialysis outlet pump failed motor direction check. Measured vs. commanded | Dialysate Pump Error | Call service to schedule a repair. | If the DPo pump set direction does not agree with hall sensors or set direction does not agree with the sign of the motor speed for more than 25 msec. | 41 | ALARM_ID_DIAL_OUT_PUMP_MC_DIRECTION_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis outlet pump failed motor direction check. Measured vs. commanded | Dialysis outlet pump failed motor direction check. Measured vs. commanded
|  |  | HD | Dialysis outlet pump failed rotor speed check. Mismatch with rotor and motor speeds | Dialysate Pump Error | Call service to schedule a repair. | If DPo pump measured rotor speed does not agreed with the measured/commanded motor speed | 42 | ALARM_ID_DIAL_OUT_PUMP_ROTOR_SPEED_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysis outlet pump failed rotor speed check. Mismatch with rotor and motor speeds | Dialysis outlet pump failed rotor speed check. Mismatch with rotor and motor speeds
|  |  | ??? | Watchdog expired error. Watchdog was not "pet" within the appropriate time | Machine Error | Call service to schedule a repair. | If the watchdog has expired | 43 | ALARM_ID_HD_BLOOD_LEAK_SELF_TEST_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood leak self test fault | Alarm ID available for use
|  |  | HD | User interface communication timeout | Machine Error | Call service to schedule a repair. | If the sub-system communication timeout has occurred. UI is 5 seconds. DG is 2 seconds. | 44 | ALARM_ID_UI_COMM_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | User interface communication timeout | User interface communication timeout
|  |  | ??? | Too many bad CRCs detected on received system messages | Machine Error | Call service to schedule a repair. | If too many bad messages CRCs are received within a set period of time | 45 | ALARM_ID_HD_COMM_TOO_MANY_BAD_CRCS | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD too many bad communications CRC | HD too many bad communications CRC
|  |  | ??? | System message that required acknowledgment was not acknowledged | Machine Error | Call service to schedule a repair. | If no retries left in the pending acknowlegement list. | 46 | ALARM_ID_CAN_MESSAGE_NOT_ACKED | ALARM_PRIORITY_HIGH | 3 | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | System message that required acknowledgment was not acknowledged | System message that required acknowledgment was not acknowledged
|  |  | HD | Ultrafiltration rate is too high error during treatment | Ultrafiltration Error | Call service to schedule a repair. | If the UF rate > 2750.0 mL per hour. | 47 | ALARM_ID_UF_RATE_TOO_HIGH_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Ultrafiltration rate is too high error during treatment | Ultrafiltration rate is too high error during treatment
|  |  | HD | Ultrafiltration volume accuracy error during treatment | Ultrafiltration Error | Call service to schedule a repair. | If the UF volume (refUFvolume - measureVolume) > 250.0 mL. | 48 | ALARM_ID_UF_VOLUME_ACCURACY_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Ultrafiltration volume accuracy error during treatment | Ultrafiltration volume accuracy error during treatment
|  |  | HD | HD FPGA communication down for too long | Machine Error | Call service to schedule a repair. | If retries for commands > 5 counts. | 49 | ALARM_ID_HD_FPGA_COMM_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD FPGA communication down for too long | HD FPGA communication down for too long
|  |  | DG | DG FPGA not accepting commanded valve states | Machine Error | Call service to schedule a repair. | If the read back FPGA valve states does not match the commanded valve states and the mismatch counter > 3 | 50 | ALARM_ID_VALVE_CONTROL_FAILURE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | FPGA not accepting commanded valve states | FPGA not accepting commanded valve states
|  |  | HD | Blood pump failed motor speed check. Measured vs. commanded | Blood Pump Error | Call service to schedule a repair. | If measured motor speed versus commanded rotor speed exceeded limits. | 51 | ALARM_ID_BLOOD_PUMP_MOTOR_SPEED_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood pump failed motor speed check. Measured vs. commanded | Blood pump failed motor speed check. Measured vs. commanded
|  |  | HD | Dialysate inlet pump failed motor speed check. Measured vs. commanded | Dialysate Pump Error | Call service to schedule a repair. | If measured motor speed versus commanded motor speed exceeded limits for more than 5 secs. | 52 | ALARM_ID_DIAL_IN_PUMP_MOTOR_SPEED_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysate inlet pump failed motor speed check. Measured vs. commanded | Dialysate inlet pump failed motor speed check. Measured vs. commanded
|  |  | HD | Dialysate outlet pump failed motor speed check. Measured vs. commanded | Dialysate Pump Error | Call service to schedule a repair. | If measured motor speed versus commanded motor speed exceeded limits for more than 5 secs. | 53 | ALARM_ID_DIAL_OUT_PUMP_MOTOR_SPEED_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysate outlet pump failed motor speed check. Measured vs. commanded | Dialysate outlet pump failed motor speed check. Measured vs. commanded
|  |  | HD | HD critical data integrity check failed | Machine Error | Call service to schedule a repair. | If the critical data failed range check. | 54 | ALARM_ID_HD_CRITICAL_DATA_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD critical data integrity check failed | HD critical data integrity check failed
|  |  | DG | DG critical data integrity check failed | Communication Failure | Call service to schedule a repair. | If the critical data failed range check. | 55 | ALARM_ID_DG_CRITICAL_DATA_ERROR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG critical data integrity check failed | DG critical data integrity check failed
|  |  | HD | HD accelerometer error (no readings or FPGA reports error) | Machine Error | Call service to schedule a repair. | If no fresh sample of accelerometer data has been detected for more than 10 msec | 56 | ALARM_ID_HD_ACCELEROMETER_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD accelerometer error (no readings or FPGA reports error) | HD accelerometer error (no readings or FPGA reports error)
|  |  | DG | DG accelerometer error (no readings or FPGA reports error) | Machine Error | Call service to schedule a repair. | If the accel sensor sampling counter as reported by the FPGA indicated 0 count | 57 | ALARM_ID_DG_ACCELEROMETER_FAILURE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG accelerometer error (no readings or FPGA reports error) | DG accelerometer error (no readings or FPGA reports error)
|  |  | HD | HD valve homing failed | Valve Error | Call service to schedule a repair. | If the number of failed homing > = 3 counts. | 58 | ALARM_ID_HD_VALVE_HOMING_FAILED | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD valve homing failed | HD valve homing failed
|  |  | HD | HD valve transition time out | Valve Error | Call service to schedule a repair. | if the valve's transition time has timed out (3000msec) | 59 | ALARM_ID_HD_VALVE_TRANSITION_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD valve transition time out | HD valve transition time out
|  |  | HD | HD valve not functional | Valve Error | Call service to schedule a repair. | If the valve bit masks are not as expected in different valve control mode. | 60 | ALARM_ID_HD_VALVE_NOT_FUNCTIONAL | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD valve not functional | HD valve not functional
|  |  | HD | HD valve current out of range | Valve Error | Call service to schedule a repair. | If the valve current is over the threshold of 1.0 A for 1000 msec. | 61 | ALARM_ID_HD_VALVE_CURRENT_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD valve current out of range | HD valve current out of range
|  |  | HD | HD valve position out of target | Valve Error | Call service to schedule a repair. | If the current position has deviated from the expected position for more than 1 second. | 62 | ALARM_ID_HD_VALVE_POSITION_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD valve position out of target | HD valve position out of target
|  |  | HD | FPGA reports failure reading arterial pressure sensor | Arterial Pressure Error | Call service to schedule a repair. | If the MSB in the FPGA arterial pressure indicated error (> 0) | 63 | ALARM_ID_ARTERIAL_PRESSURE_SENSOR_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | FPGA reports failure reading arterial pressure sensor | FPGA reports failure reading arterial pressure sensor
|  |  | HD | FPGA reports failure reading venous pressure sensor | Venous Pressure Error | Call service to schedule a repair. | If the venous pressure sensor temperature > 80.0 or < 0.0 | 64 | ALARM_ID_VENOUS_PRESSURE_SENSOR_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | FPGA reports failure reading venous pressure sensor | FPGA reports failure reading venous pressure sensor
|  |  | HD | HD requests DG command with invalid parameter fault | Machine Error | Call service to schedule a repair. | If the latest DG command response indicated invalid parameter | 65 | ALARM_ID_DG_COMMAND_INVALID_PARAMETER_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD requests DG command with invalid parameter fault | HD requests DG command with invalid parameter fault
|  |  | HD | HD blood leak sensor set point set failure |  |  | This alarm ID is not used in the HD code | 66 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_SET_POINT_SET_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood leak sensor set point set failure | HD blood leak sensor set point set failure
|  |  | HD | HD blood pump occlusion self-test failure alarm | Self Test Error | Call service to schedule a repair. | If the blood pump pressure > 5000 | 67 | ALARM_ID_HD_BP_OCCLUSION_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood pump occlusion self-test failure alarm | HD blood pump occlusion self-test failure alarm
|  |  | HD | HD active reservoir recirculation out of range | Dialysate Recirculation Error | Dialysate has recirculated for too long. | If the recirculation level has exceeded the max allowed of 1.1 | 68 | ALARM_ID_HD_ACTIVE_RESERVOIR_RECIRCULATION_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD active reservoir recirculation out of range | HD active reservoir recirculation out of range
|  |  | HD | HD blood leak sensor invalid calibration record |  |  | This alarm ID is not used in the HD code. | 69 | ALARM_ID_HD_BLOOD_LEAK_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood leak sensor invalid calibration record | HD blood leak sensor invalid calibration record
|  |  | HD | HD arterial pressure self-test failure alarm | Self Test Error | Call service to schedule a repair. | If arterial pressure < -300 or > 100 mmHg | 70 | ALARM_ID_HD_ARTERIAL_PRESSURE_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD arterial pressure self-test failure alarm | HD arterial pressure self-test failure alarm
|  |  | HD | HD venous pressure self-test failure alarm | Self Test Error | Call service to schedule a repair. | If the venous pressure < -100 or > 600 mmHg | 71 | ALARM_ID_HD_VENOUS_PRESSURE_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD venous pressure self-test failure alarm | HD venous pressure self-test failure alarm
|  |  | HD | HD blood flow meter status check self-test failure alarm |  |  | This alarm ID is not used in the HD code. | 72 | ALARM_ID_HD_BLOOD_FLOW_STATUS_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood flow meter status check self-test failure alarm | HD blood flow meter status check self-test failure alarm
|  |  | HD | HD dialysate flow meter status check self-test failure alarm |  |  | This alarm ID is not used in the HD code. | 73 | ALARM_ID_HD_DIALYSATE_FLOW_STATUS_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD dialysate flow meter status check self-test failure alarm | HD dialysate flow meter status check self-test failure alarm
|  |  | HD | HD blood leak detector self-test failure alarm | Self Test Error | Call service to schedule a repair. | If the FPGA reported no blood leak detected for more than 500 msec | 74 | ALARM_ID_HD_BLOOD_LEAK_SELF_TEST_SEND_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood leak detector self-test transmit timeout | HD blood leak detector self-test failure alarm
|  |  | HD | HD syringe pump self-test failure alarm | Self Test Error | Call service to schedule a repair. | If the calibration data that was received from NVDataMgmt is NOT legitimate. | 75 | ALARM_ID_HD_SYRINGE_PUMP_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump self-test failure alarm | HD syringe pump self-test failure alarm
|  |  | HD | HD monitored voltage is out of range | Machine Error | Call service to schedule a repair. | This alarm ID is not used in the HD code. | 76 | ALARM_ID_HD_VOLTAGE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD monitored voltage is out of range | HD monitored voltage is out of range
|  |  | DG | DG monitored voltage is out of range | Machine Error | Call service to schedule a repair. | Any of the monitor voltages is out of range for more than 20 msec. | 77 | ALARM_ID_DG_VOLTAGE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG monitored voltage is out of range | DG monitored voltage is out of range
|  |  | HD | HD syringe pump direction (from encoder) error | Syringe Pump Error | Call service to schedule a repair. | If the measured encoder direction of the syringe pump does not agree with the expected direction | 78 | ALARM_ID_HD_SYRINGE_PUMP_ENCODER_DIRECTION_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump direction (from encoder) error | HD syringe pump direction (from encoder) error
|  |  | HD | HD syringe pump direction (from controller) error | Syringe Pump Error | Call service to schedule a repair. | If the measured direction of the syringe pump motor controller does not agree with the expected direction | 79 | ALARM_ID_HD_SYRINGE_PUMP_CONTROLLER_DIRECTION_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump direction (from controller) error | HD syringe pump direction (from controller) error
|  |  | HD | HD syringe pump fault reported by FPGA | Syringe Pump Error | Call service to schedule a repair. | If the FPGA reported the syringe pump fault | 80 | ALARM_ID_HD_SYRINGE_PUMP_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump fault reported by FPGA | HD syringe pump fault reported by FPGA
|  |  | HD | HD syringe pump over travel error | Syringe Pump Error | Call service to schedule a repair. | If the syringe pump position < SYRINGE_PUMP_RETRACT_POS_MIN | 81 | ALARM_ID_HD_SYRINGE_PUMP_OVER_TRAVEL_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump over travel error | HD syringe pump over travel error
|  |  | HD | HD syringe pump DAC write failure | Syringe Pump Error | Call service to schedule a repair. | If the syringe pump DAC write error is reported by the FPGA | 82 | ALARM_ID_HD_SYRINGE_PUMP_DAC_WRITE_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump DAC write failure | HD syringe pump DAC write failure
|  |  | HD | HD syringe pump is running while the BP is off | Syringe Pump Error | Call service to schedule a repair. | If the syringe pump is on while the blood pump is off | 83 | ALARM_ID_HD_SYRINGE_PUMP_RUNNING_WHILE_BP_OFF_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump is running while the BP is off | HD syringe pump is running while the BP is off
|  |  | HD | HD venous pressure sensor not being read | Venous Pressure Error | Call service to schedule a repair. | If the stale venous pressure reading is detected | 84 | ALARM_ID_HD_VENOUS_PRESSURE_READ_TIMEOUT_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD venous pressure sensor not being read | HD venous pressure sensor not being read
|  |  | HD | HD venous pressure sensor temperature out of range error | Venous Pressure Error | Call service to schedule a repair. | If the venous pressure sensor temperature > 80.0 or < 0.0 | 85 | ALARM_ID_PUMP_TRACK_LATCH_OPENED | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD pump track latch opened alarm | HD pump track latch opened alarm
|  |  | HD | HD BP occlusion sensor not being read | Occlusion Sensor Error | Call service to schedule a repair. | If the stale occlusion reading is detected | 86 | ALARM_ID_HD_BP_OCCLUSION_READ_TIMEOUT_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD BP occlusion sensor not being read | HD BP occlusion sensor not being read
|  |  |  | This alarm ID is available for use |  |  | Available for use | 87 | ALARM_ID_DG_HEATING_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG heating invalid calibration record | DG heating invalid calibration record
|  |  |  | This alarm ID is available for use |  |  | Available for use | 88 | ALARM_ID_DG_CONC_PUMP_HALL_SENSOR_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG concentrate pumps hall sensor out of range | DG concentrate pumps hall sensor out of range
|  |  | HD | HD BP occlusion sensor error | Occlusion Sensor Error | Call service to schedule a repair. | If the blood pump occlusion sensor error is reported by the FPGA | 89 | ALARM_ID_HD_BP_OCCLUSION_SENSOR_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD BP occlusion sensor error | HD BP occlusion sensor error
|  |  |  | This alarm ID is available for use |  |  | Available for use | 90 | ALARM_ID_DG_LOAD_CELL_FPGA_READ_ERROR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG load cells FPGA read error. | DG load cells FPGA read error.
|  |  | HD | HD No dialysate flow data receive in the last 3 seconds | Dialysate Measurement Error | Call service to schedule a repair. | If not receiving new dialysate flow readings in expected interval (3 seconds) | 91 | ALARM_ID_HD_DIALYSATE_FLOW_DATA_NOT_RECEIVE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD No dialysate flow data receive in the last 3 seconds | HD No dialysate flow data receive in the last 3 seconds
|  |  |  | This alarm ID is available for use |  |  | Available for use | 92 | ALARM_ID_DG_CONDUCTIVITY_SENSOR_BAD_STATUS | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG conductivity sensor bad status. | DG conductivity sensor bad status.
|  |  |  | This alarm ID is available for use |  |  | Available for use | 93 | ALARM_ID_DG_OUTLET_PRIMARY_CONDUCTIVITY_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG outlet primary conductivity out of range | DG outlet primary conductivity out of range
|  |  |  | This alarm ID is available for use |  |  | Available for use | 94 | ALARM_ID_DG_PRESSURE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG pressure out of range | DG pressure out of range
|  |  |  | This alarm ID is available for use |  |  | Available for use | 95 | ALARM_ID_DG_WATCHDOG_EXPIRED | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG watchdog expired | DG watchdog expired
|  |  |  | This alarm ID is available for use |  |  | Available for use | 96 | ALARM_ID_INLET_WATER_TEMPERATURE_IN_HIGH_RANGE | ALARM_PRIORITY_MEDIUM | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Inlet water temperature in high range | Inlet water temperature in high range
|  |  |  | This alarm ID is available for use |  |  | Available for use | 97 | ALARM_ID_FILL_CONDUCTIVITY_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | Fill conductivity out of range | Fill conductivity out of range
|  |  | HD | HD battery communication fault | Battery Error | Call service to schedule a repair. | If battery I2C communication fault persistence count > 5 | 98 | ALARM_ID_HD_BATTERY_COMM_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD battery communication fault | HD battery communication fault
|  |  | HD | HD syringe pump stall alarm | Syringe Pump Error | Call service to schedule a repair. | If syringe pump measured rate < stall speed threshold (0.05) and the ramp up has been restarted 3 times after stall has been detected | 99 | ALARM_ID_HD_SYRINGE_PUMP_STALL | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump stall alarm | HD syringe pump stall alarm
|  |  | HD | HD no cartridge self-test timeout | Self Test Error | Call service to schedule a repair. | If during the no cartridge self-tests state, the elapsed self-test time > maximum no cartridge self test time (30secs) | 100 | ALARM_ID_HD_NO_CART_SELF_TEST_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD no cartridge self-test timeout | HD no cartridge self-test timeout
|  |  | HD | HD dry self-test timeout | Self Test Error | Call service to schedule a repair. | If the dry self test exceeded 180 seconds. | 101 | ALARM_ID_HD_DRY_SELF_TEST_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD dry self-test timeout | HD dry self-test timeout
|  |  | ??? | Real-time clock communication error |  |  | This alarm is not used in the DG/HD code | 102 | ALARM_ID_DG_HEAT_DISINFECT_TEMP_GRAD_OUT_OF_RANAGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG heat disinfect temperature gradient out of range | DG heat disinfect temperature gradient out of range
|  |  | HD | HD non-volatile calibration group invalid record CRC | Machine Error | Call service to schedule a repair. | If the computed CRC of the calibration group record is not equal to the record CRC | 103 | ALARM_ID_HD_NVDATAMGMT_CAL_GROUP_RECORD_CRC_INVALID | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD non-volatile calibration group invalid record CRC | HD non-volatile calibration group invalid record CRC
|  |  | HD | Air trap level sensors reporting illegal combination of air/fluid | Air Trap Error | Call service to schedule a repair. | If the air trap level readings are not valid for more than 2 secs | 104 | ALARM_ID_AIR_TRAP_ILLEGAL_LEVELS | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Air trap level sensors reporting illegal combination of air/fluid | Air trap level sensors reporting illegal combination of air/fluid
|  |  | DG | DG non-volatile system group invalid record CRC | Machine Error | Call service to schedule a repair. | If the DG NVDATAMGMT CAL group calculated CRC does not mached the expected CRC | 105 | ALARM_ID_DG_NVDATAMGMT_CAL_GROUP_RECORD_CRC_INVALID | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG non-volatile system group invalid record CRC | DG non-volatile system group invalid record CRC
|  |  |  | This alarm ID is available for use |  |  | Available for use | 106 | ALARM_ID_DG_DIALYSATE_FLOW_SENSOR_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG dialysate flow sensor invalid calibration record | DG inlet water conductivity is greater than threshold
|  |  | HD | HD reports DG restarted fault | Machine Error | Call service to schedule a repair. | If DG restart command has been sent and the DG is not in the expected mode (Standby and Standby Idle) | 107 | ALARM_ID_DG_RESTARTED_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD reports DG restarted fault | HD reports DG restarted fault
|  |  | HD | HD syringe pump ADC error | Syringe Pump Error | Call service to schedule a repair. | If the ADC read is not fresh after POST init and the ADC read counter has not changed | 108 | ALARM_ID_HD_SYRINGE_PUMP_ADC_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump ADC error | HD syringe pump ADC error
|  |  | HD | HD syringe pump volume check error | Syringe Pump Error | Call service to schedule a repair. | If the volume delivered is off by more than 5% or 0.1mL | 109 | ALARM_ID_HD_SYRINGE_PUMP_VOLUME_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump volume check error | HD syringe pump volume check error
|  |  | HD | HD syringe pump speed check error | Syringe Pump Error | Call service to schedule a repair. | If the syringe measured rate is off by more than 5% or 0.1mL | 110 | ALARM_ID_HD_SYRINGE_PUMP_SPEED_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump speed check error | HD syringe pump speed check error
|  |  | HD | HD syringe pump not stopped in off state error | Syringe Pump Error | Call service to schedule a repair. | If the syringe pump position is not equal to its last position | 111 | ALARM_ID_HD_SYRINGE_PUMP_NOT_STOPPED_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump not stopped in off state error | HD syringe pump not stopped in off state error
|  |  | HD | HD blood leak detector fault | Blood Leak Detector Error | Call service to schedule a repair. | If the blood leak zeroing state exceeed 500 msec | 112 | ALARM_ID_HD_BLOOD_LEAK_ZERO_CMD_SEND_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood leak zero command transmit timeout. | HD blood leak detector fault
|  |  | HD | HD arterial air bubble detector self-test failure |  |  | This alarm is not used in the HD code | 113 | ALARM_ID_HD_BLOOD_LEAK_ZERO_CMD_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood leak zero command fault. | Unused - alarm ID is available
|  |  | HD | HD venous air bubble detector self-test failure | Air Bubble Detector Error | Call service to schedule a repair. | If the venous air bubble self test did not detect bubble for more than 500 msec | 114 | ALARM_ID_HD_VENOUS_BUBBLE_SELF_TEST_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD venous air bubble detector self-test failure | HD venous air bubble detector self-test failure
|  |  | DG | DG temperature sensor out of range | Dialysate Temperature Error | Call service to schedule a repair. | If the measured temperature is out of range | 115 | ALARM_ID_DG_TEMPERATURE_SENSOR_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG temperature sensor out of range | DG temperature sensor out of range
|  |  | DG | DG temperature sensor ADC out of range | Dialysate Temperature Error | Call service to schedule a repair. | If the ADC value of the sensor is not out of range | 116 | ALARM_ID_DG_TEMPERATURE_SENSORS_ADC_FAULT | ALARM_PRIORITY_HIGH | 3 | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG temperature sensor ADC fault | DG temperature sensor ADC out of range
|  |  | DG | DG primary heater internal temperature out of range | Dialysate Temperature Error | Call service to schedule a repair. | If the primary heater's internal temperature or the cold junction temperature is above the allowed limit | 117 | ALARM_ID_DG_CAN_MESSAGE_NOT_ACKED | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG CAN message not acked | DG CAN message not acked
|  |  | DG | DG primary heater cold junction temperature out of range | Dialysate Temperature Error | Call service to schedule a repair. | If the primary heater's internal temperature or the cold junction temperature > 170.0 Deg C | 118 | ALARM_ID_DG_RTC_CONFIG_ERROR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | DG RTC config error | DG RTC config error
|  |  | DG | DG trimmer heater internal temperature out of range | Dialysate Temperature Error | Call service to schedule a repair. | If the primary heater's internal temperature or the cold junction temperature is above the allowed limit | 119 | ALARM_ID_DG_RTC_BATTERY_LOW | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | DG RTC battery low error | DG RTC battery low error
|  |  | DG | DG trimmer heater cold junction temperature out of range | Dialysate Temperature Error | Call service to schedule a repair. | If the primary heater's internal temperature or the cold junction temperature > 170.0 Deg C | 120 | ALARM_ID_PRE_TREATMENT_WET_PRIME_TEST_FAILURE | ALARM_PRIORITY_LOW | 910 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD pre-treatment mode wet self-test prime check failure | HD pre-treatment mode wet self-test prime check failure
|  |  | DG | DG main primary heater voltage out of range | Dialysate Temperature Error | Call service to schedule a repair. | If the main primary heater's voltage is out of range | 121 | ALARM_ID_DG_MAIN_PRIMARY_HEATER_VOLTAGE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG main primary heater voltage out of range | DG main primary heater voltage out of range
|  |  | DG | DG small primary heater voltage out of range | Dialysate Temperature Error | Call service to schedule a repair. | If the small primary heater's voltage is out of range | 122 | ALARM_ID_DG_SMALL_PRIMARY_HEATER_VOLTAGE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG small primary heater voltage out of range | DG small primary heater voltage out of range
|  |  | DG | DG trimmer heater voltage out of range | Dialysate Temperature Error | Call service to schedule a repair. | If the trimmer voltage is out of range | 123 | ALARM_ID_DG_TRIMMER_HEATER_VOLTAGE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG trimmer heater voltage out of range | DG trimmer heater voltage out of range
|  |  | HD | HD end of treatment alarm (high priority - user still not acting to end treatment) | Blood Stagnant | Blood has sat too long in the blood tubing set. No rinseback allowed. | In treatment end paused state for > 5 minutes | 124 | ALARM_ID_END_OF_TREATMENT_HIGH | ALARM_PRIORITY_HIGH | 4 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD end of treatment alarm (high priority - user still not acting to end treatment) | HD end of treatment alarm (high priority - user still not acting to end treatment)
|  |  | HD | Blood sitting too long after treatment stopped by user alarm (> 5 min) | Blood Stagnant | Blood has sat too long in the blood tubing set. Treatment must end. | If in treatment stop sub mode when blood is in the tubing, the blood sitting counter exceeded 5 minutes | 125 | ALARM_ID_TREATMENT_STOPPED_NO_RINSEBACK | ALARM_PRIORITY_HIGH | 5 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood sitting too long after treatment stopped by user alarm (>5 min) | Blood sitting too long after treatment stopped by user alarm (>5 min)
|  |  | HD | HD blood leak detected alarm | Dialyzer Blood Leak Detected | A blood leak has been detected. Treatment must end. | If the blood leak is detected based on the blood leak status reported by the sensor. | 126 | ALARM_ID_HD_BLOOD_LEAK_DETECTED | ALARM_PRIORITY_HIGH | 6 | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | HD blood leak detected alarm | HD blood leak detected alarm
|  |  | HD | Venous pressure too low during treatment | Venous Pressure Low | Venous pressure is low. Check to be sure your blue venous connections are secure and that your venous access is intact. | If the venous pressure < TREATMENT_PARAM_VEN_PRESSURE_LOW_LIMIT | 127 | ALARM_ID_VENOUS_PRESSURE_LOW | ALARM_PRIORITY_HIGH | 8 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Venous pressure too low during treatment | Venous pressure too low during treatment
|  |  | HD | HD venous air bubble detected alarm | Venous Air Bubble Detected | Venous air bubble detected. Check for air bubbles in the venous line. If bubbles are present, tap the tubing to move the bubbles up to the air trap. If bubbles cannot be removed, treatment must end. | If bubbles are detected during the treatment. | 128 | ALARM_ID_HD_VENOUS_BUBBLE_DETECTED | ALARM_PRIORITY_HIGH | 9 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | HD venous air bubble detected alarm | HD venous air bubble detected alarm
|  |  | HD | HD venous air bubble detected rinseback alarm | Venous Air Bubble Detected During Rinseback | Venous air bubble detected. Check for air bubbles in the venous line. If bubbles are present, tap the tubing to move the bubbles up to the air trap. If bubbles cannot be removed, treatment must end. | If bubble is detected in the treatment rinseback state | 129 | ALARM_ID_HD_VENOUS_BUBBLE_DETECTED_RINSEBACK | ALARM_PRIORITY_HIGH | 10 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | HD venous air bubble detected rinseback alarm | HD venous air bubble detected rinseback alarm
|  |  | HD | Venous pressure too high during treatment | Venous Pressure High | Venous pressure is high. Check to be sure your blue venous connections are secure and that there are no kinks or clots in the tubing. | If the venous pressure > TREATMENT_PARAM_VEN_PRESSURE_HIGH_LIMIT | 130 | ALARM_ID_VENOUS_PRESSURE_HIGH | ALARM_PRIORITY_HIGH | 11 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Venous pressure too high during treatment | Venous pressure too high during treatment
|  |  | HD | Arterial pressure too low during treatment | Arterial Pressure Low | Arterial pressure is low. Check to be sure your red arterial connections are secure and that there are no kinks or clots in the tubing. | If the arterial pressure < TREATMENT_PARAM_ART_PRESSURE_LOW_LIMIT | 131 | ALARM_ID_ARTERIAL_PRESSURE_LOW | ALARM_PRIORITY_HIGH | 12 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Arterial pressure too low during treatment | Arterial pressure too low during treatment
|  |  | HD | Arterial pressure too high during treatment | Arterial Pressure High | Arterial pressure is low. Check to be sure your red arterial connections are secure and that your venous access is intact. | If the arterial pressure > TREATMENT_PARAM_ART_PRESSURE_HIGH_LIMIT | 132 | ALARM_ID_ARTERIAL_PRESSURE_HIGH | ALARM_PRIORITY_HIGH | 13 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Arterial pressure too high during treatment | Arterial pressure too high during treatment
|  |  | DG | DG fluid leak detected alarm | Fluid Leak in Dialysate Generator | Fluid leak detected in Dialysate Generation | If the fluid leak is detected as reported by the fluid leak detector. | 133 | ALARM_ID_DG_FLUID_LEAK_DETECTED | ALARM_PRIORITY_HIGH | 14 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | TRUE | TRUE | 0 | ALARM_ID_NO_ALARM | DG fluid leak detected alarm | DG fluid leak detected alarm
|  |  | HD | HD fluid leak detected alarm | Fluid Leak in Hemodialysis Delivery Device | Fluid leak detected in Hemodialysis Delivery device. Clean up the spill before continuing. | If the fluid leak is detected as reported by the fluid leak detector. | 134 | ALARM_ID_HD_FLUID_LEAK_DETECTED | ALARM_PRIORITY_HIGH | 15 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | HD fluid leak detected alarm | HD fluid leak detected alarm
|  |  | HD | HD experienced large acceleration | Excessive Acceleration Detected in the Hemodialysis Device | Excessive movement detected with Hemodialysis Delivery device. Confirm the device is stationary and level before continuing. | If the system experienced excessive shock > 2.5 g. | 135 | ALARM_ID_HD_SHOCK | ALARM_PRIORITY_HIGH | 16 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD experienced large acceleration | HD experienced large acceleration
|  |  | DG | DG experienced large acceleration | Excessive Acceleration Detected in the Dialysate Generation Device | Excessive movement detected with Dialysis Generation device. Confirm the device is stationary and level before continuing. | If the system experienced excessive shock > 2.5 g. | 136 | ALARM_ID_DG_SHOCK | ALARM_PRIORITY_HIGH | 16 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG experienced large acceleration | DG experienced large acceleration
|  |  | HD | HD tilt exceeds maximum | Excessive Tilt Detected in the Hemodialysis Device | Excessive tilt detected with Hemodialysis Delivery device. Confirm the device is stationary and level before continuing. | If the system experienced excessive tilt x and y axis angle > 7 deg for more than 1 sec | 137 | ALARM_ID_HD_EXCESSIVE_TILT | ALARM_PRIORITY_HIGH | 17 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD tilt exceeds maximum | HD tilt exceeds maximum
|  |  | DG | DG tilt exceeds maximum | Excessive Tilt Detected in the Dialysate Generation Device | Excessive tilt detected with Dialysis Generation device. Confirm the device is stationary and level before continuing. | If the system experienced excessive tilt x and y axis angle > 7 deg for more than 1 sec | 138 | ALARM_ID_DG_EXCESSIVE_TILT | ALARM_PRIORITY_HIGH | 17 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG tilt exceeds maximum | DG tilt exceeds maximum
|  |  | HD | HD AC power lost alarm | Power Lost | Power has been lost. Check all electrical connections. The display and CPU will run off of the battery until power returns. | If AC Power is lost for more 2.2 seconds. | 139 | ALARM_ID_HD_AC_POWER_LOST | ALARM_PRIORITY_HIGH | 18 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | HD AC power lost alarm | HD AC power lost alarm
|  |  | HD | Dialysate generator communication timeout | Communication Failure with Dialysate Generator | Communication failure with Dialysis Generator. Ensure all cords are securely connected. | If DG communication timeout is detected for more than 2 seconds. | 140 | ALARM_ID_DG_COMM_TIMEOUT | ALARM_PRIORITY_HIGH | 19 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Dialysate generator communication timeout | Dialysate generator communication timeout
|  |  | HD | Air trap fill timeout during treatment | Air Trap Unable to Fill | Failure to fill the Air Trap. Check the air trap vent line and the transducer protector. If the transducer protector is wet, replace it. | If the air trap fill timeout during treatment | 141 | ALARM_ID_AIR_TRAP_FILL_DURING_TREATMENT | ALARM_PRIORITY_HIGH | 20 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Air trap fill timeout during treatment | Air trap fill timeout during treatment
|  |  | HD | Blood pump occlusion detected | Blood Pump Occlusion Detected | Occlusion detected after the blood pump. Check for clots or kinks around the dialyzer and in the venous line. | If blood occlusion is detected > OCCLUSION_THRESHOLD_OFFSET + bloodPumpOcclusionAfterCartridgeInstall | 142 | ALARM_ID_OCCLUSION_BLOOD_PUMP | ALARM_PRIORITY_HIGH | 21 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Blood pump occlusion detected | Blood pump occlusion detected
|  |  | DG | DG dialysate temperature sensors drift timeout. | Dialysate Temperature Error | Call service to schedule a repair. | If the dialysate temp sensors have drifted for more than 10 secs | 143 | ALARM_ID_DG_DIALYSATE_TEMPERATURE_SENSORS_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | DG dialysate temperature sensors out of range | DG dialysate temperature sensors out of range
|  |  |  | This alarm ID is available for use |  |  | Available for use | 144 | ALARM_ID_DG_RO_FLOW_SENSOR_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG RO flow sensor invalid calibration record | Alarm ID available for use
|  |  | DG | DG Concentrate conductivity after adding acid out of range alarm | Acid Concentrate Conductivity Out of Range | Dialysate conductivity is out of range. Check the acid dialysate concentrates and make sure the correct concentrate is connected. | If the conductivity value for the acid concentrate is out of range for more than 45 secs | 145 | ALARM_ID_ACID_CONDUCTIVITY_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Concentrate conductivity after adding acid out of range alarm | Concentrate conductivity after adding acid out of range alarm
|  |  |  | Dialysate conductivity out of range |  |  | This alarm ID is not used in the DG code. | 146 | ALARM_ID_DG_RTC_OR_TIMER_ACCURACY_FAILURE | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | DG RTC or timer accuracy failure. | DG RTC or timer accuracy failure.
|  |  | DG | DG dialysate generation conductivity fault alarm | Dialysate Conductivity Error | Call service to schedule a repair. | If dialysate conductivity > 14850 or < 13000 | 147 | ALARM_ID_CREATING_DIALYSATE_PLEASE_WAIT | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Wait for the DG to produce dialysate | Wait for the DG to produce dialysate
|  |  | DG | Inlet water temperature too high | Inlet Water Temperature High | Inlet water temperature is high. Check the incoming water for issues with cold water supply. | Inlet temperature > 35 Deg C | 148 | ALARM_ID_INLET_WATER_HIGH_TEMPERATURE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Inlet water temperature too high | Inlet water temperature too high
|  |  | DG | Inlet water temperature too low | Inlet Water Temperature Low | Inlet water temperature is low. Check the incoming water for issues with hot water supply. | Inlet temperature < 22 Deg C | 149 | ALARM_ID_HD_WATCHDOG_EXPIRED | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD watchdog expired | HD watchdog expired
|  |  | DG | Inlet water conductivity too high | Inlet Water Conductivity High | Inlet water conductivity is high. Check the incoming water for issues. | Inlet water conductivity > 2000 | 150 | ALARM_ID_INLET_WATER_CONDUCTIVITY_IN_HIGH_RANGE | ALARM_PRIORITY_MEDIUM | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Inlet water conductivity in high range | Inlet water conductivity in high range
|  |  | DG | Inlet water conductivity too low | Inlet Water Conductivity Low | Inlet water conductivity is low. Check the incoming water for issues. | Inlet water conductivity< 100 | 151 | ALARM_ID_INLET_WATER_CONDUCTIVITY_IN_LOW_RANGE | ALARM_PRIORITY_MEDIUM | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Inlet water conductivity in low range | Inlet water conductivity in low range
|  |  | DG | Inlet water pressure too low | Inlet Water Pressure Low | Inlet water pressure is low. Check the incoming water supply. | Inlet water pressure < 25 PSI? | 152 | ALARM_ID_INLET_WATER_PRESSURE_IN_LOW_RANGE | ALARM_PRIORITY_MEDIUM | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Inlet water pressure in low range | Inlet water pressure in low range
|  |  | HD | HD prime completed high priority alarm |  |  | This alarm is not used in the HD/DG code | 153 | ALARM_ID_PRIME_COMPLETED_HIGH | ALARM_PRIORITY_HIGH | 36 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD prime completed high priority alarm | HD prime completed high priority alarm
|  |  | ??? | EEPROM operations (read write erase) failure write erase) failure | Communication Failure | Communication has been lost. Ensure all cords are securely connected. | If the RTC RAM or the EEPROM has timeout | 154 | ALARM_ID_NVDATA_EEPROM_OPS_FAILURE | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | EEPROM operations (read, write, erase) failure | EEPROM operations (read, write, erase) failure
|  |  |  | This alarm ID is available for use |  |  | Available for use | 155 | ALARM_ID_DG_TEMPERATURE_SENSOR_ADC_OUT_OF_RANGE | ALARM_PRIORITY_LOW | 799 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Alarm ID available for use | Alarm ID available for use
|  |  | ??? | HW usage data (treatment time in HD and total consumed water in DG) failure |  |  | This alarm is not used in the DG/HD code | 156 | ALARM_ID_DG_TEMPERATURE_SENSOR_FAULT | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Alarm ID available for use | HW usage data (treatment time in HD and total consumed water in DG) failure
|  |  |  | This alarm ID is available for use |  |  | Available for use | 157 | ALARM_ID_DG_CHEM_DISINFECT_PRIME_ACID_LINE_TIME_OUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG chemical disinfect prime acid line timeout. | DG chemical disinfect prime acid line timeout.
|  |  | HD | Blood pump failed flow vs. motor speed check. |  |  | This alarm ID is not used in the HD code. | 158 | ALARM_ID_INLET_WATER_TEMPERATURE_IN_LOW_RANGE | ALARM_PRIORITY_MEDIUM | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | 0 | ALARM_ID_NO_ALARM | Inlet water temperature is in the low range | Inlet water temperature is in the low range
|  |  | HD | Dialysate inlet pump failed flow vs. motor speed check. | Dialysate Pump Failure | Dialysate pump failure. Check there are no kinks in the dialysate tubing. | If dialysate in pump max flow rate versus pump speed mismatched is detected for 5 seconds. | 159 | ALARM_ID_DIAL_IN_PUMP_FLOW_VS_MOTOR_SPEED_CHECK | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysate inlet pump failed flow vs. motor speed check. | Dialysate inlet pump failed flow vs. motor speed check.
|  |  | HD | Blood pump rotor speed too high | Blood Pump Speed High | Blood pump failure. Check the pump rotor for secure connection. | If blood pump motor speed exceeded 100.0 RPM. | 160 | ALARM_ID_BLOOD_PUMP_ROTOR_SPEED_TOO_HIGH | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood pump rotor speed too high | Blood pump rotor speed too high
|  |  | HD | Blood flow sensor signal strength too low TODO remove this alarm when FMB fully removed |  |  | This alarm is not used in the HD code | 161 | ALARM_ID_DG_COMM_TOO_MANY_BAD_CRCS | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG comm too many bad CRCs | DG comm too many bad CRCs
|  |  | HD | Dialysate flow sensor signal strength too low TODO remove this alarm when FMD fully removed |  |  | This alarm is not used in the HD/DG code | 162 | ALARM_ID_DG_FPGA_CLOCK_SPEED_CHECK_FAILURE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG FPGA clock speed check failure. | DG FPGA clock speed check failure.
|  |  | HD | HD sees primary load cell for reservoir 1 change too much too fast | Ultrafiltration Measurement Error | Ultrafiltration error. | If the delta volume of the active reservoir exceeded 50.0 (MAX_ACTIVE_LOAD_CELL_CHANGE_G) | 163 | ALARM_ID_HD_LOAD_CELL_ACCELERATION_RES_1_ALARM | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD sees primary load cell for reservoir 1 change too much too fast | HD sees primary load cell for reservoir 1 change too much too fast
|  |  | HD | HD sees primary load cell for reservoir 2 change too much too fast | Ultrafiltration Measurement Error | Ultrafiltration error. | If the delta volume of the active reservoir exceeded 50.0 (MAX_ACTIVE_LOAD_CELL_CHANGE_G) | 164 | ALARM_ID_HD_LOAD_CELL_ACCELERATION_RES_2_ALARM | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD sees primary load cell for reservoir 2 change too much too fast | HD sees primary load cell for reservoir 2 change too much too fast
|  |  | HD | HD in treatment rinseback operation timeout | Rinseback Timeout | Maximum Rinseback time exceeded. Rinseback must end. | If the maximum time inrinseback operation has been exceeded | 165 | ALARM_ID_TREATMENT_RINSEBACK_TIMEOUT_ALARM | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD in treatment rinseback operation timeout | HD in treatment rinseback operation timeout
|  |  | HD | HD in treatment re-circ sub-mode for too long | Recirculation Timeout | Maximum Recirculation time exceeded. Treatment must end. | If the maximum time in recirculate sub-mode has been exceeded | 166 | ALARM_ID_TREATMENT_RECIRC_TIMEOUT_ALARM | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD in treatment re-circ sub-mode for too long | HD in treatment re-circ sub-mode for too long
|  |  | HD | HD cartridge door opened alarm | Door Open | Hemodialysis Delivery device door open. | If the cartridge door is opened as reported by the FPGA | 167 | ALARM_ID_CARTRIDGE_DOOR_OPENED | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD cartridge door opened alarm | HD cartridge door opened alarm
|  |  | DG | HD active reservoir depletion timeout. | Dialysate Recirculation Error | Dialysate recirculation error. | If the active reservoir's depletion time has elapsed | 168 | ALARM_ID_HD_ACTIVE_RESERVOIR_DEPLETION_TIME_OUT | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD active reservoir depletion timeout. | HD active reservoir depletion timeout.
|  |  | DG | Dialysate flow rate out of range | Dialysate Generation Error | Call service to schedule a repair. | If the dialysate flow rate > 2.00 LPM | 169 | ALARM_ID_DIALYSATE_FLOW_RATE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysate flow rate out of range | Dialysate flow rate out of range
|  |  | HD | HD syringe empty alarm | Syringe Empty | Heparin syringe empty. Option to resume treatment without heparin. | If the syringe pump position is near empty position | 170 | ALARM_ID_HD_SYRINGE_PUMP_SYRINGE_EMPTY | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe empty alarm | HD syringe empty alarm
|  |  | HD | HD syringe pump occlusion alarm | Heparin Occlusion | Heparin pump occlusion detected. Check heparin line is unclamped and there are no kinks in the heparin line. | If the pressure sensor voltage reading > SYRINGE_FORCE_OCCLUSION_THRESHOLD_V | 171 | ALARM_ID_HD_SYRINGE_PUMP_OCCLUSION | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump occlusion alarm | HD syringe pump occlusion alarm
|  |  | HD | HD syringe pump not enough Heparin alarm | Not Enough Heparin | Not enough heparin is available for treatment. Check if syringe pump is touching syringe plunger. If syringe pump is touching syringe plunger, end treatment. If syringe pump is not touching syringe plunger, reload syringe with enough heparin. | If the estimated syringe volume needed for treatment is not sufficient | 172 | ALARM_ID_HD_SYRINGE_PUMP_NOT_ENOUGH_HEPARIN_ALARM | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump not enough Heparin alarm | HD syringe pump not enough Heparin alarm
|  |  | HD | HD arterial air bubble detected alarm |  |  | This alarm is not used in the DG or HD code | 173 | ALARM_ID_HD_RTC_CONFIG_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD RTC configuration error | HD RTC configuration error
|  |  | HD | HD arterial air bubble detected rinseback alarm |  |  | This alarm is not used in the DG or HD code | 174 | ALARM_ID_HD_RTC_OR_TIMER_ACCURACY_FAILURE | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD RTC or timer accuracy failure | HD RTC or timer accuracy failure
|  |  | HD | HD pump direction status error | Blood Pump Error | Blood pump eailure. Check the pump rotor for secure connection. | If the set pump direction does not agree with the hall sensor direction. | 175 | ALARM_ID_HD_PUMP_DIRECTION_STATUS_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD pump direction status error | HD pump direction status error
|  |  | HD | // TODO remove this alarm HD active reservoir is full during treatment and DG not ready to switch reservoirs. |  |  | This alarm ID is not used in the HD code. | 176 | ALARM_ID_HD_RESERVOIR_FULL_AND_DG_NOT_READY_TO_SWITCH | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD active reservoir is full during treatment and DG not ready to switch reservoirs. // | HD active reservoir is full during treatment and DG not ready to switch reservoirs. //
|  |  | DG | DG software fault. Software found itself in an unexpected state | Machine Error | Call service to schedule a repair. | If the specific software faults listed in alarm management SW faults has been detected | 177 | ALARM_ID_DG_SOFTWARE_FAULT | ALARM_PRIORITY_HIGH | 799 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG software fault. Software found itself in an unexpected state | DG software fault. Software found itself in an unexpected state
|  |  | DG | HD communication timeout |  |  | This alarm ID is not used in the HD code. | 178 | ALARM_ID_HD_COMM_TIMEOUT | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD communication timeout | HD communication timeout
|  |  | DG | DG FPGA communication down for too long | Communication Failure | Call service to schedule a repair. | If retries for commands exceeds limit or FPGA reports comm error | 179 | ALARM_ID_DG_FPGA_COMM_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG FPGA communication down for too long | DG FPGA communication down for too long
|  |  | DG | DG load cell ADC error | Ultrafiltration Measurement Error | Call service to schedule a repair. | If the MSB indicated DG load cell error | 180 | ALARM_ID_DG_LOAD_CELL_ADC_ERROR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG load cell ADC error | DG load cell ADC error
|  |  | DG | DG load cells weight out of range for tare | Ultrafiltration Measurement Error | Call service to schedule a repair. | If the load cell weight is out of range for tare | 181 | ALARM_ID_DG_LOAD_CELLS_TARE_WEIGHT_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG load cells weight out of range for tare | DG load cells weight out of range for tare
|  |  | DG | DG load cells invalid calibration | Ultrafiltration Measurement Error | Call service to schedule a repair. | If the load cells calibration data in the non-volatile memory is invalid | 182 | ALARM_ID_DG_LOAD_CELLS_INVALID_CAL_RECORD | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG load cells invalid calibration | DG load cells invalid calibration
|  |  | DG | DG invalid load cell value | Ultrafiltration Measurement Error | Call service to schedule a repair. | If reservoirs are NOT drained and flushed with the defined period of time | 183 | ALARM_ID_DG_INVALID_LOAD_CELL_VALUE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG invalid load cell value | DG invalid load cell value
|  |  | DG | Inlet or Outlet UV reactor not healthy | Machine Error | Call service to schedule a repair. | If the inlet or outlet reactor are not healthy | 184 | ALARM_ID_UV_REACTOR_NOT_HEALTHY | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Inlet or Outlet UV reactor not healthy | Inlet or Outlet UV reactor not healthy
|  |  | DG | DG fan RPM out of range | Fan Error | Fan failure. Check that the fans inlets are not blocked or obstructed. | If the fan RPM is out of Min/Max range. | 185 | ALARM_ID_DG_FAN_RPM_OUT_OF_RANGE | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | DG fan RPM out of range | DG fan RPM out of range
|  |  | DG | DG concentrate pump fault | Dialysate Concentrate Pump Error | Call service to schedule a repair. | If the concentrate pumps fault is reported by the FPGA | 186 | ALARM_ID_DG_CONCENTRATE_PUMP_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG concentrate pump fault | DG concentrate pump fault
|  |  | DG | DG concentrate pump CP1 speed control error | Dialysate Concentrate Pump Error | Call service to schedule a repair. | If the concentrate pump CP1 speed > 0.02 tolerance | 187 | ALARM_ID_CP1_SPEED_CONTROL_ERROR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG concentrate pump CP1 speed control error | DG concentrate pump CP1 speed control error
|  |  | DG | DG concentrate pump CP2 speed control error | Dialysate Concentrate Pump Error | Call service to schedule a repair. | If the concentrate pump CP1 speed > 0.02 tolerance | 188 | ALARM_ID_CP2_SPEED_CONTROL_ERROR | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG concentrate pump CP2 speed control error | DG concentrate pump CP2 speed control error
|  |  | DG | DG drain pump RPM out of range | Dialysate Drain Pump Error | Call service to schedule a repair. | If the drain pump RPM > 100.0 RMP | 189 | ALARM_ID_DRAIN_PUMP_RPM_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG drain pump RPM out of range | DG drain pump RPM out of range
|  |  | DG | DG drain pump off fault | Dialysate Drain Pump Error | Call service to schedule a repair. | If the pump is in off state and the RPM is greater than the minimum RPM of 300.0 | 190 | ALARM_ID_DRAIN_PUMP_OFF_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG drain pump off fault | DG drain pump off fault
|  |  | DG | DG flow rate out of upper range | Dialysate Flowrate Error | Call service to schedule a repair. | If the DG flow rate > upper range | 191 | ALARM_ID_FLOW_RATE_OUT_OF_UPPER_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG flow rate out of upper range | DG flow rate out of upper range
|  |  | DG | DG flow rate out of lower range | Dialysate Flowrate Error | Call service to schedule a repair. | If the DG flow rate < lower range | 192 | ALARM_ID_FLOW_RATE_OUT_OF_LOWER_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG flow rate out of lower range | DG flow rate out of lower range
|  |  | DG | RO pump flow rate out of range |  |  | This alarm is not used in the DG/HD code | 193 | ALARM_ID__AVAILABLE_13 | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Alarm available for use | RO pump flow rate out of range
|  |  | DG | DG RO pump duty cycle out of range | RO Pump Error | Call service to schedule a repair. | If the DG RO pump duty cycle is out of range | 194 | ALARM_ID_RO_PUMP_DUTY_CYCLE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG RO pump duty cycle out of range | DG RO pump duty cycle out of range
|  |  | DG | DG RO pump pressure out of range | RO Pump Error | Call service to schedule a repair. | If the DG RO pump pressure is out of range | 195 | ALARM_ID_RO_PUMP_PRESSURE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 799 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG RO pump pressure out of range | DG RO pump pressure out of range
|  |  | DG | DG temperature sensor error flag fault. | Dialysate Temperature Error | Call service to schedule a repair. | If the error flag (MSB bit of the last byte) is a logic '1' | 196 | ALARM_ID_DG_TEMPERATURE_SENSOR_ERROR_BIT_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG temperature sensor error bit fault | DG temperature sensor error flag fault.
|  |  | DG | DG temperature sensors ADC fault | Dialysate Temperature Error | Call service to schedule a repair. | If the FPGA count has changed and the ADC value is not the same as the previous read or there is FPGA error | 197 | ALARM_ID_DG_TEMPERATURE_SENSORS_FPGA_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG temperature sensors FPGA fault | DG temperature sensors ADC fault
|  |  | DG | DG heaters cold junction temperature out of range | Dialysate Heater Error | Call service to schedule a repair. | If the DG heaters cold junction temperature < 0 | 198 | ALARM_ID_RO_FLOW_TOO_LOW_WHILE_PRIMARY_HEATER_IS_ON | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | RO flow too low while primary heater is on | Alarm ID available for use
|  |  | DG | DG heaters fault |  |  | This alarm ID is not used in the DG code. | 199 | ALARM_ID_DIALYSATE_FLOW_TOO_LOW_WHILE_TRIMMER_HEATER_IS_ON | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Dialysate flow too low while trimmer heater is on | DG heaters fault
|  |  | DG | DG thermistors/sensors temperature out of range | Dialysate Temperature Error | Call service to schedule a repair. | If thermistor > 80.0 or < 0.0 deg C | 200 | ALARM_ID_DG_THERMISTORS_TEMPERATURE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG thermistors/sensors temperature out of range | DG thermistors/sensors temperature out of range
|  |  | DG | Inlet water quality is bad (temperature conductivity pressure) conductivity pressure) | Inlet Water Error | Inlet water quality is bad. | If the inlet water quality is NOT good for more than 30 secs | 201 | ALARM_ID_PRE_TREATMENT_WET_FLOW_TEST_FAILURE | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD pre-treatment mode wet self-test lc vs FMD failure | HD pre-treatment mode wet self-test lc vs FMD failure
|  |  | DG | Inlet water pressure fault | Inlet Water Error | Inlet water quality is bad. | If the inlet pressure | 202 | ALARM_ID_PRE_TREATMENT_DRY_PRESSURE_NORMAL_TEST_FAILURE | ALARM_PRIORITY_LOW | 910 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD pre-treatment mode dry pressure normal self-test failure | HD pre-treatment mode dry pressure normal self-test failure
|  |  | DG | Pressure sensor fault | Dialysate Pressure Error | Call service to schedule a repair. | If inlet pressure ADC reading = 0 or inlet pressure ADC reading > 4095 (12-bit full scale ) | 203 | ALARM_ID_PRESSURE_SENSOR_FAULT | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Pressure sensor fault | Pressure sensor fault
|  |  | DG | RO rejection ratio out of range | RO Rejection Error | Call service to schedule a repair. | If RO rejection ratio > 0.1 and outlet water conductivity > 15.0 | 204 | ALARM_ID_RO_REJECTION_RATIO_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | RO rejection ratio out of range | RO rejection ratio out of range
|  |  | DG | Conductivity sensor fault | Dialysate Conductivity Error | Call service to schedule a repair. | If conductivity sensor error count reported bythe FPGA > 0 for more than | 205 | ALARM_ID_CONDUCTIVITY_SENSOR_FAULT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Conductivity sensor fault | Conductivity sensor fault
|  |  | DG | DG dialysate fill runs out of time | Dialysate Production Error | Call service to schedule a repair. | If DG dialysate fill > 5 minutes | 206 | ALARM_ID_DG_DIALYSATE_FILL_OUT_OF_TIME | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG dialysate fill runs out of time | DG dialysate fill runs out of time
|  |  | DG | DG flow meter check failure alarm | Dialysate Flowrate Error | Call service to schedule a repair. | If integrated Volume To Load Cell Reading Percent > 0.1 | 207 | ALARM_ID_DG_FLOW_METER_CHECK_FAILURE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG flow meter check failure alarm | DG flow meter check failure alarm
|  |  | DG | DG primary heater on with no flow time out | Dialysate Temperature Error | Call service to schedule a repair. | If heater on with no flow for more than 120 secs | 208 | ALARM_ID_HD_VENOUS_LINE_OCCLUSION | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD venous pressure occlusion alarm | Alarm ID available for use
|  |  | DG | DG drain circulation line timeout | Dialysate Production Error | Call service to schedule a repair. | If the integrated volume < 0.1 for more than 60 secs | 209 | ALARM_ID_DG_DRAIN_CIRCULATION_LINE_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG drain circulation line timeout | DG drain circulation line timeout
|  |  | HD | HD battery pack detected an error | Battery Error | Battery pack error. | If the battery pack status error is detected from the slave device I2C address 0x0B | 210 | ALARM_ID_HD_BATTERY_PACK_ERROR_DETECTED | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD battery pack detected an error | HD battery pack detected an error
|  |  | HD | Blood sitting too long warning (> 4 min) | Blood Stagnant | Maximum time for blood stoppage has been exceeded. | During treatment stop sub-mode, if the blood is in the line for more than 4 minutes | 211 | ALARM_ID_BLOOD_SITTING_WARNING | ALARM_PRIORITY_MEDIUM | 800 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Blood sitting too long warning (>4 min) | Blood sitting too long warning (>4 min)
|  |  | HD | Blood sitting too long warning (> 4 min) | Blood Stagnant | Maximum time for blood stoppage has been exceeded. | During treatment end paused state, if the blood is in the line for more than 4 minutes | 212 | ALARM_ID_END_OF_TREATMENT_ALARM | ALARM_PRIORITY_MEDIUM | 801 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD end of treatment alarm (user not acting to end treatment) | HD end of treatment alarm (user not acting to end treatment)
|  |  | HD | HD prime completed medium priority alarm |  |  | This alarm is not used in the HD code | 213 | ALARM_ID_PRIME_COMPLETED_MEDIUM | ALARM_PRIORITY_MEDIUM | 802 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | ALM_ESC_30_MIN | ALARM_ID_PRIME_COMPLETED_HIGH | HD prime completed medium priority alarm | HD prime completed medium priority alarm
|  |  | HD | HD Saline bolus volume check failure | Saline Bolus Error | Inability to deliver saline bolus. Check for kinks or clamped lines in the saline line. | If the saline bolus is under or over delivering | 214 | ALARM_ID_SALINE_BOLUS_VOLUME_CHECK_FAILURE | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD Saline bolus volume check failure | HD Saline bolus volume check failure
|  |  | HD | HD rinseback volume check failure | Rinseback Error | Unable to complete Rinseback. Check for kinks or clamped blood lines. | If the delta between the expected rinseback volume and the measured rinseback volume > 60.0 mL | 215 | ALARM_ID_RINSEBACK_VOLUME_CHECK_FAILURE | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD rinseback volume check failure | HD rinseback volume check failure
|  |  | HD | HD end treatment sub-mode timeout alarm | Treatment Ended | Treatment has ended |  | 216 | ALARM_ID_END_TREATMENT_TIMEOUT_ALARM | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD end treatment sub-mode timeout alarm | HD end treatment sub-mode timeout alarm
|  |  | HD | HD blood prime volume check failure | Prime Error | Blood prime error. Check for kinks or clamped blood lines. | If expected blood prime volume - blood prime volume > 60.0 mL. | 217 | ALARM_ID_BLOOD_PRIME_VOLUME_CHECK_FAILURE | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD blood prime volume check failure | HD blood prime volume check failure
|  |  | HD | HD syringe detected alarm | Syringe Not Detected | Heparin syringe not detected. | If the syringe detection switch > 2.0 volts. | 218 | ALARM_ID_HD_SYRINGE_DETECTED | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe detected alarm | HD syringe detected alarm
|  |  | HD | HD syringe pump syringe removed alarm | Syringe Removed | Heparin syringe removed. | If the syringe detection switch < 2.0 volts. | 219 | ALARM_ID_HD_SYRINGE_PUMP_SYRINGE_REMOVED | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD syringe pump syringe removed alarm | HD syringe pump syringe removed alarm
|  |  | HD | HD battery pack SOC is too low. | Battery Error | Battery pack error. | If the battery remaining charge < 50%. | 220 | ALARM_ID_HD_BATTERY_PACK_CHARGE_TOO_LOW | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD battery pack SOC is too low. | HD battery pack SOC is too low.
|  |  | HD | HD Saline bag is empty | Saline Bag Empty | Saline back appears empty. Please check saline bag and line and replace if saline bag is empty. | If the measured arterial pressure < -300.0 mmHg for more than 250 msec. | 221 | ALARM_ID_EMPTY_SALINE_BAG | ALARM_PRIORITY_LOW | 900 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD Saline bag is empty | HD Saline bag is empty
|  |  | HD | HD dialysate temperature too high alarm | Dialysate Temperature High | Dialysate temperature high. | If the dialysate temperature > 42.0 deg C. | 222 | ALARM_ID_DIALYSATE_TEMPERATURE_HIGH | ALARM_PRIORITY_LOW | 901 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | ALM_ESC_5_MIN | ALARM_ID_TREATMENT_STOPPED_NO_RINSEBACK | HD dialysate temperature too high alarm | HD dialysate temperature too high alarm
|  |  | HD | HD dialysate temperature too low alarm | Dialysate Temperature Low | Dialysate temperature low. | If the dialysate temperature < 33.0 deg C. | 223 | ALARM_ID_DIALYSATE_TEMPERATURE_LOW | ALARM_PRIORITY_LOW | 902 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | ALM_ESC_5_MIN | ALARM_ID_TREATMENT_STOPPED_NO_RINSEBACK | HD dialysate temperature too low alarm | HD dialysate temperature too low alarm
|  |  | HD | Treatment stopped by user action - pressed stop button | Treatment Stopped | Treatment paused. | If the user has pressed the stop button | 224 | ALARM_ID_TREATMENT_STOPPED_BY_USER | ALARM_PRIORITY_LOW | 903 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Treatment stopped by user action - pressed stop button | Treatment stopped by user action - pressed stop button
|  |  | HD | HD end of treatment warning | End of Treatment | Treatment has ended | If the treatment duration has exceeded the prescribed treatment time entered by the user. | 225 | ALARM_ID_END_OF_TREATMENT_WARNING | ALARM_PRIORITY_LOW | 904 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD end of treatment warning | HD end of treatment warning
|  |  | HD | HD prime completed low priority alarm | Prime Complete | Prime of the bloodlines is complete. | If the submodeCompleteTransitionTimeCounter > 20 | 226 | ALARM_ID_PRIME_COMPLETED_LOW_PRIORITY | ALARM_PRIORITY_LOW | 905 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | ALM_ESC_90_MIN | ALARM_ID_PRIME_COMPLETED_MEDIUM | HD prime completed low priority alarm | HD prime completed low priority alarm
|  |  | HD | HD prime out of time alarm | Maximum Prime Time Exceeded | Maximum time for blood prime has been exceeded. | If the priming elapsed time exceeded 10 minutes. | 227 | ALARM_ID_PRIME_OUT_OF_TIME | ALARM_PRIORITY_LOW | 909 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD prime out of time alarm | HD prime out of time alarm
|  |  | HD | HD time out on prime sub-mode purge air state. | Maximum Prime Time Exceeded | Priming failure. Check blood and dialysate lines for clamps and kinks. | During the pre-treatment prime state, if the prime saline purge air state exceeded 60 secs | 228 | ALARM_ID_HD_PRIME_PURGE_AIR_TIME_OUT | ALARM_PRIORITY_LOW | 909 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD time out on prime sub-mode purge air state. | HD time out on prime sub-mode purge air state.
|  |  | HD | HD prime dialysate dialyzer time out alarm | Maximum Prime Time Exceeded | Priming failure. Check blood and dialysate lines for clamps and kinks. | If the priming of the dialyzer dialysate flow path has exceeded the time limit. | 229 | ALARM_ID_PRIME_DIALYSATE_DIALYZER_TIME_OUT | ALARM_PRIORITY_LOW | 909 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD prime dialysate dialyzer time out alarm | HD prime dialysate dialyzer time out alarm
|  |  | HD | HD prime dialysate bypass time out alarm | Maximum Prime Time Exceeded | Priming failure. Check blood and dialysate lines for clamps and kinks. | If the priming of the dialyzer bypass flow path has exceeded the time limit. | 230 | ALARM_ID_PRIME_DIALYSATE_BYPASS_TIME_OUT | ALARM_PRIORITY_LOW | 909 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD prime dialysate bypass time out alarm | HD prime dialysate bypass time out alarm
|  |  | HD | HD pre-treatment mode dry self-tests failure | Self Test Error | Self Test Failure. Please ensure the disposable cartridge is completely installed and the handle moved to the down position. | If the following dry self tests have failed: 1. handleDrySelfTestPressureSensorsState 2. handleDrySelfTestPressureSensorsNormalState | 231 | ALARM_ID_PRE_TREATMENT_DRY_PRESSURE_TEST_FAILURE | ALARM_PRIORITY_LOW | 910 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD pre-treatment mode dry pressure self-test failure | HD pre-treatment mode dry pressure self-test failure
|  |  | HD | HD pre-treatment mode wet self-tests failure | Self Test Error | Self Test Failure. Please ensure the disposable cartridge is completely installed and the handle moved to the down position. Check for closed clamps or kinks in the tubing. | If the following wet self tests have failed: 1. handleWetSelfTestPrimeCheckState 2. handleWetSelfTestFirstDisplacementVerifyState 3. handleWetSelfTestSecondDisplacementVerifyState | 232 | ALARM_ID_PRE_TREATMENT_WET_LC_TEST_FAILURE | ALARM_PRIORITY_HIGH | 799 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD pre-treatment mode wet self-test LC vs. LC failure | HD pre-treatment mode wet self-test LC vs. LC failure
|  |  | ??? | RTC battery low | Battery Error | Battery pack error. | If the controlReg3 & RTC_REG_3_BLF_MASK is logic '1'. | 233 | ALARM_ID_DG_BAROMETRIC_SENSOR_COEFFS_BAD_CRC | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG barometric sensor coefficients bad CRC | Alarm ID available for use
|  |  | ??? | RTC (i.e read time) or RAM operations failure (read or write) | Communication Failure | Communication failure. Ensure all cords are securely connected. | If the status of the RTC write to register0 is NOT successful. | 234 | ALARM_ID_RTC_RAM_OPS_ERROR | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | RTC (i.e read time) or RAM operations failure (read or write) | RTC (i.e read time) or RAM operations failure (read or write)
|  |  | HD | HD in treatment stopped sub-mode after rinseback completed (no escalation) | Rinseback Complete | Rinseback has been completed. | If there is no active alarm and rinseback has completed | 235 | ALARM_ID_TREATMENT_STOPPED_AFTER_RINSEBACK | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD in treatment stopped sub-mode after rinseback completed (no escalation) | HD in treatment stopped sub-mode after rinseback completed (no escalation)
|  |  | HD | HD needs new cartridge to be installed | New Cartridge Needed | New cartridge needs to be installed. | If the air trap sensor reported AIR and bubble sensor detected bubbles. | 236 | ALARM_ID_INSTALL_NEW_CARTRIDGE | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD needs new cartridge to be installed | HD needs new cartridge to be installed
|  |  | HD | HD prime saline dialyzer time out alarm |  |  | This alarm is available. If timeout, just transition to HD_PRIME_RESERVOIR_TWO_FILL_COMPLETE_STATE | 237 | ALARM_ID_PRIME_SALINE_DIALYZER_TIME_OUT | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD prime saline dialyzer time out alarm | HD prime saline dialyzer time out alarm
|  |  | HD | HD no cartridge loaded or installed improperly alarm | Cartridge Needed | Cartridge not loaded or loaded improperly. Check that a cartridge is properly installed. | If the measured blood pump occlucsion pressure reading is < 5000 or > 20000 | 238 | ALARM_ID_NO_CARTRIDGE_LOADED | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD no cartridge loaded or installed improperly alarm | HD no cartridge loaded or installed improperly alarm
|  |  | HD | HD fail to remove cartridge alarm | Cartridge Removal Error | Cartridge not unloaded, or unloaded improperly. Check that the cartridge has been removed. | In the post treatment state, if the cartridge has not been removed | 239 | ALARM_ID_CARTRIDGE_REMOVAL_FAILURE | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD fail to remove cartridge alarm | HD fail to remove cartridge alarm
|  |  | DG | Bicarb conductivity out of range during bicarb pump check alarm | Dialysate Conductivity Error | Dialysate conductivity out of range. Check the bicarbonate dialysate concentrates and make sure the correct concentrate is connected. | If the conductivity value for bicarbonate concentrate is out of range for more than 45 secs | 240 | ALARM_ID_BICARB_CONDUCTIVITY_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Bicarb conductivity out of range during bicarb pump check alarm | Bicarb conductivity out of range during bicarb pump check alarm
|  |  | DG | DG reservoir drain time out | Dialysate Drain Pump Error | Call service to schedule a repair. | If the DG reservoir drain is not complete in 120 secs | 241 | ALARM_ID_DG_RESERVOIR_DRAIN_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG reservoir drain time out | DG reservoir drain time out
|  |  | DG | DG reservoir fill time out | Dialysate Production Error | Call service to schedule a repair. | If the DG reservoir fill is not complete in 120 secs | 242 | ALARM_ID_DG_RESERVOIR_FILL_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG reservoir fill time out | DG reservoir fill time out
|  |  | DG | DG reservoir leak time out | Dialysate Production Error | Call service to schedule a repair. | If the reservoir level is drifted for more than 30 secs | 243 | ALARM_ID_DG_RESERVOIR_LEAK_TIMEOUT | ALARM_PRIORITY_HIGH | 799 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG reservoir leak time out | DG reservoir leak time out
|  |  | DG | DG temperature sensors difference out of range | Dialysate Temperature Error | Call service to schedule a repair. | After the flush, if any of the temperature sensors are out of tolerance | 244 | ALARM_ID_DG_TEMP_SENSORS_DIFF_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG temperature sensors difference out of range | DG temperature sensors difference out of range
|  |  | DG | DG heat disinfect target temperature time out (could not reach to temperature) | Heat Disinfection Error | Call service to schedule a repair. | If heating up to minimum temperature for heat disinfect failed | 245 | ALARM_ID_DG_HEAT_DISINFECT_TARGET_TEMP_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG heat disinfect target temperature time out (could not reach to temperature) | DG heat disinfect target temperature time out (could not reach to temperature)
|  |  | DG | DG heat disinfect inlet pressure and temperature sensors out of range | Heat Disinfection Error | Call service to schedule a repair. | If the inlet pressure and the difference in between TDi and TRo sensors and if they are not in range | 246 | ALARM_ID_DG_HEAT_DISINFECT_INLET_PRES_AND_TEMP_SNSRS_OUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG heat disinfect inlet pressure and temperature sensors out of range | DG heat disinfect inlet pressure and temperature sensors out of range
|  |  | DG | DG heat disinfect inlet conductivity and temperature out of range | Heat Disinfection Error | Call service to schedule a repair. | If the measured temperature and conductivity of water is out of range after the flush time has elapsed | 247 | ALARM_ID_DG_HEAT_DISINFECT_INLET_COND_AND_TEMP_OUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG heat disinfect inlet conductivity and temperature out of range | DG heat disinfect inlet conductivity and temperature out of range
|  |  | DG | DG chemical disinfect could not reach to target temperature | Chemical Disinfection Error | Call service to schedule a repair. | If heating up to minimum temperature for chemical disinfect failed | 248 | ALARM_ID_DG_CHEM_DISINFECT_TARGET_TEMP_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG chemical disinfect could not reach to target temperature | DG chemical disinfect could not reach to target temperature
|  |  | DG | DG chemical disinfect inlet pressure and temperature sensors out of range | Chemical Disinfection Error | Call service to schedule a repair. | If the inlet pressure is less than the threshold of 30.0 PSI and TDi and TRo difference is greater than 3 C | 249 | ALARM_ID_DG_CHEM_DISINFECT_INLET_PRES_AND_TEMP_SNSRS_OUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG chemical disinfect inlet pressure and temperature sensors out of range | DG chemical disinfect inlet pressure and temperature sensors out of range
|  |  | DG | DG chemical disinfect inlet conductivity and temperature out of range | Chemical Disinfection Error | Call service to schedule a repair. | If couldn't get a good water sample after a couple of trials and the disinfect cycle failed | 250 | ALARM_ID_DG_CHEM_DISINFECT_INLET_COND_AND_TEMP_OUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG chemical disinfect inlet conductivity and temperature out of range | DG chemical disinfect inlet conductivity and temperature out of range
|  |  | HD | HD invalid system record | Machine Error | Call service to schedule a repair. | If the calculated CRC of the HD system record image does not match the expected CRC. | 251 | ALARM_ID_HD_INVALID_SYSTEM_RECORD_CRC | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD invalid system record | HD invalid system record
|  |  | HD | HD invalid service record | Machine Error | Call service to schedule a repair. | If the calculated CRC of the HD service record image does not match the expected CRC. | 252 | ALARM_ID_HD_INVALID_SERVICE_RECORD_CRC | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD invalid service record | HD invalid service record
|  |  | DG | DG invalid system record | Machine Error | Call service to schedule a repair. | If the calculated CRC of the DG system record image does not match the expected CRC. | 253 | ALARM_ID_DG_INVALID_SYSTEM_RECORD_CRC | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG invalid system record | DG invalid system record
|  |  | DG | DG invalid service record | Machine Error | Call service to schedule a repair. | If the calculated CRC of the DG service record image does not match the expected CRC. | 254 | ALARM_ID_DG_INVALID_SERVICE_RECORD_CRC | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG invalid service record | DG invalid service record
|  |  | HD | HD and UI software builds are not compatible | Machine Error | Call service to schedule a repair. | If the compatibility revision field of the UI is not equal to SW_COMPATIBILITY_REV (1) | 255 | ALARM_ID_HD_UI_COMPATIBILITY_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD and UI software builds are not compatible | HD and UI software builds are not compatible
|  |  | HD | HD and DG software builds are not compatible |  |  | This alarm is not used in the HD/DG code | 256 | ALARM_ID_HD_DG_COMPATIBILITY_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD and DG software builds are not compatible | HD and DG software builds are not compatible
|  |  | DG | DG FPGA power out timeout | Machine Error | Call service to schedule a repair. | If the FPGA power gpio status pin has a logic '0' for more than 2 secs | 257 | ALARM_ID_DG_FPGA_POWER_OUT_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG FPGA power out timeout | DG FPGA power out timeout
|  |  | HD | HD FPGA power out timeout |  |  | This alarm ID is not used in the HD code. | 258 | ALARM_ID_HD_FPGA_POWER_OUT_TIMEOUT | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD FPGA power out timeout | HD FPGA power out timeout
|  |  | HD | HD temperatures out of range | Dialysate Temperature Error | Call service to schedule a repair. | If thermistor reports temperature > 80.0 or < 0.0 deg F. | 259 | ALARM_ID_HD_TEMPERATURES_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD temperatures out of range | HD temperatures out of range
|  |  | UI | UI POST Application Integrity (Sha256Sum) failure |  |  | This alarm is not used in DG/HD | 260 | ALARM_ID_UI_POST_FAILURE_SHASUM | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST Application Integrity (Sha256Sum) failure | UI POST Application Integrity (Sha256Sum) failure
|  |  | UI | UI POST CANBus failure |  |  | This alarm is not used in DG/HD | 261 | ALARM_ID_UI_POST_FAILURE_CANBUS | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST CANBus failure | UI POST CANBus failure
|  |  | UI | UI POST Display failure |  |  | This alarm is not used in DG/HD | 262 | ALARM_ID_UI_POST_FAILURE_DISPLAY | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST Display failure | UI POST Display failure
|  |  | UI | UI POST Touch failure |  |  | This alarm is not used in DG/HD | 263 | ALARM_ID_UI_POST_FAILURE_TOUCH | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST Touch failure | UI POST Touch failure
|  |  | UI | UI POST SD-Card failure |  |  | This alarm is not used in DG/HD | 264 | ALARM_ID_UI_POST_FAILURE_SDCARD | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST SD-Card failure | UI POST SD-Card failure
|  |  | UI | UI POST RTC failure |  |  | This alarm is not used in DG/HD | 265 | ALARM_ID_UI_POST_FAILURE_RTC | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST RTC failure | UI POST RTC failure
|  |  | UI | UI POST WiFi failure |  |  | This alarm is not used in DG/HD | 266 | ALARM_ID_UI_POST_FAILURE_WIFI | ALARM_PRIORITY_HIGH | 999 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST WiFi failure | UI POST WiFi failure
|  |  | UI | UI POST Bluetooth failure |  |  | This alarm is not used in DG/HD | 267 | ALARM_ID_UI_POST_FAILURE_BLUETOOTH | ALARM_PRIORITY_HIGH | 999 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST Bluetooth failure | UI POST Bluetooth failure
|  |  | UI | UI POST Ethernet failure |  |  | This alarm is not used in DG/HD | 268 | ALARM_ID_UI_POST_FAILURE_ETHERNET | ALARM_PRIORITY_HIGH | 999 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST Ethernet failure | UI POST Ethernet failure
|  |  | UI | UI POST Sound failure |  |  | This alarm is not used in DG/HD | 269 | ALARM_ID_UI_POST_FAILURE_SOUND | ALARM_PRIORITY_HIGH | 999 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | UI POST Sound failure | UI POST Sound failure
|  |  | HD | HD POST Safety Shutdown failure | Self Test Error | Call service to schedule a repair. | If the 24 volts power source is greater than 5.0 volts during safety shutdown self test state. If the 24 volts power source is less than 22.6 volts during safety shutdown self test recover state. | 270 | ALARM_ID_HD_SAFETY_SHUTDOWN_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD POST Safety Shutdown failure | HD POST Safety Shutdown failure
|  |  | DG | DG POST Safety Shutdown failure | Self Test Error | Call service to schedule a repair. | If the 24 volts power source is greater than 5.0 volts during safety shutdown self test state. If the 24 volts power source is less than 22.6 volts during safety shutdown self test recover state. | 271 | ALARM_ID_DG_SAFETY_SHUTDOWN_POST_TEST_FAILED | ALARM_PRIORITY_HIGH | 1 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG POST Safety Shutdown failure | DG POST Safety Shutdown failure
|  |  | HD | HD Fan RPM out of range |  | Fan failure. Check that the fans inlets are not blocked or obstructed. | If the fan RPM is out of Min/Max range. | 272 | ALARM_ID_HD_FAN_RPM_OUT_OF_RANGE | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | HD Fan RPM out of range | HD Fan RPM out of range
|  |  | HD | HD measured blood pump flow rate is out of range | Blood Pump Error | Call service to schedule a repair. | If the blood pump flow rate is > 1320.0 or < 1320.0. | 273 | ALARM_ID_HD_BLOOD_FLOW_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD measured blood pump flow rate is out of range | HD measured blood pump flow rate is out of range
|  |  | HD | HD measured dialysate inlet pump flow rate is out of range | Dialysate Pump Error | Call service to schedule a repair. | If the DPi pump flow rate is > 1320.0 or < 1320.0 | 274 | ALARM_ID_HD_DIAL_IN_FLOW_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD measured dialysate inlet pump flow rate is out of range | HD measured dialysate inlet pump flow rate is out of range
|  |  | HD | HD arterial pressure sensor is reading out of range | Arterial Pressure Error | Call service to schedule a repair. | If the arterial pressure > 2000.0 mmHg or < -1500.0 mmHg. | 275 | ALARM_ID_HD_ARTERIAL_PRESSURE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD arterial pressure sensor is reading out of range | HD arterial pressure sensor is reading out of range
|  |  | HD | HD venous pressure sensor is reading out of range | Venous Pressure Error | Call service to schedule a repair. | If the venous pressure > 2000.0 mmHg or < -1500.0 mmHg. | 276 | ALARM_ID_HD_VENOUS_PRESSURE_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD venous pressure sensor is reading out of range | HD venous pressure sensor is reading out of range
|  |  | HD | HD BP occlusion sensor is reading out of range | Occlusion Sensor Error | Call service to schedule a repair. | If the blood pump occlusion count is < 2000. | 277 | ALARM_ID_HD_BP_OCCLUSION_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD BP occlusion sensor is reading out of range | HD BP occlusion sensor is reading out of range
|  |  | HD | HD active reservoir weight out of range. | Dialysate Production Error | Ultrafiltration error. | If the active reservoir weight is > 1950.0 g. | 278 | ALARM_ID_HD_ACTIVE_RESERVOIR_WEIGHT_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD active reservoir weight out of range. | HD active reservoir weight out of range.
|  |  | DG | DG dialysate drain time out | Dialysate Drain Pump Error |  | If the DG dialysate drain exceeded 120 secs | 279 | ALARM_ID_DG_DIALYSATE_DRAIN_TIME_OUT | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG dialysate drain time out | DG dialysate drain time out
|  |  | HD | HD arterial pressure sensor read timeout error | Arterial Pressure Error | Call service to schedule a repair. | If the arterial pressure read counter has not changed or arterial error counter is > 0 | 280 | ALARM_ID_HD_ARTERIAL_PRESSURE_READ_TIMEOUT_ERROR | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD arterial pressure sensor read timeout error | HD arterial pressure sensor read timeout error
|  |  | DG | DG Acid concentration bottle low volume alarm | Acid Concentrate Low | Acid concentrate low. Replace acid concentrate bottle. | If the acid concentration bottle volume > 29900 (3000 - 100) | 281 | ALARM_ID_DG_ACID_BOTTLE_LOW_VOLUME | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Acid concentration bottle low volume alarm | Acid concentration bottle low volume alarm
|  |  | DG | DG Bicarbonate concentration bottle low volume alarm | Bicarbonate Concentrate Low | Bicarbonate concentrate low. Replace bicarbonate concentrate bottle. | If the bicarbonate concentration bottle volume > 29900 (3000 - 100) | 282 | ALARM_ID_DG_BICARB_BOTTLE_LOW_VOLUME | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | Bicarbonate concentration bottle low volume alarm | Bicarbonate concentration bottle low volume alarm
|  |  | DG | DG load cells weight out of range | Ultrafiltration Measurement Error | Call service to schedule a repair. | If the weight of the DG load cells > 4500.0 or < 0.0 | 283 | ALARM_ID_DG_LOAD_CELL_WEIGHT_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG load cells weight out of range | DG load cells weight out of range
|  |  | DG | DG load cells primary/back up drift out of range |  |  | This alarm is disabled until a better drift algorithm is known | 284 | ALARM_ID_DG_LOAD_CELL_PRIMARY_BACKUP_DRIFT_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG load cells primary/back up drift out of range | DG load cells primary/back up drift out of range
|  |  | HD | HD treatment recirculate timeout warning | Dialysate Recirculation Error | Dialysate recirculation error. | If the maximum time in recirculate sub-mode has been exceeded | 285 | ALARM_ID_HD_TREATMENT_RECIRC_TIMEOUT_WARNING | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | HD treatment recirculate timeout warning | HD treatment recirculate timeout warning
|  |  | HD | HD treatment rinseback complete timeout warning | Rinseback Complete | Rinseback complete. | If the rise back time counter checks have failed when full blood volume has been delivered back to the patient | 286 | ALARM_ID_HD_TREATMENT_RINSEBACK_TIMEOUT_WARNING | ALARM_PRIORITY_MEDIUM | 899 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | TRUE | FALSE | 0 | ALARM_ID_NO_ALARM | HD treatment rinseback complete timeout warning | HD treatment rinseback complete timeout warning
|  |  | HD | HD processor clock speed checks against FPGA clock failure | Machine Error | Call service to schedule a repair. | If during initialization mode, the difference between the processor clock speed and the FPGA clock speed > 1. | 287 | ALARM_ID_HD_FPGA_CLOCK_SPEED_CHECK_FAILURE | ALARM_PRIORITY_HIGH | 3 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD processor clock speed checks against FPGA clock failure | HD processor clock speed checks against FPGA clock failure
|  |  | DG | DG trimmer heater on with flow timeout. | Dialysate Temperature Error | Call service to schedule a repair. | If the heater is ON with no flow detected for more than 120 secs | 288 | ALARM_ID_HD_LOAD_CELL_PRIMARY_BACKUP_DRIFT_OUT_OF_RANGE | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | HD load cells primary/back up drift out of range | Alarm ID available for use
|  |  | DG | DG cap in wrong state | DG Cap State Error | Check dialysate/concentrate caps. | DG dialysate and/or concentrate cap is in the wrong state | 289 | ALARM_ID_DG_DIALYSATE_OR_CONC_CAP_NOT_IN_PROPER_POSITION | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG dialysate or concentrate caps not closed | DG dialysate or concentrate caps not closed
[8 empty] | 290 | ALARM_ID_DG_FLOW_TOO_LOW_WHILE_HEATER_ON | ALARM_PRIORITY_HIGH | 3 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | 0 | ALARM_ID_NO_ALARM | DG flow too low while heater is on | DG flow too low while heater is on

## Property Definitions
Property Name | Definition
AlarmID | Unique alarm identification number
Alarm Name | Name of the alarm
Priority | Priority of the alarm (High, Medium, Low)
Rank | Within the same priority, higher rank alarms have higher priority
Fault | The alarm activation causes an HD fault
DGFault | The alarm activation causes an DG fault
Stop | The alarm activation causes a controlled stop
Clear Immediate | Alarm condition is cleared immediately after alarm activation
Non Recoverable | Alarm cannot be cleared (unrecoverable)
No Resume | Alarm activation prevents treatment resume
No Rinseback | Alarm activation prevents rinseback
No End Treatment | Alarm activation prevents ending treatment
No New Treatment | Alarm activation prevents new treatment
No Dialysate Recirc | Alarm activation prevents dialysate re-circulation
Clear Top Only | User acknowledgment clears only displayed alarm
Escalate In | Alarm will escalate if not cleared within this time (zero indicates no escalation)
Escalate To | ID of the alarm that this alarm will be escalated to