## Alarms
Action | PRS ID2 | SA | HD SRS ID | DG SRS ID | UI SRS ID | Device | Displayed Title | Char count | Displayed Instructions | Char Count2 | Description | Trigger Conditions | AlarmID | Alarm Name | Alarm ID Check | Priority | Rank | Fault | DGFault | Stop | Clear Immediate | Non Recoverable | No Resume | No Rinseback | No End Treatment | Block Until SB Rinseback | Block Until SB End Tx | No Blood Recirc | No Dialysate Recirc | Clear Top Only | TreatmentLog | Doxygen Title
N/A | N/A |  | N/A | N/A |  | N/A | ` |  | N/A |  | Indication for when no alarms is a possible situation | N/A | 0 | ALARM_ID_NO_ALARM |  | ALARM_PRIORITY_NONE | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | Indication for when no alarms is a possible situation
x | PRS 398 | SA 213 | SRSHD 652 |  |  | HD | HD POST: Stuck Button |  | Stop or On button pressed during startup\n-Release button and power cycle\n-Call service to schedule a repair if problem persists |  | Stuck button POST failure. Stop or Off button detected to be pressed for at least 1 second during test shortly after power up | Stop or Off button are pressed for more than 1 sec during POST. | 1 | ALARM_ID_HD_STUCK_BUTTON_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD stuck button POST failure. Stop or Off button detected to be pressed for at least 1 second during test shortly after power up
x | PRS 398 | SA 207 | SRSHD 995 SRSHD 1357 |  |  | HD | HD POST: FPGA Compatibility |  | Call service to schedule a repair. |  | HD FPGA POST failure | Minimum HD FPGA major revision is not compatible with this version of HD firmware or the FPGA Id is not correct. | 2 | ALARM_ID_HD_FPGA_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD FPGA POST failure
x | PRS 863 | SA 287 |  | SRSDG 602 SRSDG 814 SRSDG 1061 |  | DG | DG POST: FPGA Compatibility |  | Call service to schedule a repair. |  | DG FPGA POST failure | Minimum DG FPGA major and minor revision are not equal to the FPGA major and minor revision in FPGA header | 3 | ALARM_ID_DG_FPGA_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG FPGA POST failure
Sean - check if DG and HD POST test is the same in code. Consider updating trigger conditions | PRS 398 | SA 208 SA 209 SA 210 SA 211 | SRSHD 650 |  |  | HD | HD POST: WD Safestate |  | Call service to schedule a repair. |  | HD Watchdog POST failure | If the following are not observed after 500ms of not petting the WD: 1. WD exp pin=low 2. 24V <= 5V 3. audio current >= 200mA If the following are not observed after 500ms of resumed WD petting at proper timing: 1. WD exp pin=high 2. 24V >= 22.6V 3. audio current <= 10 mA | 4 | ALARM_ID_HD_WATCHDOG_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD Watchdog POST failure
Sean - check if DG and HD POST test is the same in code. Consider updating trigger conditions | PRS 863 | SA 288 SA 290 SA 289 |  | SRSDG 53 |  | DG | DG POST: WD Safestate |  | Call service to schedule a repair. |  | DG Watchdog POST failure | If the following are not observed after 500ms of not petting the WD: 1. WD exp pin=low 2. 24V <= 5V If the following are not observed after 750ms of resumed WD petting at proper timing: 1. WD exp pin=high 2. 24V >= 22.6V | 5 | ALARM_ID_DG_WATCHDOG_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG Watchdog POST failure
x | PRS 398 | SA 204 | SRSHD 1484 |  |  | HD | HD POST: Proc-UI Comm |  | Call service to schedule a repair. |  | UI communication POST failure. UI failed to communicate within a reasonable time after power up | If the UI failed to checkin via the CAN bus within 40 seconds | 6 | ALARM_ID_HD_UI_COMM_POST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI communication POST failure. UI failed to communicate within a reasonable time after power up
x | PRS 872 | SA 457 | SRSHD 1595 |  |  | HD | HD Alarm (Low): RTC Bat Low |  | Call service to schedule a repair. |  | Real-time clock configuration error | If the RTC is not functional as indicated by the bits in the first 3 control registers. | 7 | ALARM_ID_HD_RTC_BATTERY_LOW |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD RTC battery low
x | PRS 398 | SA 471 | SRSHD 632 |  |  | HD | HD POST: Acc Vector |  | Call service to schedule a repair. |  | HD accelerometer failed POST | If the calculated vector length of the accelerometer x, y, z is not the nominal length. | 8 | ALARM_ID_HD_ACCELEROMETER_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD accelerometer failed POST
x | PRS 863 | SA 472 |  | SRSDG 323 |  | DG | DG POST: Acc Vector |  | Call service to schedule a repair. |  | DG accelerometer failed POST | If the calculated vector length of the accelerometer x, y, z is not the nominal length. | 9 | ALARM_ID_DG_ACCELEROMETER_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG accelerometer failed POST
x | PRS 398 | SA 480 | SRSHD 1478 |  |  | HD | HD POST: BLD Config Emb Mode |  | Call service to schedule a repair. |  | HD blood leak sensor embedded mode communication failed | If the blood leak sensor embedded mode is unresponsive. | 10 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_EMBEDDED_MODE_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood leak sensor setting embedded mode failure.
Dialysate Temp | [5 empty] | HD | Dial. Gen: Dialysate Temp Low |  | Dialysate temperature is low\n- Treatment is paused and will be ready to resume when \ndialysate is within target range. |  | Dialysate temperature is below the target temperature | If the dialysate temperature is below the target temperature for a certain period of time. | 11 | ALARM_ID_HD_DIALYSATE_TEMP_BELOW_TARGET_TEMP |  | ALARM_PRIORITY_MEDIUM | 402 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | HD dialysate temperature below target temperature or below safety temp
x | PRS 398 | SA 195 | SRSHD 658 |  |  | HD | HD POST: HD Image |  | Call service to schedule a repair. |  | HD firmware image integrity POST test failed | If the calculated CRC of the HD image does not match the expected CRC. | 12 | ALARM_ID_HD_INTEGRITY_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD firmware image integrity POST test failed
x | PRS 863 | SA 282 |  | SRSDG 813 |  | DG | DG POST: DG Image |  | Call service to schedule a repair. |  | DG firmware image integrity POST test failed | If the calculated CRC of the DG firmware image does not matched the expected CRC | 13 | ALARM_ID_DG_INTEGRITY_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG firmware image integrity POST test failed
x | PRS 863 | SA 284 |  | SRSDG 957 |  | DG | DG POST: Usage Rec NVRAM CRC |  | Call service to schedule a repair. |  | DG non-volatile usage information group invalid record CRC | If the DG usage information record's CRC does not match the calculated CRC in POST. | 14 | ALARM_ID_DG_INVALID_USAGE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid usage record CRC
Dara - remove Alarm ID 15 (Duplicate with Alarm ID 304 - SRS 1587) | [5 empty] | HD | Dial. Gen: Dialysate Temp High |  | Dialysate temperature is high\n-Treatment is paused and will be ready to resume when \ndialysate is within target range. |  | Dialysate temperature is out of high safety range | If the dialysate temperature is > 46 C for a certain period of time. | 15 | ALARM_ID_HD_DIALYSATE_TEMP_ABOVE_SAFETY_TEMP |  | ALARM_PRIORITY_MEDIUM | 401 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | HD dialysate temperature above high safety.
x | PRS 398 | SA 212 | SRSHD 656 |  |  | HD | HD POST: Primary Alarm Current |  | Call service to schedule a repair. |  | HD alarm audio failed POST | If alarm audio current is NOT in expected range for than 1 second while alarm tone is being output | 16 | ALARM_ID_HD_ALARM_AUDIO_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD alarm audio failed POST
Sean to review and update |  |  | SRSHD 1677 |  |  | HD | HD POST: UI POST Not Complete |  | Call service to schedule a repair. |  | HD UI POST failed | If UI reports POST failed but did not trigger fault OR UI has not sent POST results within 2 seconds. | 17 | ALARM_ID_HD_UI_POST_TIMEOUT |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST failed
[6 empty] | HD | Loss of DG Comm (DG ACK) |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. |  | HD to DG comm error. | HD message to DG that requires acknowledge was not acknowledged within a certain time. | 18 | ALARM_ID_HD_CAN_MESSAGE_NOT_ACKED_BY_DG |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD didn't get ACK on message to DG that required acknowledgment
Dialysate Temp |  |  | SRSHD 901 |  |  | HD | Dial. Gen: Dialysate Temp High |  | Dialysate temperature is high\n- Treatment is paused and will be ready to resume when \ndialysate is within target range. |  | Dialysate temperature is above the target temperature | If the dialysate temperature is above the target temperature for a certain period of time. | 19 | ALARM_ID_HD_DIALYSATE_TEMP_ABOVE_TARGET_TEMP |  | ALARM_PRIORITY_MEDIUM | 401 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | HD dialysate temperature above target temperature.
x | PRS 863 | SA 369 |  | SRSDG 16 |  | DG | DG POST: Conductivity Cal |  | Call service to schedule a repair. |  | DG conductivity sensors invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 20 | ALARM_ID_DG_COND_SENSORS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG conductivity sensors invalid calibration record
x | PRS 863 | SA 369 |  | SRSDG 936 |  | DG | DG POST: Drain Line Vol Cal |  | Call service to schedule a repair. |  | DG drain line volume invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 21 | ALARM_ID_DG_DRAIN_LINE_VOLUME_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain line volume invalid calibration record
x | PRS 863 | SA 369 |  | SRSDG 935 |  | DG | DG POST: Reservoir Vol Cal |  | Call service to schedule a repair. |  | DG reservoirs invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 22 | ALARM_ID_DG_RESERVOIRS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG reservoirs invalid calibration record
x | PRS 863 | SA 369 |  | SRSDG 934 |  | DG | DG POST: Acid Conc Cal |  | Call service to schedule a repair. |  | DG acid concentrate calibration record | If the non-volatile data is valid and if not raise the alarm | 23 | ALARM_ID_DG_ACID_CONCENTRATE_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG acid concentrate calibration record
x | PRS 863 | SA 369 |  | SRSDG 933 |  | DG | DG POST: Bicarb Conc Cal |  | Call service to schedule a repair. |  | DG bicarb concentrate calibration record | If the non-volatile data is valid and if not raise the alarm | 24 | ALARM_ID_DG_BICARB_CONCENTRATE_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG bicarb concentrate calibration record
x | PRS 863 | SA 369 |  | SRSDG 817 |  | DG | DG POST: Acc Cal |  | Call service to schedule a repair. |  | DG accelerometer invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 25 | ALARM_ID_DG_ACCELEROMETERS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG accelerometer invalid calibration record
x | PRS 398 | SA 368 | SRSHD 1576 |  |  | HD | HD POST: Acc Cal |  | Call service to schedule a repair. |  | HD accelerometer invalid calibration record | If the accel calibration data that was received from NVDataMgmt is NOT legitimate. | 26 | ALARM_ID_HD_ACCELEROMETERS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD accelerometer invalid calibration record
Move to wet self-test - procedural |  |  | SRSHD 1428 |  |  | HD | Process: Pre-Tx Wet Self-Test |  | Self-test detected has a problem with the blood leak sensor\n- Check cartridge installation on to machine\n- If problem persists, terminate pre-treatment and restart \nwith a new cartridge. |  | HD BLD sensor failed zero sequence | BLD fails to complete zero sequence. | 27 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_ZERO_SEQUENCE_FAILED |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD blood leak sensor zero and self test sequence failed
x | PRS 863 | SA 319 SA 320 |  | SRSDG 864 |  | DG | DG Fault: Conductivity RTD Comm |  | Call service to schedule a repair. |  | DG two wire sensors FPGA fault | If the FPGA does not report fresh data for a certain period of time | 28 | ALARM_ID_DG_TWO_WIRE_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG two wire sensors FPGA fault
x | PRS 398 | SA 368 | SRSHD 1482 |  |  | HD | HD POST: SP DAC Cal |  | Call service to schedule a repair. |  | HD heparin force sensor invalid calibration record | If the heparin force sensor calibration data that was received from NVDataMgmt is NOT legitimate. | 29 | ALARM_ID_HD_HEPARIN_FORCE_SENSOR_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD heparin force sensor invalid calibration record
x | PRS 494 | SA 416 | SRSHD 1704 |  |  | HD | HD Fault: SW Invalid State |  | Call service to schedule a repair. |  | HD Software fault. Software found itself in an unexpected state | If the specific software faults listed in alarm management SW faults has been detected | 30 | ALARM_ID_HD_SOFTWARE_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD Software fault. Software found itself in an unexpected state
x | PRS 494 | SA 241 SA 242 | SRSHD 515 |  |  | HD | HD Fault: BP Current |  | Call service to schedule a repair. |  | Blood pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | If the measured MC current is not greater than BP_MAX_CURR_WHEN_STOPPED_MA or is not greater than BP_MAX_CURR_WHEN_RUNNING_MA. | 31 | ALARM_ID_HD_BLOOD_PUMP_MC_CURRENT_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running
x | PRS 494 | SA 239 | SRSHD 1087 |  |  | HD | HD Fault: BP Speed |  | Call service to schedule a repair. |  | Blood pump failed motor off check. Measured speed while commanded off | If the measured motor speed is greater than BP_MAX_MOTOR_SPEED_WHILE_OFF_RPM | 32 | ALARM_ID_HD_BLOOD_PUMP_OFF_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed motor off check. Measured speed while commanded off
x | PRS 494 | SA 240 SA 244 | SRSHD 1079 |  |  | HD | HD Fault: BP Direction |  | Call service to schedule a repair. |  | Blood pump failed motor direction check. Measured vs. commanded | If the set direction versus hall sensors direction versus sign of motor speed DOES NOT agree. | 33 | ALARM_ID_HD_BLOOD_PUMP_MC_DIRECTION_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed motor direction check. Measured vs. commanded
x | PRS 494 | SA 246 | SRSHD 1081 |  |  | HD | HD Fault: BP Rotor |  | Blood pump rotor failure.\n- Check the pump rotor for secure connection. |  | Blood pump failed rotor speed check. Mismatch with rotor and motor speeds | If the delta of the measured motor speed and the commanded motor speed is not within tolerance. | 34 | ALARM_ID_HD_BLOOD_PUMP_ROTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed rotor speed check. Mismatch with rotor and motor speeds
x | PRS 494 | SA 250 SA 251 | SRSHD 542 |  |  | HD | HD Fault: DPi Current |  | Call service to schedule a repair. |  | Dialysis inlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | If the Motor Controller measured DPi pump current > 150.0 mA for more than 250.0 msec. | 35 | ALARM_ID_HD_DIAL_IN_PUMP_MC_CURRENT_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis inlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running
x | PRS 494 | SA 248 | SRSHD 1088 |  |  | HD | HD Fault: DPi Speed |  | Call service to schedule a repair. |  | Dialysis inlet pump failed motor off check. Measured speed while commanded off | If during command off state, the DPi pump motor speed is measured > 100.0 RMP. | 36 | ALARM_ID_HD_DIAL_IN_PUMP_OFF_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis inlet pump failed motor off check. Measured speed while commanded off
x | PRS 494 | SA 249 SA 253 | SRSHD 1050 |  |  | HD | HD Fault: DPi Direction |  | Call service to schedule a repair. |  | Dialysis inlet pump failed motor direction check. Measured vs. commanded | If the set direction versus hall sensors direction DO NOT agree for more than 250.0 msec. | 37 | ALARM_ID_HD_DIAL_IN_PUMP_MC_DIRECTION_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis inlet pump failed motor direction check. Measured vs. commanded
x | PRS 494 | SA 255 | SRSHD 1082 |  |  | HD | HD Fault: DPi Rotor |  | Dialysate in pump rotor failure.\n- Check the pump rotor for secure connection. |  | Dialysis inlet pump failed rotor speed check. Mismatch with rotor and motor speeds | If the measured motor speed versus the measured motor speed while controlling the target > 5.0 for more than 12 seconds. | 38 | ALARM_ID_HD_DIAL_IN_PUMP_ROTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis inlet pump failed rotor speed check. Mismatch with rotor and motor speeds
x | PRS 494 | SA 258 SA 259 | SRSHD 554 |  |  | HD | HD Fault: DPo Current |  | Call service to schedule a repair. |  | Dialysis outlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | If the DPo pump current > 150.0 mA while stopped or Dpo current > 2000.0 mA for more than 2 seconds while running. | 39 | ALARM_ID_HD_DIAL_OUT_PUMP_MC_CURRENT_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis outlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running
x | PRS 494 | SA 256 | SRSHD 1089 |  |  | HD | HD Fault: DPo Speed |  | Call service to schedule a repair. |  | Dialysis outlet pump failed motor off check. Measured speed while commanded off | If the Dpo pump current > 100.0 RPM while commanded off for more than 5000 msec. | 40 | ALARM_ID_HD_DIAL_OUT_PUMP_OFF_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis outlet pump failed motor off check. Measured speed while commanded off
x | PRS 494 | SA 257 SA 261 | SRSHD 1051 |  |  | HD | HD Fault: DPo Direction |  | Call service to schedule a repair. |  | Dialysis outlet pump failed motor direction check. Measured vs. commanded | If the DPo pump set direction does not agree with hall sensors or set direction does not agree with the sign of the motor speed for more than 25 msec. | 41 | ALARM_ID_HD_DIAL_OUT_PUMP_MC_DIRECTION_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis outlet pump failed motor direction check. Measured vs. commanded
x | PRS 494 | SA 263 | SRSHD 1083 |  |  | HD | HD Fault: DPo Rotor |  | Dialysate out pump rotor failure.\n- Check the pump rotor for secure connection. |  | Dialysis outlet pump failed rotor speed check. Mismatch with rotor and motor speeds | If DPo pump measured rotor speed does not agreed with the measured/commanded motor speed | 42 | ALARM_ID_HD_DIAL_OUT_PUMP_ROTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis outlet pump failed rotor speed check. Mismatch with rotor and motor speeds
x | PRS 863 | SA 357 |  | SRSDG 954 |  | DG | DG Fault: Iso ADC Comm |  | Call service to schedule a repair. |  | Main primary heater FPGA fault | If the FPGA does not report fresh data for a certain period of time | 43 | ALARM_ID_DG_MAIN_PRIMARY_HEATER_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG main primary heater FPGA fault
x | PRS 494 | SA 226 | SRSHD 1001 |  |  | HD | HD Fault: UI Comm |  | Call service to schedule a repair. |  | User interface communication timeout | If the sub-system communication timeout has occurred. UI is 5 seconds. DG is 2 seconds. | 44 | ALARM_ID_HD_UI_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI communication timeout
x | PRS 494 | SA 226 SA 477 | SRSHD 967 |  |  | HD | HD Fault: CAN CRC Com |  | Call service to schedule a repair. |  | Too many bad CRCs detected on received system messages | If too many bad messages CRCs are received within a set period of time | 45 | ALARM_ID_HD_COMM_TOO_MANY_BAD_CRCS |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD too many bad communications CRC
x | PRS 494 | SA 226 SA 477 | SRSHD 964 |  |  | HD | HD Fault: CAN ACK UI Com |  | Call service to schedule a repair. |  | System message that required acknowledgment was not acknowledged | If no retries left in the pending acknowlegement list. | 46 | ALARM_ID_HD_CAN_MESSAGE_NOT_ACKED_BY_UI |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD didn't get ACK on message to UI that required acknowledgment
x | PRS 494 | SA 476 | SRSHD 907 | [9 empty] | 47 | ALARM_ID_AVAILABLE_16 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | Available for use
x | PRS 494 | SA 452 | SRSHD 1029 |  |  | HD | HD Fault: UF Rate Tare Err |  | Call service to schedule a repair. |  | Ultrafiltration volume accuracy error during treatment | If the UF volume (refUFvolume - measureVolume) > 100.0 mL. | 48 | ALARM_ID_HD_UF_VOLUME_ACCURACY_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD ultrafiltration volume accuracy error during treatment
x | PRS 494 | SA 224 | SRSHD 1356 SRSHD 1490 |  |  | HD | HD Fault: FPGA Com |  | Call service to schedule a repair. |  | HD FPGA communication down for too long | If there are more than 3 comm errors in a 1 min period. | 49 | ALARM_ID_HD_FPGA_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD FPGA communication down for too long
x | PRS 863 | SA 481 |  | SRSDG 26 |  | DG | DG Fault: Valve Control |  | Call service to schedule a repair. |  | DG FPGA not accepting commanded valve states | If the read back FPGA valve states does not match the commanded valve states for a certain amount of time. | 50 | ALARM_ID_DG_VALVE_CONTROL_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG FPGA not accepting commanded valve states
x | PRS 494 | SA 239 SA 243 | SRSHD 1098 |  |  | HD | HD Fault: BP Speed |  | Call service to schedule a repair. |  | Blood pump failed motor speed check. Measured vs. commanded | If measured motor speed versus commanded rotor speed exceeded limits. | 51 | ALARM_ID_HD_BLOOD_PUMP_MOTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed motor speed check. Measured vs. commanded
x | PRS 494 | SA 252 | SRSHD 1099 |  |  | HD | HD Fault: DPi Speed |  | Call service to schedule a repair. |  | Dialysate inlet pump failed motor speed check. Measured vs. commanded | If measured motor speed versus commanded motor speed exceeded limits for more than 5 secs. | 52 | ALARM_ID_HD_DIAL_IN_PUMP_MOTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysate inlet pump failed motor speed check. Measured vs. commanded
x | PRS 494 | SA 260 SA 256 | SRSHD 1198 SRSHD 1431 SRSHD 1670 |  |  | HD | HD Fault: DPo Speed |  | Call service to schedule a repair. |  | Dialysate outlet pump failed motor speed check. Measured vs. commanded | If measured motor speed versus commanded motor speed exceeded limits for more than 5 secs. | 53 | ALARM_ID_HD_DIAL_OUT_PUMP_MOTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysate outlet pump failed motor speed check. Measured vs. commanded
x | PRS 494 | SA 501 | SRSHD 1705 |  |  | HD | HD Fault: Critical Data Err |  | Call service to schedule a repair. |  | HD critical data integrity check failed | If the critical data failed range check. | 54 | ALARM_ID_HD_CRITICAL_DATA_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD critical data integrity check failed
| PRS 863 | SA 500 |  | SRSDG 921 |  | DG | DG Fault: Critical Data Err |  | Call service to schedule a repair. |  | DG critical data integrity check failed | If the critical data failed range check. | 55 | ALARM_ID_DG_CRITICAL_DATA_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG critical data integrity check failed
x | PRS 494 | SA 232 | SRSHD 1526 |  |  | HD | HD Fault: Acc Com |  | Call service to schedule a repair. |  | HD accelerometer error (no readings or FPGA reports error) | If no fresh sample of accelerometer data has been detected for more than 100 msec. | 56 | ALARM_ID_HD_ACCELEROMETER_FAILURE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD accelerometer error (no readings or FPGA reports error)
x | PRS 863 | SA 346 |  | SRSDG 869 SRSDG 870 |  | DG | DG Fault: Acc Com |  | Call service to schedule a repair. |  | DG accelerometer error (no readings or FPGA reports error) | If no fresh sample of accelerometer data has been detected for more than 100 msec. | 57 | ALARM_ID_DG_ACCELEROMETER_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG accelerometer error (no readings or FPGA reports error)
x | PRS 494 | SA 384 SA 385 SA 386 SA 387 | SRSHD 1340 |  |  | HD | HD Fault: Valve Home |  | Call service to schedule a repair. |  | HD valve homing failed | If the number of failed homing > = 3 counts. | 58 | ALARM_ID_HD_VALVE_HOMING_FAILED |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD valve homing failed
x | PRS 494 | SA 264 SA 266 SA 268 SA 270 | SRSHD 1338 |  |  | HD | HD Fault: Valve Transition Err |  | Call service to schedule a repair. |  | HD valve transition time out | if the valve's transition time has timed out (3000msec) | 59 | ALARM_ID_HD_VALVE_TRANSITION_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD valve transition time out
x | PRS 494 | SA 502 SA 505 SA 503 SA 504 | SRSHD 1515 SRSHD 1516 SRSHD 1632 |  |  | HD | HD Fault: Valve Control |  | Call service to schedule a repair. |  | HD valve not functional | If the valve bit masks are not as expected in different valve control mode. | 60 | ALARM_ID_HD_VALVE_NOT_FUNCTIONAL |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD valve not functional
x | PRS 494 | SA 265 SA 267 SA 269 SA 271 | SRSHD 1339 |  |  | HD | HD Fault: Valve Current |  | Call service to schedule a repair. |  | HD valve current out of range | If the valve current is over the threshold of 1.0 A for 1000 msec. | 61 | ALARM_ID_HD_VALVE_CURRENT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD valve current out of range
x | PRS 494 | SA 264 SA 266 SA 268 SA 270 | SRSHD 1336 |  |  | HD | HD Fault: Valve Position |  | Call service to schedule a repair. |  | HD valve position out of target | If the current position has deviated from the expected position for more than 1 second. | 62 | ALARM_ID_HD_VALVE_POSITION_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD valve position out of target
x | PRS 494 | SA 453 | SRSHD 1586 |  |  | HD | HD Fault: BLD Com |  | Call service to schedule a repair. |  | FPGA reports error | If the FPGA error counter increments for a certain period of time. | 63 | ALARM_ID_HD_BLOOD_LEAK_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood leak FPGA fault
x | PRS 863 | SA 350 |  | SRSDG 874 |  | DG | DG Fault: Baro Range |  | Call service to schedule a repair. |  | Barometric pressure is out of range | If the barometric pressure is > 10.1 psi or < 15.4 psi for a certain period of time. | 64 | ALARM_ID_DG_BARO_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG barometric pressure sensor out of range
x | PRS 400 |  | SRSHD 1531 |  |  | HD | Loss of DG Comm (Reject) |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. |  | HD requests DG command with invalid parameter fault | If the latest DG command response indicated invalid parameter | 65 | ALARM_ID_HD_DG_COMMAND_INVALID_PARAMETER_FAULT |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD DG requests DG command with invalid parameter fault
x | PRS 398 | SA 480 | SRSHD 1427 SRSHD 1429 SRSHD 1631 |  |  | HD | HD POST: BLD Config Set Point |  | Call service to schedule a repair. |  | HD blood leak sensor set point set failure | If the set point of the blood leak sensor does not match the set point in the calibration record. | 66 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_SET_POINT_SET_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood leak sensor set point set failure
No cartridge self-test (procedural) |  |  | SRSHD 1511 |  |  | HD | Process: Pre-Tx Remove Cartridge |  | Cartridge has been detected.\n- Open door\n- Remove cartridge\n- Close door\n- Resume |  | HD blood pump occlusion self-test failure alarm | If the blood pump pressure > 5000 | 67 | ALARM_ID_HD_BP_OCCLUSION_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 700 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD blood pump occlusion self-test failure alarm
agree to change rank and re-assess similar to temperature alarm Blaine |  |  | SRSHD 1426 |  |  | HD | Dial. Gen: Dialysate not ready |  | Dialysate is not yet ready.\n- Treatment is paused and will be ready to resume when \ndialysate is ready for use. |  | HD active reservoir recirculation out of range | If the recirculation level has exceeded the max allowed | 68 | ALARM_ID_HD_ACTIVE_RESERVOIR_RECIRCULATION_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 421 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | HD active reservoir recirculation out of range
x | PRS 398 | SA 368 | SRSHD 532 |  |  | HD | HD POST: BLD Cal |  | Call service to schedule a repair. |  | HD blood leak sensor invalid calibration record | If the BLD data that was received from NVDataMgmt is NOT legitimate. | 69 | ALARM_ID_HD_BLOOD_LEAK_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood leak sensor invalid calibration record
This should be in-treatment alarm? Not machine alarm? See alarm ID 275 for machine alarm (valid range check) agree - make this in-treatment alarms not faults Blaine to update - not a fault Sean to update the range to what is in the SRS |  |  | SRSHD 737 |  |  | HD | Process: Pre-Tx Remove Cartridge |  | Cartridge has been detected.\n- Open door\n- Remove cartridge\n- Close door\n- Resume |  | HD arterial pressure self-test failure alarm | If arterial pressure < -300 or > 100 mmHg | 70 | ALARM_ID_HD_ARTERIAL_PRESSURE_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 700 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD arterial pressure self-test failure alarm
This should be in-treatment alarm? Not machine alarm? See alarm ID 275 for machine alarm (valid range check) agree - make this in-treatment alarms not faults Blaine to update - not a fault Sean to update the range to what is in the SRS |  |  | SRSHD 738 |  |  | HD | Process: Pre-Tx Remove Cartridge |  | Cartridge has been detected.\n- Open door\n- Remove cartridge\n- Close door\n- Resume |  | HD venous pressure self-test failure alarm | If the venous pressure < -100 or > 600 mmHg | 71 | ALARM_ID_HD_VENOUS_PRESSURE_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 700 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD venous pressure self-test failure alarm
x | PRS 400 |  | SRSHD 1126 |  |  | HD | Loss of DG Comm (LC) |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. |  | HD No load cell data message receive at least once every 2 seconds | No new load cell data for 2 seconds. | 72 | ALARM_ID_HD_NEW_LOAD_CELL_DATA_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | HD No load cell data message receive at least once every 2 seconds
x | PRS 400 |  | SRSHD 1573 |  |  | HD | Loss of DG Comm (TDI/TRO) |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. |  | HD No dialysate temperature data message received at least once every 2 seconds | No new dialysate temperature data for 2 seconds. | 73 | ALARM_ID_HD_NEW_DIALYSATE_TEMP_DATA_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | HD No dialysate temperature data message receive at least once every 2 seconds
x | PRS 863 | SA 364 |  | SRSDG 600 |  | DG | DG Fault: UVi On No Flow |  | Call service to schedule a repair. |  | DG inlet UV reactor is on while there is no flow | If the inlet UV reactor is on while VPi is closed for a certain period of time. | 74 | ALARM_ID_DG_INLET_UV_REACTOR_ON_WITH_NO_FLOW |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG inlet UV reactor is on with no flow
x | PRS 494 | SA 469 | SRSHD 1488 |  |  | HD | HD Fault: Syringe DAC |  | Call service to schedule a repair. |  | HD syringe pump self-test failure alarm | If the calibration data that was received from NVDataMgmt and the DAC voltage with syringe retracted are not within tol. | 75 | ALARM_ID_HD_SYRINGE_PUMP_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump self-test failure alarm
Sean to update code | PRS 494 | SA 217 SA 218 SA 219 | SRSHD 635 SRSHD 1678 SRSHD 1681 |  |  | HD | HD Fault: Voltage Monitor |  | Call service to schedule a repair. |  | HD monitored voltage is out of range | If the HD monitored voltages are out of range for a certain period of time. | 76 | ALARM_ID_HD_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD monitored voltage is out of range
Sean to update code | PRS 863 | SA 295 SA 296 SA 297 |  | SRSDG 46 SRSDG 866 SRSDG 1088 |  | DG | DG Fault: Voltage Monitor |  | Call service to schedule a repair. |  | DG monitored voltage is out of range | If the DG monitored voltages are out of range for a certain period of time. | 77 | ALARM_ID_DG_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG monitored voltage is out of range
x | PRS 494 | SA 275 | SRSHD 1324 |  |  | HD | HD Fault: Syringe Pump Enc Dir |  | Call service to schedule a repair. |  | HD syringe pump direction (from encoder) error | If the measured encoder direction of the syringe pump does not agree with the expected direction | 78 | ALARM_ID_HD_SYRINGE_PUMP_ENCODER_DIRECTION_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump direction (from encoder) error
x | PRS 494 | SA 489 | SRSHD 1507 |  |  | HD | HD Fault: Syringe Pump Drv Dir |  | Call service to schedule a repair. |  | HD syringe pump direction (from controller) error | If the measured direction of the syringe pump motor controller does not agree with the expected direction | 79 | ALARM_ID_HD_SYRINGE_PUMP_CONTROLLER_DIRECTION_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump direction (from controller) error
x | PRS 494 | SA 272 | SRSHD 1317 |  |  | HD | HD Fault: Syringe Pump Driver |  | Call service to schedule a repair. |  | HD syringe pump fault reported by FPGA | If the FPGA reported the syringe pump fault | 80 | ALARM_ID_HD_SYRINGE_PUMP_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump fault reported by FPGA
x | PRS 494 | SA 488 | SRSHD 1330 |  |  | HD | HD Fault: Syringe Pump OT |  | Call service to schedule a repair. |  | HD syringe pump over travel error | If the syringe pump position is deemed to have overtraveled | 81 | ALARM_ID_HD_SYRINGE_PUMP_OVER_TRAVEL_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump over travel error
x | PRS 494 | SA 277 | SRSHD 1506 |  |  | HD | HD Fault: Syringe Pump DAC Err |  | Call service to schedule a repair. |  | HD syringe pump DAC write failure | If the syringe pump DAC write error is reported by the FPGA | 82 | ALARM_ID_HD_SYRINGE_PUMP_DAC_WRITE_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump DAC write failure
x | PRS 494 | SA 415 | SRSHD 1328 |  |  | HD | HD Fault: Syringe Pump On Err |  | Call service to schedule a repair. |  | HD syringe pump is running while the BP is off | If the syringe pump is on while the blood pump is off | 83 | ALARM_ID_HD_SYRINGE_PUMP_RUNNING_WHILE_BP_OFF_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump is running while the BP is off
x | PRS 863 | SA 286 |  | SRSDG 815 |  | DG | DG POST: HD RTC Time |  | Call service to schedule a repair. |  | DG RTC set year is invalid | If DG RTC set year is < 2021. | 84 | ALARM_ID_DG_SET_RTC_YEAR_INVALID |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG set RTC year is invalid
revisit - PRS has this as low priority - why did SRS make this high priority? Probably needs to be high priority as blood pump needs to stop - update PRS to make high priority | [5 empty] | HD | Blood Stop: Cartridge Lever Open |  | Cartridge Lock Lever is not closed.\n- Open front door and push lever all the way down.\n- Close door and press RESUME button to continue. |  | HD pump track lever is not fully latched | If the pump track lever is detected as unlatched when it should be. | 85 | ALARM_ID_HD_PUMP_TRACK_LATCH_OPENED |  | ALARM_PRIORITY_HIGH | 209 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD pump track latch opened alarm
x | PRS 398 | SA 199 | SRSHD 1479 |  |  | HD | HD POST: HD RTC Time |  | Call service to schedule a repair. |  | HD RTC set year is invalid | If HD RTC set year is < 2021. | 86 | ALARM_ID_HD_SET_RTC_YEAR_INVALID |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD set RTC year is invalid
x | PRS 863 | SA 369 |  | SRSDG 929 |  | DG | DG POST: Heat Cal |  | Call service to schedule a repair. |  | DG heating parameters invalid calibration record | If the heating paramters' non-volatile record's CRC does not match the calculated CRC. | 87 | ALARM_ID_DG_HEATING_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG heating invalid calibration record
[6 empty] | DG | DG Fault: CP1 CP2 Speed Range |  | Call service to schedule a repair. |  | DG concentrate pump hall sensor out of range. | Hall sensor pulse to pulse timing is < 1ms (speed too fast). | 88 | ALARM_ID_DG_CONC_PUMP_HALL_SENSOR_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG concentrate pumps hall sensor out of range
x | PRS 863 | SA 365 |  | SRSDG 599 |  | DG | DG Fault: UVo On No Flow |  | Call service to schedule a repair. |  | DG outlet UV reactor is on while there is no flow | If the outlet UV reactor is on while the RO flow is less than 200 mL/min. | 89 | ALARM_ID_DG_OUTLET_UV_REACTOR_ON_WITH_NO_FLOW |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG outlet UV reactor on with no flow
x | PRS 863 | SA 307 SA 308 |  | SRSDG 604 |  | DG | DG Fault: LC A1/B1 Comm |  | Call service to schedule a repair. |  | DG load cells A1/B1 FPGA fault | If the FPGA does not report fresh data for a certain period of time. | 90 | ALARM_ID_DG_LOAD_CELL_A1_B1_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells A1/B1 FPGA fault
x | PRS 400 |  | SRSHD 1408 |  |  | HD | Loss of DG Comm (FMD) |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. |  | HD No dialysate flow data received in the last 3 seconds | If not receiving new dialysate flow readings in expected interval (3 seconds) | 91 | ALARM_ID_HD_DIALYSATE_FLOW_DATA_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | HD No dialysate flow data receive in the last 3 seconds
Remove (redundant with SRSDG 913) | [12 empty] | 92
[13 empty] | 92 | ALARM_ID_AVAILABLE_1 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Available for use
x | PRS 863 | SA 380 |  | SRSDG 430 |  | DG | DG Fault: RO Permeate Cond High |  | Call service to schedule a repair. |  | DG outlet primary conductivity sensor out of range | If the outlet primary conductivity sensor, CPo is greater than 30 uS/cm for a certain period of time. | 93 | ALARM_ID_DG_OUTLET_PRIMARY_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG outlet primary conductivity out of range
x | PRS 863 | SA 303 SA 304 SA 305 SA 306 |  | SRSDG 629 |  | DG | DG Fault: Pressure Sensor Range |  | Call service to schedule a repair. |  | DG pressure sensor out of range | If either of the pressure sensors is < 0 psia or > 200 psia for DG pressure sensors. | 94 | ALARM_ID_DG_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG pressure out of range
x | PRS 863 | SA 486 |  | SRSDG 907 |  | DG | DG Fault: WD Error |  | Call service to schedule a repair. |  | DG watchdog expired | If the watchdog signal is low while not in POST. | 95 | ALARM_ID_DG_WATCHDOG_EXPIRED |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG watchdog expired
x | PRS 406 |  |  | SRSDG 410 |  | DG | Dial. Gen: Inlet Water Temp High |  | Inlet water temperature is high.\n- Treatment is paused\n- Check the incoming water supply. |  | DG inlet water temperature out of range | If the inlet water temperature is > 37 C. | 96 | ALARM_ID_DG_INLET_WATER_TEMPERATURE_IN_HIGH_RANGE |  | ALARM_PRIORITY_MEDIUM | 432 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water temperature in high range
Check that software behavior and PRS behavior are aligned | PRS 543 |  |  | SRSDG 397 |  | DG | Dial. Gen: Dialysate Cond. |  | Dialysate conductivity is out of range.\n- Check acid and bicarb concentrate is correct and not empty. |  | DG fill conductivity out of range | If the average acid or bicard conductivity is > 5% of the normal acid and bicarb conductivity values. | 97 | ALARM_ID_DG_FILL_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 420 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | DG fill conductivity out of range
x | PRS 872 | SA 227 SA 228 | SRSHD 1395 |  |  | HD | HD Alarm (Low): Battery Com |  | Call service to schedule a repair. |  | HD battery or battery charger communication fault | If battery I2C communication fault persistence count > 5. | 98 | ALARM_ID_HD_BATTERY_COMM_FAULT |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | HD battery communication fault
x | PRS 494 | SA 274 | SRSHD 1683 |  |  | HD | HD Fault: Syringe Pump Stall |  | Call service to schedule a repair. |  | HD syringe pump stall alarm | If syringe pump measured rate < stall speed threshold (0.05) and the ramp up has been restarted 3 times after stall has been detected. | 99 | ALARM_ID_HD_SYRINGE_PUMP_STALL |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump stall alarm
x | PRS 494 | SA 416 | SRSHD 1679 |  |  | HD | HD Fault: Timeout No Cart ST |  | Call service to schedule a repair. |  | HD no cartridge self-test timeout | If during the no cartridge self-tests state, the elapsed self-test time > maximum no cartridge self test time (30secs) | 100 | ALARM_ID_HD_NO_CART_SELF_TEST_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD no cartridge self-test timeout
[6 empty] | UI | HD POST: UI-Proc Comm |  | Call service to schedule a repair. |  | UI POST FW Communication failure | If the HD is not communicating with the UI during POST. | 101 | ALARM_ID_UI_POST_HD_COMM |  | ALARM_PRIORITY_HIGH | 0 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | UI POST HD communication failure
cleaning | PRS 863 | SA 482 |  | SRSDG 856 SRSDG 930 |  | DG | Cleaning: Tpo THD Temp Delta |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. |  | DG heat disinfection temperature gradient out of range | If temperature gradient between Tpo and Thd exceeds maximum 15°C. | 102 | ALARM_ID_DG_HEAT_DISINFECT_TEMP_GRAD_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG heat disinfect temperature gradient out of range
x | PRS 398 | SA 196 | SRSHD 1059 |  |  | HD | HD POST: Cal NVRAM CRC |  | Call service to schedule a repair. |  | HD non-volatile calibration group invalid record CRC | If the computed CRC of the calibration group record is not equal to the record CRC. | 103 | ALARM_ID_HD_NVDATAMGMT_CAL_GROUP_RECORD_CRC_INVALID |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD invalid calibration CRC
x | PRS 494 | SA 478 | SRSHD 1115 |  |  | HD | HD Fault: Air Trap Sensor |  | Call service to schedule a repair. |  | Air trap level sensors reporting illegal combination of air/fluid | If the air trap level readings are not valid for more than 2 seconds. | 104 | ALARM_ID_HD_AIR_TRAP_ILLEGAL_LEVELS |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD air trap level sensors reporting illegal combination of air/fluid
x | PRS 863 | SA 283 |  | SRSDG 608 |  | DG | DG POST: Cal NVRAM CRC |  | Call service to schedule a repair. |  | DG non-volatile system group invalid record CRC | If the DG NVDATAMGMT CAL group calculated CRC does not mached the expected CRC. | 105 | ALARM_ID_DG_NVDATAMGMT_CAL_GROUP_RECORD_CRC_INVALID |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid calibration CRC
x | PRS 863 | SA 369 |  | SRSDG 820 SRSDG 821 SRSDG 1087 |  | DG | DG POST: Flow Sensor Cal |  | Call service to schedule a repair. |  | DG flow sensors calibration error | If the calculated CRC of the flow sensors calibration record does not match the calibration records' CRC. | 106 | ALARM_ID_DG_FLOW_SENSORS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG dialysate flow sensor invalid calibration record
No SRS ID - need to add then assess Move to DG machine alarm? (allow properties table to be aligned) Sean - to revisit code to assess | [5 empty] | HD | DG Fault: DG Restarted |  | Call service to schedule a repair. |  | HD reports DG restarted fault | If DG reports being in any mode other than idle, drain, fill or fault during pre-tx, tx, or post-tx modes - means DG restarted unexpectedly and is no longer in sync w/ HD. | 107 | ALARM_ID_HD_DG_RESTARTED_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD reports DG restarted fault
x | PRS 494 | SA 276 | SRSHD 1318 |  |  | HD | HD Fault: Syringe Pump ADC Com |  | Call service to schedule a repair. |  | HD syringe pump ADC error | If the ADC read is not fresh after POST init and the ADC read counter has not changed. | 108 | ALARM_ID_HD_SYRINGE_PUMP_FPGA_ADC_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump ADC FPGA fault
x | PRS 494 | SA 490 | SRSHD 580 |  |  | HD | HD Fault: Syringe Pump Vol |  | Call service to schedule a repair. |  | HD syringe pump volume check error | If the volume delivered is off by more than 5% or 0.1mL. | 109 | ALARM_ID_HD_SYRINGE_PUMP_VOLUME_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump volume check error
x | PRS 494 | SA 274 | SRSHD 579 |  |  | HD | HD Fault: Syringe Pump Speed |  | Call service to schedule a repair. |  | HD syringe pump speed check error | If the syringe measured rate is off by more than 5% or 0.1mL. | 110 | ALARM_ID_HD_SYRINGE_PUMP_SPEED_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump speed check error
x | PRS 494 | SA 274 | SRSHD 1333 |  |  | HD | HD Fault: Syringe Pump Speed |  | Call service to schedule a repair. |  | HD syringe pump not stopped in off state error | If the syringe pump position is not equal to its last position. | 111 | ALARM_ID_HD_SYRINGE_PUMP_NOT_STOPPED_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump not stopped in off state error
x | PRS 863 | SA 312 SA 313 |  | SRSDG 953 |  | DG | DG Fault: LC A2/B2 Comm |  | Call service to schedule a repair. |  | DG load cells A2/B2 FPGA fault | If the FPGA does not report fresh data for a certain period of time. | 112 | ALARM_ID_DG_LOAD_CELL_A2_B2_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells A1/B1 FPGA fault
x | PRS 863 | SA 341 |  | SRSDG 890 |  | DG | DG Fault: DRP Current |  | Call service to schedule a repair. |  | DG drain pump current out of range | If the DG drain pump current > 2.2 A when it is in open loop or if the current > 0.1 A when it is off. | 113 | ALARM_ID_DG_DRAIN_PUMP_CURRENT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG Drain pump current out of range
move to wet self-test? | PRS 494 | SA 479 | SRSHD 1532 |  |  | HD | Process: Pre-Tx Wet Self-Test |  | Self-test has detected a problem with the air bubble sensor\n- Check cartridge installation on to machine\n- If problem persists, terminate pre-treatment and \nrestart with a new cartridge. |  | HD venous air bubble detector self-test failure | If the venous air bubble self test did not detect bubble for more than 500 msec. | 114 | ALARM_ID_HD_VENOUS_BUBBLE_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD venous air bubble detector self-test failure
Sean - update code to remove temperatures from code that are no longer needed (see updated SRSDG 224) | PRS 863 | SA 321 SA 322 SA 323 SA 324 SA 328 SA 332 SA 336 SA 298 SA 460 |  | SRSDG 224 |  | DG | DG Fault: Temp Sensor Range |  | Call service to schedule a repair. |  | DG temperature sensor out of range | If the measured temperature < 0 C or > 80 C for the non-fluid inline sensors and > 120 C for the fluid sensors. | 115 | ALARM_ID_DG_TEMPERATURE_SENSOR_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG temperature sensor out of range
[6 empty] | UI HD | HD Alarm (Low): SD-Card Error |  | SD-Card Failure \n- Power cycle the system.\n- Call service to schedule a repair if problem persists. |  | SD card is non-operational | SD card is non-operational | 116 | ALARM_ID_HD_UI_SDCARD_FAILURE |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD UI SD card failure
|  |  |  | SRSDG 909 |  | DG | Loss of HD Comm (HD ACK) |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. |  | System message that required acknowledgment was not acknowledged | If no retries left in the pending acknowlegement list. | 117 | ALARM_ID_DG_CAN_MESSAGE_NOT_ACKED |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CAN message not acked
x | PRS 863 | SA 458 |  | SRSDG 43 |  | DG | DG Fault: RTC Config |  | Call service to schedule a repair. |  | DG RTC configuration error | If the RTC configuration registers report an error. | 118 | ALARM_ID_DG_RTC_CONFIG_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RTC config error
x | PRS 874 | SA 499 |  | SRSDG 908 |  | DG | DG Alarm (Low): RTC Bat Low |  | Call service to schedule a repair. |  | DG RTC battery low | If the RTC battery bit report a low alarm. | 119 | ALARM_ID_DG_RTC_BATTERY_LOW |  | ALARM_PRIORITY_LOW | 851 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | DG RTC battery low error
|  |  | SRSHD 777 |  |  | HD | Process: Pre-Tx Wet Self-Test |  | Self-test has detected a bubble in the venous bubble detector\n- The cartridge should be primed\n- Check cartridge installation on to machine\n- If problem persists, terminate pre-treatment and \nrestart with a new cartridge. |  | HD pre-treatment wet prime test failure | If the HD pre-treatment wet self test failed. | 120 | ALARM_ID_HD_PRE_TREATMENT_WET_PRIME_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode wet self-test prime check failure
x | PRS 863 | SA 363 |  | SRSDG 611 |  | DG | DG Fault: Prim Heater Lg Control |  | Call service to schedule a repair. |  | DG main primary heater voltage out of range | If the main primary heater's voltage is out of range. | 121 | ALARM_ID_DG_MAIN_PRIMARY_HEATER_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG main primary heater voltage out of range
x | PRS 863 | SA 362 |  | SRSDG 612 |  | DG | DG Fault: Prim Heater Sm Control |  | Call service to schedule a repair. |  | DG small primary heater voltage out of range | If the small primary heater's voltage is out of range. | 122 | ALARM_ID_DG_SMALL_PRIMARY_HEATER_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG small primary heater voltage out of range
x | PRS 863 | SA 361 |  | SRSDG 613 |  | DG | DG Fault: Trimmer Heater Control |  | Call service to schedule a repair. |  | DG trimmer heater voltage out of range | If the trimmer voltage is out of range. | 123 | ALARM_ID_DG_TRIMMER_HEATER_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG trimmer heater voltage out of range
x | PRS 390 |  | SRSHD 954 |  |  | HD | Blood Stop: Blood Clot Risk |  | Blood pump has been off for too long.\n- Treatment must be terminated. Blood clot risk.\n- NO RINSEBACK ALLOWED. |  | HD end of treatment alarm (high priority - user still not acting to end treatment) | In treatment end paused state for > 5 minutes. | 124 | ALARM_ID_HD_END_OF_TREATMENT_HIGH |  | ALARM_PRIORITY_HIGH | 11 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD end of treatment alarm (high priority - user still not acting to end treatment)
x | PRS 390 |  | SRSHD 948 |  |  | HD | Blood Stop: Blood Clot Risk |  | Blood pump has been off for too long.\n- Treatment must be terminated. Blood clot risk.\n- NO RINSEBACK ALLOWED. |  | Blood sitting too long after treatment stopped by user alarm (> 5 min) | If in treatment stop sub mode when blood is in the tubing, the blood sitting counter exceeded 5 minutes. | 125 | ALARM_ID_HD_TREATMENT_STOPPED_NO_RINSEBACK |  | ALARM_PRIORITY_HIGH | 11 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood sitting too long after treatment stopped by user alarm (>5 min)
x | PRS 387 |  | SRSHD 925 |  |  | HD | Blood Stop: Dialyzer Blood Leak |  | A blood leak has been detected. Treatment must end. |  | HD blood leak detected alarm | If the blood leak is detected based on the blood leak status reported by the sensor. | 126 | ALARM_ID_HD_BLOOD_LEAK_DETECTED |  | ALARM_PRIORITY_HIGH | 150 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | HD blood leak detected alarm
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 911 |  |  | HD | Blood Stop: Venous Pres. Low |  | Venous pressure is low.\n- Check to be sure your blue venous connections are secure.\n- Check that your venous access is connected and secured \nproperly. |  | Venous pressure too low during treatment | If the venous pressure < TREATMENT_PARAM_VEN_PRESSURE_LOW_LIMIT. Or expanded pressure limits based on mode of operation. | 127 | ALARM_ID_HD_VENOUS_PRESSURE_LOW |  | ALARM_PRIORITY_HIGH | 200 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | HD venous pressure too low during treatment
x | PRS 494 | SA 506 | SRSHD 919 |  |  | HD | Blood Stop: Venous Air Bubble |  | Venous air bubble detected indicating failure of air trap \ncontrol system.\n- Treatment must be terminated without option for rinseback.\n- Check access for dislodgement. |  | HD venous air bubble detected alarm | If bubbles are detected during the treatment. | 128 | ALARM_ID_HD_VENOUS_BUBBLE_DETECTED |  | ALARM_PRIORITY_HIGH | 12 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | HD venous air bubble detected alarm
x | PRS 494 | SA 507 | SRSHD 921 | [9 empty] | 129 | ALARM_ID_AVAILABLE_17 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | Available for use
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 913 |  |  | HD | Blood Stop: Venous Pres. High |  | Venous pressure is high.\n- Check for access or patient discomfort.\n- Resolve access issue if applicable.\n- Check that there are no kinks or clots in the tubing.\n- Unkink the lines if applicable.\n- End treatment and do not return blood if clots are present. |  | Venous pressure too high during treatment | If the venous pressure > TREATMENT_PARAM_VEN_PRESSURE_HIGH_LIMIT. Or expanded pressure limits based on mode of operation. | 130 | ALARM_ID_HD_VENOUS_PRESSURE_HIGH |  | ALARM_PRIORITY_HIGH | 201 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | HD venous pressure too high during treatment
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 915 |  |  | HD | Blood Stop: Arterial Pres. Low |  | Arterial pressure is low.\n- Check that the access is properly connected and secured.\n- Resolve access issue if applicable.\n- Check to be sure the red arterial connections are secure.\n- Check that there are no kinks or clots in the tubing.\n- Unkink lines if applicable.\n- End treatment and do not return blood if clots are present. |  | Arterial pressure too low during treatment | If the arterial pressure < TREATMENT_PARAM_ART_PRESSURE_LOW_LIMIT. Or expanded pressure limits based on mode of operation. | 131 | ALARM_ID_HD_ARTERIAL_PRESSURE_LOW |  | ALARM_PRIORITY_HIGH | 202 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | HD arterial pressure too low during treatment
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 917 |  |  | HD | Blood Stop: Arterial Pres. High |  | Arterial pressure is high.\n- Check that the access is properly connected and secured.\n- Resolve access issue if applicable.\n- Check to be sure your red arterial connections are secure.\n- Check that there are no kinks or clots in the tubing.\n- Unkink lines if applicable.\n- End treatment and do not return blood if clots are present. |  | Arterial pressure too high during treatment | If the arterial pressure > TREATMENT_PARAM_ART_PRESSURE_HIGH_LIMIT. Or expanded pressure limits based on mode of operation. | 132 | ALARM_ID_HD_ARTERIAL_PRESSURE_HIGH |  | ALARM_PRIORITY_HIGH | 203 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | HD arterial pressure too high during treatment
x | PRS 863 | SA 419 |  | SRSDG 328 |  | DG | DG Fault: Fluid Leak in DG |  | Fluid leak detected in Dialysate Generation Delivery Device.\n-Treatment must be terminated.\n- Shut off water source and disconnect water and \nelectrical connections from system.\n- Call service to schedule a repair. |  | DG fluid leak detected alarm | If the fluid leak is detected as reported by the fluid leak detector. | 133 | ALARM_ID_DG_FLUID_LEAK_DETECTED |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | DG fluid leak detected alarm
x | PRS 409 |  | SRSHD 927 |  |  | HD | Blood Stop: Fluid Leak in HD |  | Fluid leak detected in Hemodialysis Delivery device.\n- If in treatment, BLOOD STOPPED.\n-Identify and correct the cause of the spill.\n- Clean up the spill and dry leak sensor.\n- If in treatment, resume treatment. |  | HD fluid leak detected alarm | If the fluid leak is detected as reported by the fluid leak detector. | 134 | ALARM_ID_HD_FLUID_LEAK_DETECTED |  | ALARM_PRIORITY_HIGH | 208 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | HD fluid leak detected alarm
x | PRS 411 |  | SRSHD 938 |  |  | HD | Blood Stop: HD Shock Detected |  | Excessive movement detected with Hemodialysis Delivery \ndevice.\n- Confirm the device is stationary and level before continuing. |  | HD experienced large acceleration | If the system experienced excessive shock > 2.5 g. | 135 | ALARM_ID_HD_SHOCK |  | ALARM_PRIORITY_HIGH | 210 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD experienced large acceleration
x | PRS 686 |  |  | SRSDG 327 |  | DG | Dial. Gen: Shock-Vibration |  | Excessive movement detected with Dialysis Generation \ndevice.\n- Treatment is paused.\n- Confirm the device is stationary and level before continuing. |  | DG experienced large acceleration | If the system experienced excessive shock > 2.5 g. | 136 | ALARM_ID_DG_SHOCK |  | ALARM_PRIORITY_MEDIUM | 440 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG experienced large acceleration
x | PRS 412 |  | SRSHD 936 |  |  | HD | Blood Stop: HD Tilted |  | Excessive tilt detected with Hemodialysis Delivery device.\n- Confirm the device is stationary and level before continuing. |  | HD tilt exceeds maximum | If the system experienced excessive tilt x and y axis angle > 7 deg for more than 1 sec. | 137 | ALARM_ID_HD_EXCESSIVE_TILT |  | ALARM_PRIORITY_HIGH | 211 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD tilt exceeds maximum
x | PRS 687 |  |  | SRSDG 326 |  | DG | Dial. Gen: Tilt |  | Excessive tilt detected with Dialysis Generation device.\n- Treatment is paused.\n-Confirm the device is stationary and level before continuing. |  | DG tilt exceeds maximum | If the system experienced excessive tilt x and y axis angle > 7 deg for more than 1 sec. | 138 | ALARM_ID_DG_EXCESSIVE_TILT |  | ALARM_PRIORITY_MEDIUM | 441 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG tilt exceeds maximum
x | PRS 823 |  | SRSHD 1593 |  |  | HD | Loss of AC Power |  | Power has been lost. Check all electrical connections.\n- The display and CPU will run off of the battery until \npower returns. |  | HD AC power lost alarm | The HD Software shall trigger an alarm if an AC power loss is detected while the HD system is in standby, service, or treatment parameters mode. | 139 | ALARM_ID_HD_AC_POWER_LOST |  | ALARM_PRIORITY_LOW | 852 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD AC power lost alarm
x | PRS 400 |  | SRSHD 1124 |  |  | HD | Loss of DG Comm (Timeout) |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. |  | Dialysate generator communication timeout | If DG communication timeout is detected for more than 2 seconds. | 140 | ALARM_ID_HD_DG_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD DG communication timeout
|  |  | SRSHD 1169 |  |  | HD | Blood Stop: Air Trap Fill |  | Failure to fill the Air Trap.\n- Treatment must be terminated without option for rinseback.\n- Check access for dislodgement. |  | Air trap fill timeout during treatment | If the air trap fill timeout during treatment. | 141 | ALARM_ID_HD_AIR_TRAP_FILL_DURING_TREATMENT |  | ALARM_PRIORITY_HIGH | 13 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | HD air trap fill timeout during treatment
x | PRS 496 |  | SRSHD 1361 SRSHD 1362 |  |  | HD | Blood Stop: Blood Pump Occlusion |  | Occlusion detected after the blood pump.\n- Check for clots or kinks around the dialyzer and in the \nvenous line.\n- Resolve access issue if applicable.\n- End treatment and do not return blood if clots are present.\n- Check for access or patient discomfort.\n- Unkink the lines if applicable. |  | Blood pump occlusion detected | If blood occlusion is detected > OCCLUSION_THRESHOLD_OFFSET + bloodPumpOcclusionAfterCartridgeInstall | 142 | ALARM_ID_HD_OCCLUSION_BLOOD_PUMP |  | ALARM_PRIORITY_HIGH | 206 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | HD blood pump occlusion detected
x | PRS 863 | SA 338 |  | SRSDG 912 |  | DG | DG Fault: TRo TDi Delta |  | Call service to schedule a repair. |  | DG dialysate temperature sensors drift timeout. | If TDi and TRo have drifted for more 1 C for a certain period of time. | 143 | ALARM_ID_DG_DIALYSATE_TEMPERATURE_SENSORS_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG dialysate temperature sensors out of range
x | PRS 853 |  |  | SRSDG 1066 SRSDG 1054 SRSDG 970 |  | DG | Cleaning: Inlet Water Cond High |  | Cleaning Mode Inlet water conductivity too high.\n- Cleaning mode process has terminated \n- Check inlet water before restarting cleaning mode. |  | DG cleaning mode inlet water conductivity out of high range | Inlet water conductivity > 2000 uS/cm. Inlet water conductivity > 100 uS/cm (PURE WATER MODE). In cleaning mode | 144 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_COND_TOO_HIGH |  | ALARM_PRIORITY_MEDIUM | 603 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water conductivity out of high range.
x | PRS 394 |  |  | SRSDG 846 |  | DG | Dial. Gen: Acid Conc. Cond. |  | Acid concentrate conductivity during prime is out of range.\n- Check correct acid concentrate is connected and full.\n- Resume to reprime concentrate. |  | DG Concentrate conductivity after adding acid out of range alarm | If the conductivity value for the acid concentrate is out of range for more than 45 secs. | 145 | ALARM_ID_DG_ACID_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 410 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG concentrate conductivity after adding acid out of range alarm
x | PRS 863 | SA 285 |  | SRSDG 312 |  | DG | DG POST: DG RTC Timer |  | Call service to schedule a repair. |  | DG RTC vs. Timer accuracy error | If RTC and timer are not aligned on 1 second elapsed time. | 146 | ALARM_ID_DG_RTC_OR_TIMER_ACCURACY_FAILURE |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RTC or timer accuracy failure.
x | PRS 543 |  |  | SRSDG 398 |  | DG | Dial. Gen: Dialysate not ready |  | Dialysate is not yet ready.\n- Treatment is paused and will be ready to resume when \ndialysate is ready for use. |  | DG dialysate generation conductivity fault alarm | If dialysate conductivity is not with 5% of the TCD | 147 | ALARM_ID_DG_CREATING_DIALYSATE_PLEASE_WAIT |  | ALARM_PRIORITY_MEDIUM | 421 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | DG wait for the DG to produce dialysate
x | PRS 557 |  |  | SRSDG 1065 SRSDG 1053 SRSDG 969 |  | DG | Cleaning: Inlet Water Temp High |  | Cleaning Mode Inlet water temperature too high.\n- Cleaning mode process has terminated.\n- Check inlet water before restarting cleaning mode. |  | Cleaning Mode Inlet water temperature too high | Inlet temperature > 37 C. | 148 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_TEMP_TOO_HIGH |  | ALARM_PRIORITY_MEDIUM | 601 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG Cleaning mode inlet water temperature too high
x | PRS 494 | SA 475 | SRSHD 646 |  |  | HD | HD Fault: WD Error |  | Call service to schedule a repair. |  | HD watchdog expired | If the watchdog signal is low while not in POST. | 149 | ALARM_ID_HD_WATCHDOG_EXPIRED |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD watchdog expired
x | PRS 404 |  |  | SRSDG 391 |  | DG | Dial. Gen: Inlet Water Cond High |  | Inlet water conductivity is high.\n- Treatment is paused.\n- Check the incoming water supply. |  | Inlet water conductivity too high | Inlet water conductivity > 2000 uS/cm. Inlet water conductivity > 100 uS/cm (PURE WATER MODE). | 150 | ALARM_ID_DG_INLET_WATER_CONDUCTIVITY_IN_HIGH_RANGE |  | ALARM_PRIORITY_MEDIUM | 434 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water conductivity in high range
x | PRS 403 |  |  | SRSDG 826 |  | DG | Dial. Gen: Inlet Water Cond Low |  | Inlet water conductivity is low.\n- Treatment is paused.\n- Check the incoming water supply. |  | Inlet water conductivity too low | Inlet water conductivity < 200 uS/cm. Inlet water conductivity low disabled (PURE WATER MODE). | 151 | ALARM_ID_DG_INLET_WATER_CONDUCTIVITY_IN_LOW_RANGE |  | ALARM_PRIORITY_MEDIUM | 435 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water conductivity in low range
x | PRS 401 |  |  | SRSDG 417 |  | DG | Dial. Gen: Inlet Water Pres Low |  | Inlet water pressure is low.\n- Treatment is paused.\n- Check the incoming water supply. |  | Inlet water pressure too low | Inlet water pressure < 25 PSI. | 152 | ALARM_ID_DG_INLET_WATER_PRESSURE_IN_LOW_RANGE |  | ALARM_PRIORITY_MEDIUM | 431 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water pressure in low range
x | PRS 529 |  | SRSHD 1122 |  |  | HD | Process: Pre-Tx Terminated |  | Pre-Treatment timeout.\n- Patient was not connected to device\n- Pre-Treatment has been terminated\n- Proceed to end treatment to remove cartridge from system. |  | HD prime completed high priority alarm | In Pre-Treatment Recirculation for 120 minutes (escalation of alarm 213) | 153 | ALARM_ID_HD_PRIME_COMPLETED_HIGH |  | ALARM_PRIORITY_HIGH | 292 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD prime completed high priority alarm
x | PRS 854 |  |  | SRSDG 1067 SRSDG 1055 SRSDG 971 |  | DG | Cleaning: Inlet Water Cond Low |  | Cleaning Mode Inlet water conductivity too low.\n- Cleaning mode process has terminated \n- Check inlet water before restarting cleaning mode. |  | Cleaning mode inlet water conductivity too low | Inlet water conductivity < 200 uS/cm. Inlet water conductivity low disabled (PURE WATER MODE). In cleaning mode | 154 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_COND_TOO_LOW |  | ALARM_PRIORITY_MEDIUM | 604 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water conductivity out of low range.
x | PRS 400 |  | SRSHD 1574 |  |  | HD | Loss of DG Comm (Reservoir) |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. |  | HD No reservoir data message receive at least once every 2 seconds | No new reservoir data for 2 seconds. | 155 | ALARM_ID_HD_NEW_RESERVOIRS_DATA_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | HD No reservoirs data message receive at least once every 2 seconds
x | PRS 400 |  | SRSHD 1575 |  |  | HD | Loss of DG Comm (Op Mode) |  | Communication failure with Dialysis Generator.\n-Treatment must be terminated.\n- Ensure all cords are securely connected. |  | HD No DG op mode message receive at least once every 2 seconds | No DG op mode for 2 seconds. | 156 | ALARM_ID_HD_DG_NEW_OPERATION_MODE_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | HD No DG operation mode message receive at least once every 2 seconds
cleaning |  |  |  | SRSDG 988 |  | DG | Cleaning: Chem Disinf PAA Prime |  | Unable to sense proper chemical disinfectant.\n- Cleaning mode process has terminated.\n- Check that proper chemical disinfectant is connected to the \nbicarb port.\n- Check chemical disinfectant bottle is not empty. |  | Chemical disinfect priming the acid line timeout | If the disinfectant acid line is not primed within a certain period of time. | 157 | ALARM_ID_DG_CHEM_DISINFECT_PRIME_ACID_LINE_TIME_OUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG chemical disinfect prime acid line timeout.
x | PRS 405 |  |  | SRSDG 408 |  | DG | Dial. Gen: Inlet Water Temp Low |  | Inlet water temperature is low.\n- Treatment is paused.\n- Check the incoming water supply. |  | Inlet water temperature low | Inlet temperature < 24 °C for 5 seconds. | 158 | ALARM_ID_DG_INLET_WATER_TEMPERATURE_IN_LOW_RANGE |  | ALARM_PRIORITY_MEDIUM | 433 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water temperature is in the low range
cleaning |  |  |  | SRSDG 1044 |  | DG | Cleaning: Chem Disinf Connect Cap |  | Remove chemical disinfectant connection and close cap. |  | Close concentrate line be able to flush it | Start of chemical disinfect flush. | 159 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_REMOVE_ACID |  | ALARM_PRIORITY_LOW | 751 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG chemical disinfect flush remove acid and close the concentrate cap.
x | PRS 494 | SA 239 | SRSHD 972 |  |  | HD | HD Fault: BP Rotor Max Speed |  | Call service to schedule a repair. |  | Blood pump rotor speed too high | If blood pump motor speed exceeded 100.0 RPM. | 160 | ALARM_ID_HD_BLOOD_PUMP_ROTOR_SPEED_TOO_HIGH |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood pump rotor speed too high
x | PRS 863 | SA 497 |  | SRSDG 122 |  | DG | DG Fault: HD-DG Comm |  | Call service to schedule a repair. |  | Too many bad CRCs detected on received system messages | If >= 5 bad CRCs detected on received CAN messages within 10 minutes. | 161 | ALARM_ID_DG_COMM_TOO_MANY_BAD_CRCS |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG comm too many bad CRCs
x | PRS 863 | SA 300 |  | SRSDG 867 |  | DG | DG Fault: Proc-FPGA Clock |  | Call service to schedule a repair. |  | FPGA clock speed failure | If the FPGA clock speed is relatively close to the processor clock for a certain period of time. | 162 | ALARM_ID_DG_FPGA_CLOCK_SPEED_CHECK_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG FPGA clock speed check failure.
revisit - determine if this is blood stopped or dialysate may need to add to PRS |  |  | SRSHD 1260 |  |  | HD | Dial. Gen: UF Rate Check |  | Ultrafiltration control detected and incorrect rate.\n- This can be caused by excessive movement of system\n(shock/vibration)\n- Check system is stable and level\n- Resume and If problem persists, call service to schedule a\nrepair. |  | HD sees primary load cell for reservoir 1 change too much too fast | If the delta volume of the active reservoir exceeded 50.0 g (MAX_ACTIVE_LOAD_CELL_CHANGE_G). | 163 | ALARM_ID_HD_LOAD_CELL_ACCELERATION_RES_1_ALARM |  | ALARM_PRIORITY_MEDIUM | 442 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD sees primary load cell for reservoir 1 change too much too fast
revisit - determine if this is blood stopped or dialysate may need to add to PRS |  |  | SRSHD 1601 |  |  | HD | Dial. Gen: UF Rate Check |  | Ultrafiltration control detected and incorrect rate.\n- This can be caused by excessive movement of system \n(shock/vibration).\n- Check system is stable and level\n- Resume and If problem persists, call service to schedule a \nrepair. |  | HD sees primary load cell for reservoir 2 change too much too fast | If the delta volume of the active reservoir exceeded 50.0 g (MAX_ACTIVE_LOAD_CELL_CHANGE_G). | 164 | ALARM_ID_HD_LOAD_CELL_ACCELERATION_RES_2_ALARM |  | ALARM_PRIORITY_MEDIUM | 442 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD sees primary load cell for reservoir 2 change too much too fast
sean remove from code 15min limint (SRS 1556 has been updated) |  |  | SRSHD 1228 SRSHD 1399 SRSHD 1556 SRSHD 1709 |  |  | HD | Blood Stop: Blood Clot Risk |  | Blood pump has been off for too long.\n- Treament must be terminated. Blood clot risk.\n- NO RINSEBACK ALLOWED. |  | HD in treatment rinseback operation timeout | If the maximum time inrinseback operation has been exceeded. | 165 | ALARM_ID_HD_TREATMENT_RINSEBACK_TIMEOUT_ALARM |  | ALARM_PRIORITY_HIGH | 11 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD in treatment rinseback operation timeout
[13 empty] | 166 | ALARM_ID_AVAILABLE_3 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | Available for use
needs to be high priority due to blood stopped | [5 empty] | HD | Blood Stop: Door Open |  | Device door is open.\n - Close door and press RESUME button to continue. |  | HD cartridge door opened alarm | If the cartridge door is opened as reported by the FPGA. | 167 | ALARM_ID_HD_CARTRIDGE_DOOR_OPENED |  | ALARM_PRIORITY_HIGH | 209 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD cartridge door opened alarm
[13 empty] | 168 | ALARM_ID__AVAILABLE_70 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | Available for use
x | PRS 863 | SA 493 |  | SRSDG 643 |  | DG | DG Fault: Dialysate Flow Sens Rg |  | Call service to schedule a repair. |  | Dialysate flow rate out of range | If the dialysate flow rate > 2.00 L/min. | 169 | ALARM_ID_DG_DIALYSATE_FLOW_RATE_OUT_OF_MAX_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG Dialysate flow rate out of maximum range
Currently in CBIT section - should be somewhere else? | [5 empty] | HD | Heparin Syringe Empty |  | Heparin syringe empty.\n- Treatment will continue\n- If additional heparin is required to complete treatment, \nadminister manually. |  | HD syringe empty alarm | If the syringe pump position is near empty position. | 170 | ALARM_ID_HD_SYRINGE_PUMP_SYRINGE_EMPTY |  | ALARM_PRIORITY_LOW | 730 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | HD syringe empty alarm
|  |  | SRSHD 1525 |  |  | HD | Process: Pre-Tx Hep. Pump Occ. |  | Heparin pump occlusion detected.\n- Check heparin line is unclamped.\n- Check there are no kinks in the heparin line. |  | HD syringe pump occlusion alarm | If the pressure sensor voltage reading > SYRINGE_FORCE_OCCLUSION_THRESHOLD_V. | 171 | ALARM_ID_HD_SYRINGE_PUMP_OCCLUSION |  | ALARM_PRIORITY_LOW | 716 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD syringe pump occlusion alarm
|  |  | SRSHD 1383 |  |  | HD | Process: Pre-Tx Not Enough Hep. |  | Not enough heparin is available for treatment.\n- Check if syringe pump is touching syringe plunger.\n- If syringe pump is touching syringe plunger, end treatment.\n- If syringe pump is not touching syringe plunger, reload \nsyringe with enough heparin. |  | HD syringe pump not enough Heparin alarm | If the estimated syringe volume needed for treatment is not sufficient | 172 | ALARM_ID_HD_SYRINGE_PUMP_NOT_ENOUGH_HEPARIN_ALARM |  | ALARM_PRIORITY_LOW | 715 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD syringe pump not enough Heparin alarm
x | PRS 494 | SA 474 | SRSHD 659 |  |  | HD | HD Fault: RTC Config |  | Call service to schedule a repair. |  | HD RTC configuration error | If the RTC configuration registers report an error. | 173 | ALARM_ID_HD_RTC_CONFIG_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD RTC configuration error
x | PRS 398 | SA 198 | SRSHD 630 |  |  | HD | HD POST: HD RTC Timer |  | Call service to schedule a repair. |  | HD RTC vs. Timer accuracy error | If RTC and timer are not aligned on 1 second elapsed time. | 174 | ALARM_ID_HD_RTC_OR_TIMER_ACCURACY_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD RTC or timer accuracy failure
x | PRS 494 | SA 245 SA 254 SA 262 | SRSHD 1502 SRSHD 1503 SRSHD 1568 |  |  | HD | HD Fault: Pump Commutation Err |  | Call service to schedule a repair. |  | HD pump commutation error | If the set pump direction does not agree with the hall sensor direction. | 175 | ALARM_ID_HD_PUMP_DIRECTION_STATUS_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD pump direction status error
Dialysate Temp |  |  | SRSHD 895 | [9 empty] | 176 | ALARM_ID_AVAILABLE_4 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD dialysate temperature above high safety range
x | PRS 863 | SA 417 |  | SRSDG 919 |  | DG | DG Fault: Software Invalid State |  | Call service to schedule a repair. |  | DG software fault. Software found itself in an unexpected state | If the specific software faults listed in alarm management SW faults has been detected | 177 | ALARM_ID_DG_SOFTWARE_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG software fault. Software found itself in an unexpected state
x | PRS 494 | SA 225 |  |  | SRSUI 1066 | UI | HD Fault: UI-Proc Comm |  | Call service to schedule a repair. |  | HD communication timeout | If the HD stops communicating with the UI. | 178 | ALARM_ID_HD_COMM_TIMEOUT | OK | ALARM_PRIORITY_HIGH | 0 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD communication timeout
|  |  |  | SRSDG 605 |  | DG | DG Fault: FPGA Com |  | Call service to schedule a repair. |  | DG FPGA communication down for too long | If retries for commands exceeds limit or FPGA reports comm error. | 179 | ALARM_ID_DG_FPGA_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG FPGA communication down for too long
x | PRS 863 | SA 494 |  | SRSDG 883 |  | DG | DG Fault: RO Flow Sensor Range |  | Call service to schedule a repair. |  | RO flow rate out of range | If the RO flow rate > 2.00 L/min. | 180 | ALARM_ID_DG_RO_FLOW_RATE_OUT_OF_MAX_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO flow out of maximum range
|  |  |  | SRSDG 849 SRSDG 838 |  | DG | DG Fault: Tare Out of Range |  | Call service to schedule a repair. |  | DG load cells weight out of range for tare | If the load cell weight is out of range for tare. | 181 | ALARM_ID_DG_LOAD_CELLS_TARE_WEIGHT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells weight out of range for tare
x | PRS 863 | SA 369 |  | SRSDG 162 |  | DG | DG POST: LC Cal |  | Call service to schedule a repair. |  | DG load cells invalid calibration | If the CRC of load cells calibration data does not match the calculated CRC. | 182 | ALARM_ID_DG_LOAD_CELLS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells invalid calibration
cleaning |  |  |  | SRSDG 851 |  | DG | Cleaning: Invalid Resv Transfer |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. |  | DG invalid load cell value | If the half full reservoir is filled prior to filling the reservoir that is supposed to overflow to the half full reservoir. | 183 | ALARM_ID_DG_INVALID_LOAD_CELL_VALUE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG invalid load cell value
x | PRS 863 | SA 355 SA 356 |  | SRSDG 559 |  | DG | DG Fault: UVi Error |  | Call service to schedule a repair. |  | Inlet reactor not healthy | If the inlet reactor are not healthy. | 184 | ALARM_ID_DG_INLET_UV_REACTOR_NOT_HEALTHY |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG inlet UV reactor not healthy
x | PRS 874 | SA 299 |  | SRSDG 628 |  | DG | DG Alarm (Low): Fan Speed |  | Fan failure.\n- Check that the fans inlets are not blocked or obstructed. |  | DG fan RPM out of range | If the fan RPM is out of Min/Max range. | 185 | ALARM_ID_DG_FAN_RPM_OUT_OF_RANGE |  | ALARM_PRIORITY_LOW | 851 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | DG fan RPM out of range
x | PRS 863 | SA 351 SA 353 |  | SRSDG 891 |  | DG | DG Fault: CP1/CP2 Fault |  | Call service to schedule a repair. |  | DG concentrate pump fault | If the concentrate pumps fault is reported by the FPGA. | 186 | ALARM_ID_DG_CONCENTRATE_PUMP_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG concentrate pump fault
x | PRS 863 | SA 352 |  | SRSDG 905 |  | DG | DG Fault: CP1 Speed Fault |  | Call service to schedule a repair. |  | DG concentrate pump CP1 speed control error | If the concentrate pump CP1 speed > 0.02 tolerance for a certain period of time. | 187 | ALARM_ID_DG_CP1_SPEED_CONTROL_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG concentrate pump CP1 speed control error
x | PRS 863 | SA 354 |  | SRSDG 906 |  | DG | DG Fault: CP2 Speed Fault |  | Call service to schedule a repair. |  | DG concentrate pump CP2 speed control error | If the concentrate pump CP1 speed > 0.02 tolerance for a certain period of time. | 188 | ALARM_ID_DG_CP2_SPEED_CONTROL_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG concentrate pump CP2 speed control error
x | PRS 863 | SA 339 SA 343 |  | SRSDG 448 |  | DG | DG Fault: DRP Speed Fault |  | Call service to schedule a repair. |  | DG drain pump RPM out of range | If the drain pump RPM > 100.0 of the target RPM. | 189 | ALARM_ID_DG_DRAIN_PUMP_RPM_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain pump RPM out of range
[6 empty] | DG | DG Fault: DRP Speed Not Off |  | Call service to schedule a repair. |  | DG drain pump off fault | If the pump is in off state and the RPM is greater than the minimum RPM of 300.0 RPM. | 190 | ALARM_ID_DG_DRAIN_PUMP_OFF_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain pump off fault
x | PRS 863 | SA 498 |  | SRSDG 885 |  | DG | DG Fault: RO Flow High |  | Call service to schedule a repair. |  | DG flow rate out of upper range | If the RO flow rate > 10% from the target flow rate. | 191 | ALARM_ID_DG_FLOW_RATE_OUT_OF_UPPER_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG flow rate out of upper range
x | PRS 863 | SA 498 |  | SRSDG 945 |  | DG | DG Fault: RO Flow Low |  | Call service to schedule a repair. |  | DG flow rate out of lower range | If the RO flow rate < 10% from the target flow rate. | 192 | ALARM_ID_DG_FLOW_RATE_OUT_OF_LOWER_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG flow rate out of lower range
Dialysate Temp |  |  | SRSHD 897 | [9 empty] | 193 | ALARM_ID_AVAILABLE_5 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD dialysate temperature out of low safety range.
x | PRS 863 | SA 360 |  | SRSDG 881 |  | DG | DG Fault: RO Driver Feedback |  | Call service to schedule a repair. |  | DG RO pump duty cycle out of range | If the DG RO pump duty cycle is out of range by 5% of the target duty cycle for a certain period of time. | 194 | ALARM_ID_DG_RO_PUMP_DUTY_CYCLE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO pump duty cycle out of range
x | PRS 863 | SA 487 |  | SRSDG 925 |  | DG | DG Fault: RO Pump High Pressure |  | Call service to schedule a repair. |  | DG RO pump pressure out of range | If the DG RO pump pressure > 135 psi for a certain period of time. | 195 | ALARM_ID_DG_RO_PUMP_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO pump pressure out of range
x | PRS 863 | SA 348 |  | SRSDG 955 |  | DG | DG Fault: CPi/CPo Comm |  | Call service to schedule a repair. |  | DG CPi/CPo conductivity sensors FPGA fault | If the FPGA reports error for a certain period of time. | 196 | ALARM_ID_DG_CPI_CPO_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CPi/CPo sensors FPGA fault
x | PRS 863 | SA 347 |  | SRSDG 956 |  | DG | DG Fault: CD1/CD2 Comm |  | Call service to schedule a repair. |  | DG CD1/CD2 conductivity sensors FPGA fault | If the FPGA reports error for a certain period of time. | 197 | ALARM_ID_DG_CD1_CD2_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CD1/CD2 sensors FPGA fault
x | PRS 863 | SA 367 |  | SRSDG 872 |  | DG | DG Fault: HP On With No Flow |  | Call service to schedule a repair. |  | DG primary heater while RO flow is low | If the RO flow is < 0.2 L/min for a certain period of time. | 198 | ALARM_ID_DG_RO_FLOW_TOO_LOW_WHILE_PRIMARY_HEATER_IS_ON |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO flow too low while primary heater is on
x | PRS 863 | SA 366 |  | SRSDG 871 |  | DG | DG Fault: HR On With No Flow |  | Call service to schedule a repair. |  | DG trimmer heater while dialysate flow is low | If the FMD flow is ≤ 0 mL/min (raw) for a certain period of time. | 199 | ALARM_ID_DG_DIALYSATE_FLOW_TOO_LOW_WHILE_TRIMMER_HEATER_IS_ON |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG Dialysate flow too low while trimmer heater is on
x | PRS 863 | SA 293 SA 358 SA 359 |  | SRSDG 178 |  | DG | DG Fault: Thermistor Range |  | Call service to schedule a repair. |  | DG thermistor or temperature sensor out of range | If thermistor > 80.0 C or < 0.0 C. | 200 | ALARM_ID_DG_THERMISTORS_TEMPERATURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG thermistors/sensors temperature out of range
|  |  | SRSHD 746 SRSHD 747 |  |  | HD | Process: Pre-Tx Wet Self-Test |  | Self-test has detected a problem in the dialysate flow \nmeasurement or there is a leak in the dialysate in/out lines.\n- The cartridge should be primed\n- Check cartridge installation on to machine\n- If problem persists, terminate pre-treatment and restart \nwith a new cartridge. |  | HD pre-treatment wet flow test failure | If the pre-treatment wet flow test has failed. | 201 | ALARM_ID_HD_PRE_TREATMENT_WET_FLOW_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode wet self-test lc vs FMD failure
[6 empty] | HD | Process: Pre-Tx Dry Self-Test |  | Self-test has detected blood circuit pressure cannot be \nrelieved.\n- Check PBo connection to HD for kinks\n -Check PBo transducer protector for occlusion\n- If problem persists, install a new cartridge and \nPBo transducer protector. |  | HD pre-treatment dry pressure normal test failure | If the pre-treatment dry pressure normal test failed. | 202 | ALARM_ID_HD_PRE_TREATMENT_DRY_PRESSURE_NORMAL_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode dry pressure normal self-test failure
[13 empty] | 203 | ALARM_ID_AVAILABLE_6 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Available for use
x | PRS 863 | SA 373 |  | SRSDG 429 |  | DG | DG Fault: RO Rejection |  | Call service to schedule a repair. |  | RO rejection ratio out of range | If RO rejection ratio > 0.1 | 204 | ALARM_ID_DG_RO_REJECTION_RATIO_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO rejection ratio out of range
|  |  |  | SRSDG 913 |  | DG | DG Fault: Cond Sensors Error |  | Call service to schedule a repair. |  | Conductivity sensor fault | If conductivity sensor reports 0x11 for more than 2 consecutive reads. | 205 | ALARM_ID_DG_CONDUCTIVITY_SENSOR_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG conductivity sensor fault
|  |  |  | SRSDG 437 |  | DG | DG Fault: Reservoir Fill Timeout |  | Call service to schedule a repair. |  | DG dialysate fill runs out of time | If DG dialysate fill > 5 minutes. | 206 | ALARM_ID_DG_DIALYSATE_FILL_OUT_OF_TIME |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG dialysate fill runs out of time
|  |  |  | SRSDG 240 |  | DG | DG Fault: FMP Flow Rate Error |  | Call service to schedule a repair. |  | DG flow meter check failure alarm | If integrated Volume To Load Cell Reading Percent > 0.1. | 207 | ALARM_ID_DG_FLOW_METER_CHECK_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG flow meter check failure alarm
blaine to revisit redundant with venous pressure alarm limit? | [5 empty] | HD | Blood Stop: Venous Pres. High |  | Venous pressure is high.\n- Check for access or patient discomfort.\n- Resolve access issue if applicable.\n- Check that there are no kinks or clots in the tubing.\n- Unkink the lines if applicable.\n- End treatment and do not return blood if clots are present. |  | HD venous line occlusion fault | If the venous pressure is > set limit + 50 mmHg if HD is in treatment or pre-treatment mode. | 208 | ALARM_ID_HD_VENOUS_LINE_OCCLUSION |  | ALARM_PRIORITY_HIGH | 201 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | HD venous pressure occlusion alarm
|  |  |  | SRSDG 980 |  | DG | Cleaning: Flush Circ Error |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. |  | DG drain circulation line timeout | If the integrated volume < 0.1 for more than 60 secs. | 209 | ALARM_ID_DG_DRAIN_CIRCULATION_LINE_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG drain circulation line timeout
x | PRS 872 | SA 229 | SRSHD 1397 |  |  | HD | HD Alarm (Low): Battery Fault |  | Call service to schedule a repair. |  | HD battery pack detected an error | If the battery pack status error is detected from the slave device I2C address 0x0B. | 210 | ALARM_ID_HD_BATTERY_PACK_ERROR_DETECTED |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | HD battery pack detected an error
need to check with sean how this will escalate with other alarm active that will affect rank | PRS 389 |  | SRSHD 946 |  |  | HD | Blood Stopped: Warning |  | Blood pump is stopped.\n- Resume treatment before countdown timer expires. |  | Blood sitting too long warning (> 4 min) | During treatment stop sub-mode, if the blood is in the line for more than 4 minutes. | 211 | ALARM_ID_HD_BLOOD_SITTING_WARNING |  | ALARM_PRIORITY_HIGH | 290 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD blood sitting too long warning (>4 min)
need to check with sean how this will escalate with other alarm active that will affect rank | PRS 389 |  | SRSHD 952 |  |  | HD | Blood Stopped: Warning |  | Blood pump is stopped.\n- Resume treatment before countdown timer expires. |  | Blood sitting too long warning (> 4 min) | During treatment end paused state, if the blood is in the line for more than 4 minutes. | 212 | ALARM_ID_HD_END_OF_TREATMENT_ALARM |  | ALARM_PRIORITY_HIGH | 290 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD end of treatment alarm (user not acting to end treatment)
sean to advise on properties | PRS 530 |  | SRSHD 1120 |  |  | HD | Process: Pre-Tx Prime Complete |  | Prime of the cartridge is complete.\n- Proceed to connect to patient. |  | HD prime completed medium priority alarm | In Pre-Treatment Recirculation for 90 minutes (escalation of alarm 226) | 213 | ALARM_ID_HD_PRIME_COMPLETED_MEDIUM |  | ALARM_PRIORITY_MEDIUM | 590 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | HD prime completed medium priority alarm
|  |  | SRSHD 909 | [9 empty] | 214 | ALARM_ID_AVAILABLE_30 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | Available for use
|  |  | SRSHD 826 | [9 empty] | 215 | ALARM_ID_AVAILABLE_31 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | Available for use
x | PRS 393 |  | SRSHD 1405 |  |  | HD | Process: End of Dialysis Tx |  | Dialysis Treatment has completed\n- Proceed to end treatment and rinseback. |  | HD end treatment sub-mode timeout alarm | If the treatment ended successfully. | 216 | ALARM_ID_HD_END_TREATMENT_TIMEOUT_ALARM |  | ALARM_PRIORITY_HIGH | 291 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD end treatment sub-mode timeout alarm
|  |  | SRSHD 1454 | [9 empty] | 217 | ALARM_ID_AVAILABLE_32 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | Available for use
|  |  | SRSHD 1294 |  |  | HD | Process: Remove Syringe |  | Syringe not removed from machine.\n- Remove heparin syringe. |  | HD syringe detected alarm | If the syringe detection switch > 2.0 volts. | 218 | ALARM_ID_HD_SYRINGE_DETECTED |  | ALARM_PRIORITY_LOW | 745 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD syringe detected alarm
Currently in CBIT section - should be somewhere else? | [5 empty] | HD | Blood Stop: Syringe Not Detected |  | Heparin is needed for treatment but syringe has been removed \nfrom syringe pump.\n- Install Heparin syringe. |  | HD syringe pump syringe removed alarm | If the syringe detection switch < 2.0 volts. | 219 | ALARM_ID_HD_SYRINGE_PUMP_SYRINGE_REMOVED |  | ALARM_PRIORITY_HIGH | 207 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | HD syringe pump syringe removed alarm
x | PRS 863 | SA 334 SA 335 |  | SRSDG 947 |  | DG | DG Fault: THD RTD Comm |  | Call service to schedule a repair. |  | DG THd sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 220 | ALARM_ID_DG_THD_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG THd sensors FPGA fault
|  |  | SRSHD 1306 SRSHD 1307 SRSHD 1308 |  |  | HD | Blood Stop: Saline Bag Empty |  | Saline bag appears empty.\n- Check saline bag and line and replace if saline bag is \nempty. |  | HD Saline bag is empty | If the measured arterial pressure < -300.0 mmHg for more than 250 msec. | 221 | ALARM_ID_HD_EMPTY_SALINE_BAG |  | ALARM_PRIORITY_HIGH | 204 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | HD Saline bag is empty
x | PRS 494 | SA 230 | SRSHD 1585 |  |  | HD | HD Fault: BP Occ Com |  | Call service to schedule a repair. |  | HD occlusion sensor FPGA fault | If the FPGA does not report fresh data or indicates an error for a certain period of time. | 222 | ALARM_ID_HD_OCCLUSION_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD occlusion sensor FPGA fault
x | PRS 494 | SA 236 | SRSHD 1583 |  |  | HD | HD Fault: Art. Pres. Com |  | Call service to schedule a repair. |  | HD arterial sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 223 | ALARM_ID_HD_ARTERIAL_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD arterial sensor FPGA fault
x | PRS 388 |  | SRSHD 944 |  |  | HD | Blood Stopped: Treatment Pause |  | Treatment paused by user. |  | Treatment stopped by user action - pressed stop button | If the user has pressed the stop button. | 224 | ALARM_ID_HD_TREATMENT_STOPPED_BY_USER |  | ALARM_PRIORITY_MEDIUM | 300 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD treatment stopped by user action - pressed stop button
revisit priority - should be low | PRS 391 |  | SRSHD 810 |  |  | HD | Process: End of Treatment |  | Dialysis Treatment has completed\n- Proceed to end treatment and rinseback. |  | HD end of treatment warning | If the treatment duration has exceeded the prescribed treatment time entered by the user. | 225 | ALARM_ID_HD_END_OF_TREATMENT_WARNING |  | ALARM_PRIORITY_LOW | 739 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD end of treatment warning
sean to advise on properties | PRS 531 |  | SRSHD 1118 |  |  | HD | Process: Pre-Tx Priming Complete |  | Prime of the cartridge is complete.\n- Proceed to connect to patient. |  | HD prime completed low priority alarm | Start of Pre-Treatment Recirculation - notifiy user to proceed to patient connection | 226 | ALARM_ID_HD_PRIME_COMPLETED_LOW_PRIORITY |  | ALARM_PRIORITY_LOW | 729 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | HD prime completed low priority alarm
[13 empty] | 227 | ALARM_ID_AVAILABLE_15 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | Available for use
|  |  | SRSHD 1571 |  |  | HD | Process: Pre-Tx Prime Timeout |  | Blood line prime with saline timeout.\n- Check blood and dialysate lines for clamps, kinks or leaks. |  | HD time out on prime sub-mode purge air state. | During the pre-treatment prime state, if the prime saline purge air state exceeded 60 secs. | 228 | ALARM_ID_HD_PRIME_SALINE_PURGE_AIR_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD time out on prime saline purge air state.
|  |  | SRSHD 1189 |  |  | HD | Process: Pre-Tx Prime Timeout |  | Dialyzer prime timeout.\n- Check blood and dialysate lines for clamps, kinks or leaks. |  | HD prime dialysate dialyzer time out alarm | If the priming of the dialyzer dialysate flow path has exceeded the time limit. | 229 | ALARM_ID_HD_PRIME_DIALYSATE_DIALYZER_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD prime dialysate dialyzer time out alarm
|  |  | SRSHD 1190 |  |  | HD | Process: Pre-Tx Prime Timeout |  | Dialyzer prime timeout.\n- Check blood and dialysate lines for clamps, kinks or leaks. |  | HD prime dialysate bypass time out alarm | If the priming of the dialyzer bypass flow path has exceeded the time limit. | 230 | ALARM_ID_HD_PRIME_DIALYSATE_BYPASS_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD prime dialysate bypass time out alarm
|  |  | SRSHD 771 SRSHD 1599 SRSHD 1579 SRSHD 1580 SRSHD 772 |  |  | HD | Process: Pre-Tx Dry Self-Test |  | Self-test has detected a leak in the circuit\n- Check blood line connections\n- Check dialysate line connections\n- Check PBo and PBa transducer protectors for occlusion.\n- Check clamps.\n- If problem persists, install a new cartridge. |  | HD pre-treatment mode dry self-tests failure | If the following dry self tests have failed: 1. handleDrySelfTestPressureSensorsState 2. handleDrySelfTestPressureSensorsNormalState | 231 | ALARM_ID_HD_PRE_TREATMENT_DRY_PRESSURE_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode dry pressure self-test failure
|  |  | SRSHD 1534 SRSHD 1535 |  |  | HD | Process: Pre-Tx Wet Self-Test |  | Self-test has detected a problem in the load cell \nmeasurement or there is a leak in the dialysate in/out lines\n- The cartridge should be primed\n- Check cartridge installation onto machine\n- If problem persists, terminate pre-treatment and restart \nwith a new cartridge. |  | HD pre-treatment mode wet self-tests failure | If the following wet self tests have failed: 1. handleWetSelfTestPrimeCheckState 2. handleWetSelfTestFirstDisplacementVerifyState 3. handleWetSelfTestSecondDisplacementVerifyState | 232 | ALARM_ID_HD_PRE_TREATMENT_WET_LC_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode wet self-test LC vs. LC failure
x | PRS 863 | SA 461 |  | SRSDG 964 SRSDG 949 SRSDG 1089 |  | DG | DG Fault: Baro Cal CRC |  | Call service to schedule a repair. |  | DG barometric pressure sensor bad coefficients CRC | If the following happens: 1. The stored CRC in the sensor does not match the calculated CRC. 2. The CRC and coefficients have not been received from the sensor after 20 seconds from the power up. | 233 | ALARM_ID_DG_BAROMETRIC_SENSOR_COEFFS_BAD_CRC |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG barometric sensor coefficients bad CRC
x | PRS 841 |  |  | SRSDG 1037 |  | DG | Dial. Gen: Inlet Water Pres High |  | Inlet water pressure is high.\n- Treatment is paused\n- Check the incoming water supply. |  | DG inlet water pressure is too high | If the inlet water pressure is > 80 psig. | 234 | ALARM_ID_DG_INLET_WATER_PRESSURE_IN_HIGH_RANGE |  | ALARM_PRIORITY_MEDIUM | 430 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water pressure in high range.
[6 empty] | HD | Process: Rinseback Complete |  | Rinseback has been completed. |  | HD in treatment stopped sub-mode after rinseback completed (no escalation) | If there is no active alarm and rinseback has completed. | 235 | ALARM_ID_HD_TREATMENT_STOPPED_AFTER_RINSEBACK |  | ALARM_PRIORITY_LOW | 738 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | HD in treatment stopped sub-mode after rinseback completed (no escalation)
|  |  | SRSHD 1205 |  |  | HD | Process: Pre-Tx Dry Self-Test |  | Self-test has detected fluid in the cartridge.\n- Cartridges are single-use only\n- Install new cartridge to resume. |  | HD needs new cartridge to be installed | If the air trap sensor reported AIR and bubble sensor detected bubbles. | 236 | ALARM_ID_HD_INSTALL_NEW_CARTRIDGE |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD needs new cartridge to be installed
|  |  | SRSHD 1188 |  |  | HD | Process: Pre-Tx Prime Timeout |  | Dialyzer prime timeout.\n- Check blood and dialysate lines for clamps, kinks or leaks. |  | HD prime saline dialyzer time out alarm | If timeout, just transition to HD_PRIME_RESERVOIR_TWO_FILL_COMPLETE_STATE | 237 | ALARM_ID_HD_PRIME_SALINE_DIALYZER_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD prime saline dialyzer time out alarm
|  |  | SRSHD 1359 |  |  | HD | Process: Pre-Tx Dry Self-Test |  | Self-test has detected cartridge is not installed.\n- Check that a cartridge is properly installed. |  | HD no cartridge loaded or installed improperly alarm | If the measured blood pump occlucsion pressure reading is < 5000 or > 24500. | 238 | ALARM_ID_HD_NO_CARTRIDGE_LOADED |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD no cartridge loaded or installed improperly alarm
|  |  | SRSHD 1542 |  |  | HD | Process: Post-Tx Remove Cartridge |  | Cartridge not unloaded, or unloaded improperly.\n- Check that the cartridge has been removed. |  | HD fail to remove cartridge alarm | In the post treatment state, if the cartridge has not been removed. | 239 | ALARM_ID_HD_CARTRIDGE_REMOVAL_FAILURE |  | ALARM_PRIORITY_LOW | 746 | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD fail to remove cartridge alarm
x | PRS 542 |  |  | SRSDG 482 |  | DG | Dial. Gen: Bicarb Conc. Cond. |  | Bicarbonate conductivity during prime is out of range.\n- Check correct concentrate is connected and full.\n- Resume to reprime concentrate. |  | Bicarb conductivity out of range during bicarb pump check alarm | If the conductivity value for bicarbonate concentrate is out of range for more than 45 secs. | 240 | ALARM_ID_DG_BICARB_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 411 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | DG bicarb conductivity out of range during bicarb pump check alarm
x | PRS 873 |  |  | SRSDG 848 SRSDG 90 SRSDG 561 SRSDG 577 SRSDG 590 SRSDG 591 SRSDG 336 SRSDG 879 SRSDG 340 SRSDG 496 SRSDG 501 SRSDG 505 SRSDG 532 SRSDG 535 SRSDG 556 SRSDG 557 SRSDG 1048 SRSDG 1049 SRSDG 651 SRSDG 655 SRSDG 744 SRSDG 745 SRSDG 1004 SRSDG 1008 SRSDG 1019 SRSDG 1023 SRSDG 1033 SRSDG 1034 |  | DG | Cleaning: Resv Drain Timeout |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. |  | DG reservoir drain time out | If the DG reservoir drain is not complete in 120 secs. | 241 | ALARM_ID_DG_RESERVOIR_DRAIN_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG reservoir drain time out
cleaning | PRS 873 |  |  | SRSDG 850 SRSDG 573 SRSDG 574 SRSDG 579 SRSDG 580 SRSDG 492 SRSDG 493 SRSDG 498 SRSDG 499 SRSDG 508 SRSDG 509 SRSDG 519 SRSDG 690 SRSDG 691 SRSDG 703 SRSDG 711 SRSDG 1020 SRSDG 1024 |  | DG | Cleaning: Resv Fill Timeout |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. |  | DG reservoir fill time out | If the DG reservoir fill is not complete in 120 secs. | 242 | ALARM_ID_DG_RESERVOIR_FILL_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG reservoir fill time out
cleaning |  |  |  | SRSDG 853 |  | DG | Cleaning: Fluid Loss Detected |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. |  | DG reservoir leak time out | If the reservoir level is drifted by 100 mL for more than 30 seconds. | 243 | ALARM_ID_DG_RESERVOIR_LEAK_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG reservoir leak time out
cleaning |  |  |  | SRSDG 852 |  | DG | Cleaning: Temp Sensor Check |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. |  | DG temperature sensors difference out of range | If any of the temperature sensors are out of the average of the sensors by more than 3 C. | 244 | ALARM_ID_DG_CLEANING_MODE_TEMP_SENSORS_DIFF_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode temperature sensors difference out of range
cleaning |  |  |  | SRSDG 854 |  | DG | Cleaning: Heat Disinfect Temp TO |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. |  | DG heat disinfect target temperature time out (could not reach to temperature) | If heating up to minimum temperature for heat disinfect failed after 4 hours. | 245 | ALARM_ID_DG_HEAT_DISINFECT_TARGET_TEMP_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG heat disinfect target temperature time out (could not reach to temperature)
cleaning | [5 empty] | DG | Cleaning: RO Water Cond Error |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. |  | DG chemical disinfect inlet conducitivty out of range | If CPo and CD2 conductivity values are above 100 uS/cm. | 246 | ALARM_ID_DG_CLEANING_MODE_COND_SENSORS_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode conductivity sensors out of range
cleaning | [12 empty] | 247 | ALARM_ID_AVAILABLE_7 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Alarm Id available
cleaning | [5 empty] | DG | Cleaning: Chem Disinfect Proc |  | Cleaning mode process terminated unexpectedly.\n-Restart cleaning mode.\n-If problem persists, call service to schedule a repair. |  | DG chemical disinfect out of target temperature | If TPo > 50 C after 1 second. | 248 | ALARM_ID_DG_CHEM_DISINFECT_TARGET_TEMP_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG chemical disinfect could not reach to target temperature out of range
cleaning |  |  |  | SRSDG 992 |  | DG | Cleaning: Chem Disinf PAA Cond |  | Unable to sense proper chemical disinfectant.\n- Cleaning mode process has terminated.\n- Check that proper chemical disinfectant is connected to the \nbicarb port.\n- Check chemical disinfectant bottle is not empty. |  | DG chemical disinfect out of target conductivity | If CD2 < 150 uS/cm or CD2 > 650 uS/cm after 10 seconds. | 249 | ALARM_ID_DG_CHEM_DISINFECT_TARGET_COND_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG chemical disinfect target conductivity out of range
cleaning |  |  |  | SRSDG 1046 |  | DG | Cleaning: Chem Disinf Connect PAA |  | Connect chemical disinfectant bottle to bicarb port. |  | Insert the acid to start chemical disinfect | Start of chemical disinfect | 250 | ALARM_ID_DG_CHEM_DISINFECT_INSERT_ACID |  | ALARM_PRIORITY_LOW | 750 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG chemical disinfect insert acid and remove the concentrate cap
x | PRS 398 | SA 196 | SRSHD 1058 |  |  | HD | HD POST: Sys Rec NVRAM CRC |  | Call service to schedule a repair. |  | HD invalid system record | If the calculated CRC of the HD system record image does not match the expected CRC. | 251 | ALARM_ID_HD_INVALID_SYSTEM_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD invalid system record
x | PRS 398 | SA 196 | SRSHD 1062 |  |  | HD | HD POST: Serv Rec NVRAM CRC |  | Call service to schedule a repair. |  | HD invalid service record | If the calculated CRC of the HD service record image does not match the expected CRC. | 252 | ALARM_ID_HD_INVALID_SERVICE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD invalid service record
x | PRS 863 | SA 283 |  | SRSDG 609 |  | DG | DG POST: Sys Rec NVRAM CRC |  | Call service to schedule a repair. |  | DG invalid system record | If the calculated CRC of the DG system record image does not match the expected CRC. | 253 | ALARM_ID_DG_INVALID_SYSTEM_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid system record
x | PRS 863 | SA 283 |  | SRSDG 610 |  | DG | DG POST: Serv Rec NVRAM CRC |  | Call service to schedule a repair. |  | DG invalid service record | If the calculated CRC of the DG service record image does not match the expected CRC. | 254 | ALARM_ID_DG_INVALID_SERVICE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid service record
x | PRS 398 | SA 206 | SRSHD 1358 |  |  | HD | HD POST: UI Compatibility |  | Call service to schedule a repair. |  | HD and UI software builds are not compatible | If the compatibility revision field of the UI is not equal to SW_COMPATIBILITY_REV. | 255 | ALARM_ID_HD_UI_COMPATIBILITY_ERROR |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD and UI software builds are not compatible
x | PRS 398 | SA 205 |  |  | SRSUI 497 | HD | Cleaning: Chem Flush In Progress |  | Chemical disinfect flush in progress\n- Please wait for the process to complete\n- Treatment will not be able to be started until process \ncompletes.\n- Process may be terminated by pressing the HD stop button. |  | DG flush started | If DG flush successfully started from the UI. | 256 | ALARM_ID_HD_DISINFECT_CHEM_FLUSH | OK | ALARM_PRIORITY_LOW | 843 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD processor is in Disinfect Chemical Flush mode
[13 empty] | 257 | ALARM_ID_AVAILABLE_8 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Alarm ID available for use
[13 empty] | 258 | ALARM_ID_AVAILABLE_9 |  | ALARM_PRIORITY_LOW | 999 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | Alarm Id available
x | PRS 494 | SA 220 SA 215 SA 216 SA 235 SA 238 | SRSHD 1390 SRSHD 1388 SRSHD 1387 SRSHD 1497 SRSHD 1499 SRSHD 1682 |  |  | HD | HD Fault: Temp Sensors Error |  | Call service to schedule a repair. |  | HD temperatures out of range | If thermistor reports temperature > 80.0 C or < 0.0 C for a certain period of time. | 259 | ALARM_ID_HD_TEMPERATURES_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD temperatures out of range
x | PRS 398 | SA 200 | SRSHD 977 |  | SRSUI 1067 | UI HD | HD POST: UI Image |  | Call service to schedule a repair. |  | UI POST Application Integrity (Sha256Sum) failure | If the UI POST reports a shasum faillure. | 260 | ALARM_ID_HD_UI_POST_FAILURE_SHASUM | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST Application Integrity (Sha256Sum) failure
x | PRS 398 | SA 201 | SRSHD 1655 |  | SRSUI 1068 | UI HD | HD POST: UI CAN Driver |  | Call service to schedule a repair. |  | UI POST CANBus failure | If the UI POST reports a can bus failure. | 261 | ALARM_ID_HD_UI_POST_FAILURE_CANBUS | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST CANBus failure
x | PRS 398 | SA 201 | SRSHD 1656 |  | SRSUI 1069 | UI HD | HD POST: UI Disp Driver |  | Call service to schedule a repair. |  | UI POST Display failure | If the UI POST reports a display failure. | 262 | ALARM_ID_HD_UI_POST_FAILURE_DISPLAY | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST Display failure
x | PRS 398 | SA 201 | SRSHD 1657 |  | SRSUI 1070 | UI HD | HD POST: UI Touch Driver |  | Call service to schedule a repair. |  | UI POST Touch failure | If the UI POST reports a touch failure. | 263 | ALARM_ID_HD_UI_POST_FAILURE_TOUCH | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST Touch failure
x | PRS 398 | SA 201 | SRSHD 1658 |  | SRSUI 1071 | UI HD | HD POST: UI SD-Card |  | SD-Card Failure \n- Power cycle the system.\n- Call service to schedule a repair if problem persists. |  | UI POST SD-Card failure | If the UI POST reports SD card failure. | 264 | ALARM_ID_HD_UI_POST_FAILURE_SDCARD | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST SD-Card failure
x | PRS 398 | SA 201 SA 202 SA 203 | SRSHD 1659 |  | SRSUI 1072 | UI HD | HD POST: UI RTC |  | Call service to schedule a repair. |  | UI POST RTC failure | If the UI POST reports RTC failure. | 265 | ALARM_ID_HD_UI_POST_FAILURE_RTC | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST RTC failure
x | PRS 398 | SA 201 | SRSHD 1660 |  | SRSUI 1073 | UI HD | HD POST: UI WiFi |  | Call service to schedule a repair. |  | UI POST WiFi failure | If the UI POST reports WiFi failure. | 266 | ALARM_ID_HD_UI_POST_FAILURE_WIFI | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST WiFi failure
x | PRS 398 | SA 201 | SRSHD 1661 |  | SRSUI 1074 | UI HD | HD POST: UI Bluetooth |  | Call service to schedule a repair. |  | UI POST Bluetooth failure | If the UI POST reports bluetooth failure. | 267 | ALARM_ID_HD_UI_POST_FAILURE_BLUETOOTH | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST Bluetooth failure
x | PRS 398 | SA 201 | SRSHD 1662 |  | SRSUI 1075 | UI HD | HD POST: UI Ethernet |  | Call service to schedule a repair. |  | UI POST Ethernet failure | If the UI POST reports ethernet failure. | 268 | ALARM_ID_HD_UI_POST_FAILURE_ETHERNET | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST Ethernet failure
x | PRS 398 | SA 201 | SRSHD 1663 |  | SRSUI 1076 | UI HD | HD POST: UI Sound |  | Call service to schedule a repair. |  | UI POST Sound failure | If the UI POST reports sound failure. | 269 | ALARM_ID_HD_UI_POST_FAILURE_SOUND | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST Sound failure
x | PRS 398 | SA 459 | SRSHD 1481 |  |  | HD | HD POST: Proc Safestate |  | Call service to schedule a repair. |  | HD POST Safety Shutdown failure | If the 24 volts power source is greater than 5.0 volts during safety shutdown self test state. If the 24 volts power source is less than 22.6 volts during safety shutdown self test recover state. | 270 | ALARM_ID_HD_SAFETY_SHUTDOWN_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD POST Safety Shutdown failure
x | PRS 863 | SA 292 |  | SRSDG 822 |  | DG | DG POST: Proc Safestate |  | Call service to schedule a repair. |  | DG POST Safety Shutdown failure | If the 24 volts power source is greater than 5.0 volts during safety shutdown self test state. If the 24 volts power source is less than 22.6 volts during safety shutdown self test recover state. | 271 | ALARM_ID_DG_SAFETY_SHUTDOWN_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG POST Safety Shutdown failure
x | PRS 872 | SA 221 | SRSHD 1394 |  |  | HD | HD Alarm (Low): Fan Speed |  | Fan failure.\n- Check that the fans inlets are not blocked or obstructed. |  | HD Fan RPM out of range | If the fan RPM is out of Min/Max range. | 272 | ALARM_ID_HD_FAN_RPM_OUT_OF_RANGE |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | HD Fan RPM out of range
[13 empty] | 273 | ALARM_ID_AVAILABLE_10 |  | ALARM_PRIORITY_LOW | 999 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | Available for use
Blaine - Remove SRS (Done) Sean - Remove from code |  |  | Removed | [9 empty] | 274 | ALARM_ID_AVAILABLE_11 |  | ALARM_PRIORITY_LOW | 999 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | Available for use
x | PRS 494 | SA 237 | SRSHD 1706 |  |  | HD | HD Fault: Arterial Press Range |  | Call service to schedule a repair. |  | HD arterial pressure sensor is reading out of range | If the arterial pressure > 2000.0 mmHg or < -1500.0 mmHg. | 275 | ALARM_ID_HD_ARTERIAL_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD arterial pressure sensor is reading out of range
x | PRS 494 | SA 234 | SRSHD 1707 |  |  | HD | HD Fault: Venous Press Range |  | Call service to schedule a repair. |  | HD venous pressure sensor is reading out of range | If the venous pressure > 2000.0 mmHg or < -1500.0 mmHg. | 276 | ALARM_ID_HD_VENOUS_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD venous pressure sensor is reading out of range
Sean to remove from code check of 32767 check (align to SRS) | PRS 494 | SA 231 | SRSHD 1708 |  |  | HD | HD Fault: Occlusion Sens Range |  | Call service to schedule a repair. |  | HD BP occlusion sensor is reading out of range | If the blood pump occlusion count is < 2000. | 277 | ALARM_ID_HD_BP_OCCLUSION_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD BP occlusion sensor is reading out of range
Dara - update code to 1900g |  |  | SRSHD 793 |  |  | HD | DG Fault: Reservoir Overfill |  | Call service to schedule a repair. |  | HD active reservoir weight out of range. | If the active reservoir weight is > 1900.0 g. | 278 | ALARM_ID_HD_ACTIVE_RESERVOIR_WEIGHT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD active reservoir weight out of range.
|  |  |  | SRSDG 752 |  | DG | DG Fault: Reservoir Drain |  | Call service to schedule a repair. |  | DG dialysate drain time out | If the DG dialysate drain exceeded 120 seconds. | 279 | ALARM_ID_DG_DIALYSATE_DRAIN_TIME_OUT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG dialysate drain time out
| PRS 863 | SA 340 SA 344 |  | SRSDG 888 |  | DG | DG Fault: DRP Direction Fault |  | Call service to schedule a repair. |  | DG drain pump invalid direction | If the drain pump's direction is not 1 when draining. | 280 | ALARM_ID_DG_DRAIN_PUMP_DIRECTION_INVALID |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain pump direction invalid
| PRS 827 |  |  | SRSDG 836 |  | DG | Dial. Gen: Acid Conc. Low |  | Acid concentrate low.\n- Replace acid concentrate bottle. |  | DG Acid concentration bottle low volume alarm | If the acid concentration bottle volume > 3240 (3340 - 100). | 281 | ALARM_ID_DG_ACID_BOTTLE_LOW_VOLUME |  | ALARM_PRIORITY_MEDIUM | 412 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | Acid concentration bottle low volume alarm
| PRS 828 |  |  | SRSDG 837 |  | DG | Dial. Gen: Bicarb Conc. Low |  | Bicarbonate concentrate low.\n- Replace bicarbonate concentrate bottle. |  | DG Bicarbonate concentration bottle low volume alarm | If the bicarbonate concentration bottle volume > 3680 (3780 - 100). | 282 | ALARM_ID_DG_BICARB_BOTTLE_LOW_VOLUME |  | ALARM_PRIORITY_MEDIUM | 413 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | Bicarbonate concentration bottle low volume alarm
x | PRS 863 | SA 309 SA 310 SA 314 SA 315 |  | SRSDG 615 |  | DG | DG Fault: LC Range |  | Call service to schedule a repair. |  | DG load cells weight out of range | If the weight of the DG load cells > 4500.0 g or < 0.0 g. | 283 | ALARM_ID_DG_LOAD_CELL_WEIGHT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells weight out of range
x | PRS 863 | SA 317 SA 318 |  | SRSDG 383 |  | DG | DG Fault: Redundant LC Check |  | Call service to schedule a repair. |  | DG load cells primary/back up drift out of range | If redundant load cells deviate by more than 60g after tare for 5 seconds. | 284 | ALARM_ID_DG_LOAD_CELL_PRIMARY_BACKUP_DRIFT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells primary/back up drift out of range
[13 empty] | 285 | ALARM_ID_AVAILABLE_12 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | Available for use
[6 empty] | DG | Cleaning: Conc Cap Removed |  | Concentrate cap removed from DG during cleaning.\n- Cleaning Process Terminated. |  | DG cap in wrong state | If the concentrate cap is not in the right position in the cleaning modes. | 286 | ALARM_ID_DG_CONCENTRATE_CAP_NOT_IN_PROPER_POSITION |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG concentrate cap not in proper position
x | PRS 494 | SA 222 | SRSHD 1403 |  |  | HD | HD Fault: Proc-FPGA Clock |  | Call service to schedule a repair. |  | HD processor clock speed checks against FPGA clock failure | If during initialization mode, the difference between the processor clock speed and the FPGA clock speed > 1. | 287 | ALARM_ID_HD_FPGA_CLOCK_SPEED_CHECK_FAILURE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD processor clock speed checks against FPGA clock failure
|  | SA 491 SA 492 | SRSHD 562 |  |  | HD | DG Fault: LC Drift Check |  | Call service to schedule a repair. |  | HD load cell inconsistency | If redundant load cells deviate from start to end of a reservoir use during treatment by more than 10g. | 288 | ALARM_ID_HD_LOAD_CELL_PRIMARY_BACKUP_DRIFT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | HD load cells primary/back up drift out of range
Cleaning |  |  |  | SRSDG 928 |  | DG | Cleaning: Dialysate Cap Removed |  | Dialysate cap removed from DG during cleaning.\n- Cleaning Process Terminated. |  | DG cap in wrong state | If the dialysate cap is not in the right position in the cleaning modes. | 289 | ALARM_ID_DG_DIALYSATE_CAP_NOT_IN_PROPER_POSITION |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG dialysate cap not closed
x | PRS 494 | SA 233 | SRSHD 1584 |  |  | HD | HD Fault: Ven. Pres. Com |  | Call service to schedule a repair. |  | HD venous sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 290 | ALARM_ID_HD_VENOUS_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD venous pressure sensor FPGA fault
Cleaning | [5 empty] | HD | Cleaning: Flush In Progress |  | System flush in progress\n- Please wait for the process to complete.\n- Treatment will not be able to be started until process \ncompletes.\n- Process may be terminated by pressing the HD stop button. |  | DG flush started | If DG flush successfully started from the UI. | 291 | ALARM_ID_HD_DISINFECT_FLUSH |  | ALARM_PRIORITY_LOW | 840 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD processor is in Disinfect Flush mode
Cleaning | [5 empty] | HD | Cleaning: Heat Disinf In Progress |  | Heat disinfect in progress\n- Please wait for the process to complete.\n- Treatment will not be able to be started until process \ncompletes.\n- Process may be terminated by pressing the HD stop button. |  | DG heat disinfect started | If DG heat disinfect successfully started from the UI. | 292 | ALARM_ID_HD_DISINFECT_HEAT |  | ALARM_PRIORITY_LOW | 841 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD processor is in Disinfect Heat mode
Cleaning | [5 empty] | HD | Cleaning: Chem Disinf In Progress |  | Chemical disinfect in progress\n- Please wait for the process to complete.\n- Treatment will not be able to be started until process \ncompletes.\n- Process may be terminated by pressing the HD stop button. |  | DG chemical disinfect started | If DG chemical disinfect successfully started from the UI. | 293 | ALARM_ID_HD_DISINFECT_CHEM |  | ALARM_PRIORITY_LOW | 842 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD processor is in Disinfect Chemical mode
remove | [12 empty] | 294 | ALARM_ID_AVAILABLE_13 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Available for use
x | PRS 398 | SA 470 | SRSHD 1664 |  | SRSUI 1238 | UI HD | HD POST: UI CloudSync |  | Call service to schedule a repair. |  | UI cloud sync POST failure | If the UI POST reports cloud sync failure. | 295 | ALARM_ID_HD_UI_POST_FAILURE_CLOUDSYNC | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST CloudSync failure
x | PRS 863 | SA 330 SA 331 |  | SRSDG 951 |  | DG | DG Fault: TDi RTD Comm |  | Call service to schedule a repair. |  | DG TDi sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 296 | ALARM_ID_DG_TDI_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG TDi sensors FPGA fault
x | PRS 863 | SA 326 SA 327 |  | SRSDG 950 |  | DG | DG Fault: TRo RTD Comm |  | Call service to schedule a repair. |  | DG TRo sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 297 | ALARM_ID_DG_TRO_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG TRo sensors FPGA fault
x | PRS 863 | SA 349 |  | SRSDG 948 |  | DG | DG Fault: Baro Pres Comm |  | Call service to schedule a repair. |  | DG barometric sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 298 | ALARM_ID_DG_BARO_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG baro sensor FPGA fault
x | PRS 863 | SA 483 |  | SRSDG 952 |  | DG | DG POST: Serial Number Check |  | Call service to schedule a repair. |  | DG invalid serial number | If the DG serial number is empty and does not have any characters. | 299 | ALARM_ID_DG_INVALID_SERIAL_NUMBER |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid serial number
x | PRS 398 | SA 473 | SRSHD 1582 |  |  | HD | HD POST: Serial Number Check |  | Call service to schedule a repair. |  | HD invalid serial number | If the HD serial number is empty and does not have any characters. | 300 | ALARM_ID_HD_INVALID_SERIAL_NUMBER |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD invalid serial number
remove redundant with SRSDG 866 | [12 empty] | 301 | ALARM_ID_AVAILABLE_14 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Alarm ID available
x | PRS 863 | SA 485 |  | SRSDG 126 |  | DG | DG Fault: Processor RAM error |  | Call service to schedule a repair. |  | DG_CPU RAM error | If the RAM status indicates any of the following: Double bit error, Address decode failure, Address decode logic element failure, Read address parity failure, Write address parity failure. | 302 | ALARM_ID_DG_CPU_RAM_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG processor RAM error
x | PRS 863 | SA 345 |  | SRSDG 958 |  | DG | DG Fault: DRP Hall Commutation |  | Call service to schedule a repair. |  | DG drain pump direction FPGA fault | If the FPGA indicates an error for a certain period of time. | 303 | ALARM_ID_DG_DRAIN_PUMP_DIRECTION_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain pump direction FPGA fault
x | PRS 398 | SA 197 | SRSHD 1587 |  |  | HD | HD POST: Usage Rec NVRAM CRC |  | Call service to schedule a repair. |  | HD invalid usage reocord | If the CRC of the HD usage record does not match the calculated CRC. | 304 | ALARM_ID_HD_INVALID_USAGE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD invalid usage record
x | PRS 494 | SA 484 | SRSHD 1577 |  |  | HD | HD Fault: Processor RAM error |  | Call service to schedule a repair. |  | HD_CPU RAM error | If the RAM status indicates any of the following: Double bit error, Address decode failure, Address decode logic element failure, Read address parity failure, Write address parity failure. | 305 | ALARM_ID_HD_CPU_RAM_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD processor RAM error
[6 empty] | DG | DG Fault: CD1 CD2 Check |  | Call service to schedule a repair. |  |  | CD1 and CD2 conductivity measurement during acid check is not within 5% of each other | 306 | ALARM_ID_DG_CD1_CD2_CHECK_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CD1 CD2 check fault
x | PRS 396 |  | SRSHD 1594 |  |  | HD | Blood Stop: Loss of AC Power |  | HD power has been lost. \n- Treatment must be terminated.\n-Check all electrical connections.\n- The display and CPU will run off of the battery until \npower returns.\n-If power is returned prior to blood stopped countdown elapses, \nrinseback will be allowed. |  | HD AC power lost alarm | The HD Software shall trigger an alarm if an AC power loss is detected while the HD system is in pre-treatment, treatment, or post-treatment mode. | 307 | ALARM_ID_HD_AC_POWER_LOST_IN_TREATMENT |  | ALARM_PRIORITY_HIGH | 100 | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | HD AC power lost second alarm
x | PRS 863 | SA 348 |  | SRSDG 959 |  | DG | DG Fault: CPi Comm |  | Call service to schedule a repair. |  | DG CPi conductivity sensor transmits invalid characters | If the CPi conductivity sensor transmits invalid characters for a certain period of time. | 308 | ALARM_ID_DG_CPI_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CPi conductivity sensor invalid character received
x | PRS 863 | SA 348 |  | SRSDG 960 |  | DG | DG Fault: CPo Comm |  | Call service to schedule a repair. |  | DG CPo conductivity sensor transmits invalid characters | If the CPo conductivity sensor transmits invalid characters for a certain period of time. | 309 | ALARM_ID_DG_CPO_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CPo conductivity sensor invalid character received
x | PRS 863 | SA 347 |  | SRSDG 962 |  | DG | DG Fault: CD1 Comm |  | Call service to schedule a repair. |  | DG CD1 conductivity sensor transmits invalid characters | If the CD1 conductivity sensor transmits invalid characters for a certain period of time. | 310 | ALARM_ID_DG_CD1_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CD1 conductivity sensor invalid character received
x | PRS 863 | SA 347 |  | SRSDG 961 |  | DG | DG Fault: CD2 Comm |  | Call service to schedule a repair. |  | DG CD2 conductivity sensor transmits invalid characters | If the CD2 conductivity sensor transmits invalid characters for a certain period of time. | 311 | ALARM_ID_DG_CD2_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CD2 conductivity sensor invalid character received
Sean - review code and requirements use raw FMD flow (add DG to publish if not doing so already) long persistance |  |  | SRSHD 1600 |  |  | HD | Dial. Gen: Dialysate In Flow |  | Dialysate in flow rate is low\n- Check dialysate line connection to the DG. |  | The dialysate inlet flow is too low. | If the dialysate inlet flow < 10 mL/min (raw) while in treatment mode or pre-treatment mode after prime and DPi is running. | 312 | ALARM_ID_HD_DIAL_IN_FLOW_CHECK_FAILURE |  | ALARM_PRIORITY_MEDIUM | 400 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | HD Dialysate In flow rate to too low.
x | PRS 852 |  |  | SRSDG 1064 SRSDG 1052 SRSDG 968 |  | DG | Cleaning: Inlet Water Temp Low |  | Cleaning Mode Inlet water temperature too low.\n- Cleaning mode process has terminated.\n- Check inlet water before restarting cleaning mode. |  | DG cleaning mode inlet water temperature too low | If inlet water temperature is < 24 C. | 313 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_TEMP_TOO_LOW |  | ALARM_PRIORITY_MEDIUM | 602 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water temperature too low.
x | PRS 855 |  |  | SRSDG 1063 SRSDG 1051 SRSDG 967 |  | DG | Cleaning: Inlet Water Press High |  | Cleaning Mode Inlet water pressure too high.\n- Cleaning mode process has terminated.\n- Check inlet water before restarting cleaning mode. |  | DG cleaning mode inlet water pressure too high | If inlet water temperature is > 80 psig. | 314 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_PRESSURE_TOO_HIGH |  | ALARM_PRIORITY_MEDIUM | 605 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water pressure too high.
x | PRS 856 |  |  | SRSDG 1068 SRSDG 1056 SRSDG 966 |  | DG | Cleaning: Inlet Water Press Low |  | Cleaning Mode Inlet water pressure too low.\n- Cleaning mode process has terminated.\n- Check inlet water before restarting cleaning mode. |  | DG cleaning mode inlet water pressure too low | If inlet water pressure is < 14 psig. | 315 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_PRESSURE_TOO_LOW |  | ALARM_PRIORITY_MEDIUM | 606 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water pressure too low.
x | PRS 863 | SA 495 |  | SRSDG 983 |  | DG | DG Fault: Acid CP Pump Park |  | Call service to schedule a repair. |  | Acid concentrate pump did not park successfully | If the park took more than 1 second. | 316 | ALARM_ID_DG_ACID_CONCENTRATE_PUMP_PARK_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG acid concentrate pump park fault.
x | PRS 863 | SA 496 |  | SRSDG 984 |  | DG | DG Fault: Bicarb CP Pump Park |  | Call service to schedule a repair. |  | Bicarb concentrate pump did not park successfully | If the park took more than 1 second. | 317 | ALARM_ID_DG_BICARB_CONCENTRATE_PUMP_PARK_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG bicarb concentrate pump part fault.
Blaine to check with michael |  |  | SRSHD 1609 |  |  | HD | HD Fault: Air Pump Timeout |  | Unable to reduce fluid level.\n- Air trap may be overfilled.\n- Pressure transducer protector may be contaminated.\n- Treatment must be terminated without option for rinseback.\n- Call service to schedule a repair. |  | HD Air Pump operation timeout | If air pump has been turned on for greater than 10s. | 318 | ALARM_ID_HD_AIR_PUMP_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | HD Air Pump timeout alarm
Cleaning |  |  |  | SRSDG 1040 |  | DG | Cleaning: Chem Residual Sampling |  | Collect chemical residuals water sample from drain line. |  | Collect sample to check the chemical disinfectant acid | Chemical disinfect flush sample alarm is prompted. | 319 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_FLUSH_SAMPLE |  | ALARM_PRIORITY_LOW | 752 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG chemical disinfect flush flush sample.
Cleaning |  |  |  | SRSDG 1043 |  | DG | Cleaning: Chem Residual Timeout |  | During Chemical Disinfection Flush\n-user residual sampling timed out. |  | The user did not proceed with collecting chemical disinfect flush sample for more than 30 minutes | Chemical disinfect flush sample was not collected form more than 30 minutes. | 320 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_SAMPLE_TIMEOUT |  | ALARM_PRIORITY_LOW | 752 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG chemical disinfect flush sample timeout.
x | PRS 863 | SA 356 |  | SRSDG 1069 |  | DG | DG Fault: UVo Error |  | Call service to schedule a repair. |  | Outlet reactor not healthy | If the outlet reactor are not healthy. | 321 | ALARM_ID_DG_OUTLET_UV_REACTOR_NOT_HEALTHY |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG outlet UV reactor not healthy
[6 empty] | UI HD | HD POST Alarm: UI Invalid Year |  | Call service to schedule a repair. |  | UI POST Invalid Year | If the UI POST reports a invalid year. | 322 | ALARM_ID_HD_UI_POST_FAILURE_INVALID_YEAR |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST Year check failure

## Property Definitions
Property Name | Definition |  | Column | Description
AlarmID | Unique alarm identification number |  | PRS ID | PRS requirement ID associated with the alarm.
Alarm Name | Name of the alarm |  | HD SRS ID | HD SRS requirement ID (if applicable) associated with the alarm.
Priority | Priority of the alarm (High, Medium, Low) |  | DG SRS ID | DG SRS requirement ID (if applicable) associated with the alarm.
Rank | Within the same priority, higher rank alarms have higher priority |  | Device | Sub-system that detects the alarm.
Fault | The alarm activation causes an HD fault |  | Description | Detailed description of the alarm.
DGFault | The alarm activation causes an DG fault |  | Displayed Title | Alarm title displayed prominantly at top of alarm dialog when the alarm is top priority.
Stop | The alarm activation causes a controlled stop |  | Displayed Instructions | Instructions displayed in alarm dialog when the alarm is top priority.
Clear Immediate | Alarm condition is cleared immediately after alarm activation |  | Trigger Conditions | Condition(s) required to trigger the alarm.
Non Recoverable | Alarm cannot be cleared (unrecoverable) |  | AlarmID | Unique alarm ID #.
No Resume | Alarm activation prevents treatment resume |  | Alarm Name | Enumeration name (code reference) for the alarm.
No Rinseback | Alarm activation prevents rinseback |  | Priority | Priority of the alarm (high, medium or low per 60601-1-8.
No End Treatment | Alarm activation prevents ending treatment |  | Rank | Rank of the alarm (used to determine which alarm is top priority) (low rank=high priority).
No New Treatment | Alarm activation prevents new treatment |  | Fault | Does HD f/w consider the alarm to be a system fault.
No Dialysate Recirc | Alarm activation prevents dialysate re-circulation |  | DGFault | Does DG f/w consider the alarm to be a system fault (will not prevent HD from some activities).
Clear Top Only | User acknowledgment clears only displayed alarm |  | Stop | Does the HD f/w need to pause treatment (or perhaps other activities) when alarm is active.
Escalate In | Alarm will escalate if not cleared within this time (zero indicates no escalation) |  | Clear Immediate | Should the f/w immediately clear the alarm condition after trigger (not clear the alarm itself).
Escalate To | ID of the alarm that this alarm will be escalated to |  | Non Recoverable | Is the alarm recoverable (can user clear the alarm).
|  |  | No Resume | Should treatment resume be prevented when the alarm is active.
|  |  | No Rinseback | Should rinseback be prevented when the alarm is active.
|  |  | No End Treatment | Should ending treatment be prevented when the alarm is active.
Category | Priority |  | No Blood Recirc | Should blood recirculation be prevented when the alarm is active (treatment is paused). Dialyzer is bypassed regardless.
HD POST | High |  | No Dialysate Recirc | Should dialysate recirculation be prevented when the alarm is active (treatment is paused). Dialyzer is bypassed regardless.
HD Machine Alarm | High |  | Clear Top Only | If user acknowledges alarm, should all active alarms be cleared or just the alarm top (displayed) alarm.
DG POST DG Machine Alarm | High |  | TreatmentLog | Should alarm be included in treatment log.
Blood Stopped - No Rinseback Allowed | HIgh |  | Escalate In | Duration (in seconds) until alarm escalates to a higher priority alarm (0 indicates no escalation).
Blood Stopped - Rinseback Allowed | High |  | Escalate To | Higher priority alarm to escalate to when escalation period expires (None indicates no escalation).
End Pre-Tx | Medium |  | Doxygen Title | Used for design documentation purposes.
Concentrate Warning | Low
Concentrate Alarm | Medium
Saline Warning | Low
Saline Alarm | Medium
Dialysate Alarms (Temp, conductivity, recirc?) | Medium
Other Procedural (setup, pre-tx, etc) | Medium/Low
Disinfection Alarms | Medium
HD Machine Alert | Low
DG Machine Alert | Low
Blood Pressure Alarms | N/A - not supported
Priority | Type | Mode