## Alarms
Action | PRS ID2 | SA | HD SRS ID | DG SRS ID | UI SRS ID | Device | Displayed Title | Char count | Displayed Instructions | Alarm List Title | Char Count2 | Description | Trigger Conditions | AlarmID | Alarm Name | Alarm ID Check | Priority | Rank | Fault | DGFault | Stop | Clear Immediate | Non Recoverable | No Resume | No Rinseback | No End Treatment | Block in Rinseback | Block in Post Tx | Block Until SB Rinseback | Block Until SB End Tx | No Blood Recirc | No Dialysate Recirc | Clear Top Only | TreatmentLog | Doxygen Title
N/A | N/A |  | N/A | N/A |  | N/A | [5 empty] | Indication for when no alarms is a possible situation | N/A | 0 | ALARM_ID_NO_ALARM |  | ALARM_PRIORITY_NONE | 999 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | Indication for when no alarms is a possible situation
x | PRS 398 | SA 213 | SRSHD 652 |  |  | HD | Stuck Button |  | A button was pressed during start up and appears stuck. \n- Press and release the power and stop buttons on the\nfront of the device.\n- Turn device off, then back on.\n- If problem continues, call service to schedule a repair. | HD POST: Stuck Button |  | Stuck button POST failure. Stop or Off button detected to be pressed for at least 1 second during test shortly after power up | Stop or Off button are pressed for more than 1 sec during POST. | 1 | ALARM_ID_HD_STUCK_BUTTON_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD stuck button POST failure. Stop or Off button detected to be pressed for at least 1 second during test shortly after power up
x | PRS 398 | SA 207 | SRSHD 995 SRSHD 1357 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: FPGA Compatibility |  | HD FPGA POST failure | Minimum HD FPGA major revision is not compatible with this version of HD firmware or the FPGA Id is not correct. | 2 | ALARM_ID_HD_FPGA_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD FPGA POST failure
x | PRS 863 | SA 287 |  | SRSDG 602 SRSDG 814 SRSDG 1061 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: FPGA Compatibility |  | DG FPGA POST failure | Minimum DG FPGA major and minor revision are not equal to the FPGA major and minor revision in FPGA header | 3 | ALARM_ID_DG_FPGA_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG FPGA POST failure
Sean - check if DG and HD POST test is the same in code. Consider updating trigger conditions | PRS 398 | SA 208 SA 209 SA 210 SA 211 | SRSHD 650 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: WD Safestate |  | HD Watchdog POST failure | If the following are not observed after 500ms of not petting the WD: 1. WD exp pin=low 2. 24V <= 5V 3. audio current >= 200mA If the following are not observed after 500ms of resumed WD petting at proper timing: 1. WD exp pin=high 2. 24V >= 22.6V 3. audio current <= 10 mA | 4 | ALARM_ID_HD_WATCHDOG_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD Watchdog POST failure
Sean - check if DG and HD POST test is the same in code. Consider updating trigger conditions | PRS 863 | SA 288 SA 290 SA 289 |  | SRSDG 53 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: WD Safestate |  | DG Watchdog POST failure | If the following are not observed after 500ms of not petting the WD: 1. WD exp pin=low 2. 24V <= 5V If the following are not observed after 750ms of resumed WD petting at proper timing: 1. WD exp pin=high 2. 24V >= 22.6V | 5 | ALARM_ID_DG_WATCHDOG_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG Watchdog POST failure
x | PRS 398 | SA 204 | SRSHD 1484 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: Proc- UI Comm |  | UI communication POST failure. UI failed to communicate within a reasonable time after power up | If the UI failed to checkin via the CAN bus within 40 seconds | 6 | ALARM_ID_HD_UI_COMM_POST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI communication POST failure. UI failed to communicate within a reasonable time after power up
x | PRS 872 | SA 457 | SRSHD 1595 |  |  | HD | Service Required: Hemodialysis Device |  | A non- critical problem was detected with the hemodialysis\n device.\n- Call service to report the issue and schedule a repair. | HD Alarm (Low): RTC Bat Low |  | Real-time clock configuration error | If the RTC is not functional as indicated by the bits in the first 3 control registers. | 7 | ALARM_ID_HD_RTC_BATTERY_LOW |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | HD RTC battery low
x | PRS 398 | SA 471 | SRSHD 632 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: Acc Vector |  | HD accelerometer failed POST | If the calculated vector length of the accelerometer x, y, z is not the nominal length. | 8 | ALARM_ID_HD_ACCELEROMETER_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD accelerometer failed POST
x | PRS 863 | SA 472 |  | SRSDG 323 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Acc Vector |  | DG accelerometer failed POST | If the calculated vector length of the accelerometer x, y, z is not the nominal length. | 9 | ALARM_ID_DG_ACCELEROMETER_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG accelerometer failed POST
x | PRS 398 | SA 480 | SRSHD 1478 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: BLD Config Emb Mode |  | HD blood leak sensor embedded mode communication failed | If the blood leak sensor embedded mode is unresponsive. | 10 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_EMBEDDED_MODE_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood leak sensor setting embedded mode failure.
Dialysate Temp | [5 empty] | HD | Low Dialysate Temperature |  | Dialysate temperature is low. \n- Treatment is currently paused.\n- When dialysate temperature is within target range,\n treatment will need to be resumed. | Dial. Gen: Dialysate Temp Low |  | Dialysate temperature is below the target temperature | If the dialysate temperature is below the target temperature for a certain period of time. | 11 | ALARM_ID_HD_DIALYSATE_TEMP_BELOW_TARGET_TEMP |  | ALARM_PRIORITY_MEDIUM | 402 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | HD dialysate temperature below target temperature or below safety temp
x | PRS 398 | SA 195 | SRSHD 658 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: HD Image |  | HD firmware image integrity POST test failed | If the calculated CRC of the HD image does not match the expected CRC. | 12 | ALARM_ID_HD_INTEGRITY_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD firmware image integrity POST test failed
x | PRS 863 | SA 282 |  | SRSDG 813 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: DG Image |  | DG firmware image integrity POST test failed | If the calculated CRC of the DG firmware image does not matched the expected CRC | 13 | ALARM_ID_DG_INTEGRITY_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG firmware image integrity POST test failed
x | PRS 863 | SA 284 |  | SRSDG 957 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Usage Rec NVRAM CRC |  | DG non-volatile usage information group invalid record CRC | If the DG usage information record's CRC does not match the calculated CRC in POST. | 14 | ALARM_ID_DG_INVALID_USAGE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid usage record CRC
Dara - remove Alarm ID 15 (Duplicate with Alarm ID 304 - SRS 1587) | [5 empty] | HD | High Dialysate Temperature |  | Dialysate temperature is high. \n- Treatment is currently paused.\n- When dialysate temperature is within target range,\n treatment will need to be resumed. | Dial. Gen: Dialysate Temp High |  | Dialysate temperature is out of high safety range | If the dialysate temperature is > 46 C for a certain period of time. | 15 | ALARM_ID_HD_DIALYSATE_TEMP_ABOVE_SAFETY_TEMP |  | ALARM_PRIORITY_MEDIUM | 401 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | HD dialysate temperature above high safety.
x | PRS 398 | SA 212 | SRSHD 656 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: Primary Alarm Current |  | HD alarm audio failed POST | If alarm audio current is NOT in expected range for than 1 second while alarm tone is being output | 16 | ALARM_ID_HD_ALARM_AUDIO_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD alarm audio failed POST
Sean to review and update |  |  | SRSHD 1677 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: UI POST Not Complete |  | HD UI POST failed | If UI reports POST failed but did not trigger fault OR UI has not sent POST results within 2 seconds. | 17 | ALARM_ID_HD_UI_POST_TIMEOUT |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST failed
[6 empty] | HD | Communication Error |  | A communication problem has been detected.\n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | Loss of DG Comm (DG ACK) |  | HD to DG comm error. | HD message to DG that requires acknowledge was not acknowledged within a certain time. | 18 | ALARM_ID_HD_CAN_MESSAGE_NOT_ACKED_BY_DG |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD didn't get ACK on message to DG that required acknowledgment
Dialysate Temp |  |  | SRSHD 901 |  |  | HD | High Dialysate Temperature |  | Dialysate temperature is high. \n- Treatment is currently paused.\n- When dialysate temperature is within target range,\n treatment will need to be resumed. | Dial. Gen: Dialysate Temp High |  | Dialysate temperature is above the target temperature | If the dialysate temperature is above the target temperature for a certain period of time. | 19 | ALARM_ID_HD_DIALYSATE_TEMP_ABOVE_TARGET_TEMP |  | ALARM_PRIORITY_MEDIUM | 401 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | HD dialysate temperature above target temperature.
x | PRS 863 | SA 369 |  | SRSDG 16 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Conductivity Cal |  | DG conductivity sensors invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 20 | ALARM_ID_DG_COND_SENSORS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG conductivity sensors invalid calibration record
x | PRS 863 | SA 369 |  | SRSDG 936 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Drain Line Vol Cal |  | DG drain line volume invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 21 | ALARM_ID_DG_DRAIN_LINE_VOLUME_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain line volume invalid calibration record
x | PRS 863 | SA 369 |  | SRSDG 935 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Reservoir Vol Cal |  | DG reservoirs invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 22 | ALARM_ID_DG_RESERVOIRS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG reservoirs invalid calibration record
x | PRS 863 | SA 369 |  | SRSDG 934 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Acid Conc Cal |  | DG acid concentrate calibration record | If the non-volatile data is valid and if not raise the alarm | 23 | ALARM_ID_DG_ACID_CONCENTRATE_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG acid concentrate calibration record
x | PRS 863 | SA 369 |  | SRSDG 933 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Bicarb Conc Cal |  | DG bicarb concentrate calibration record | If the non-volatile data is valid and if not raise the alarm | 24 | ALARM_ID_DG_BICARB_CONCENTRATE_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG bicarb concentrate calibration record
x | PRS 863 | SA 369 |  | SRSDG 817 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Acc Cal |  | DG accelerometer invalid calibration record | If the calibration data that was received from NVDataMgmt is NOT legitimate | 25 | ALARM_ID_DG_ACCELEROMETERS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG accelerometer invalid calibration record
x | PRS 398 | SA 368 | SRSHD 1576 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: Acc Cal |  | HD accelerometer invalid calibration record | If the accel calibration data that was received from NVDataMgmt is NOT legitimate. | 26 | ALARM_ID_HD_ACCELEROMETERS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD accelerometer invalid calibration record
Move to wet self-test - procedural |  |  | SRSHD 1428 |  |  | HD | Self Test: Blood Leak |  | A problem has been detected with the blood leak sensor.\n- Confirm cartridge is properly installed.\n- If problem persists, terminate treatment preparation and \n begin again using a new cartridge. | Process: Pre- Tx Wet Self- Test |  | HD BLD sensor failed zero sequence | BLD fails to complete zero sequence. | 27 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_ZERO_SEQUENCE_FAILED |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD blood leak sensor zero and self test sequence failed
x | PRS 863 | SA 319 SA 320 |  | SRSDG 864 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Conductivity RTD Comm |  | DG two wire sensors FPGA fault | If the FPGA does not report fresh data for a certain period of time | 28 | ALARM_ID_DG_TWO_WIRE_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG two wire sensors FPGA fault
x | PRS 398 | SA 368 | SRSHD 1482 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: SP DAC Cal |  | HD heparin force sensor invalid calibration record | If the heparin force sensor calibration data that was received from NVDataMgmt is NOT legitimate. | 29 | ALARM_ID_HD_HEPARIN_FORCE_SENSOR_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD heparin force sensor invalid calibration record
x | PRS 494 | SA 416 | SRSHD 1704 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: SW Invalid State |  | HD Software fault. Software found itself in an unexpected state | If the specific software faults listed in alarm management SW faults has been detected | 30 | ALARM_ID_HD_SOFTWARE_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD Software fault. Software found itself in an unexpected state
x | PRS 494 | SA 241 SA 242 | SRSHD 515 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: BP Current |  | Blood pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | If the measured MC current is not greater than BP_MAX_CURR_WHEN_STOPPED_MA or is not greater than BP_MAX_CURR_WHEN_RUNNING_MA. | 31 | ALARM_ID_HD_BLOOD_PUMP_MC_CURRENT_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running
x | PRS 494 | SA 239 | SRSHD 1087 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: BP Speed |  | Blood pump failed motor off check. Measured speed while commanded off | If the measured motor speed is greater than BP_MAX_MOTOR_SPEED_WHILE_OFF_RPM | 32 | ALARM_ID_HD_BLOOD_PUMP_OFF_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed motor off check. Measured speed while commanded off
x | PRS 494 | SA 240 SA 244 | SRSHD 1079 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: BP Direction |  | Blood pump failed motor direction check. Measured vs. commanded | If the set direction versus hall sensors direction versus sign of motor speed DOES NOT agree. | 33 | ALARM_ID_HD_BLOOD_PUMP_MC_DIRECTION_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed motor direction check. Measured vs. commanded
x | PRS 494 | SA 246 | SRSHD 1081 |  |  | HD | Blood Pump Error |  | A problem has been detected with the blood pump.\n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue. | HD Fault: BP Rotor |  | Blood pump failed rotor speed check. Mismatch with rotor and motor speeds | If the delta of the measured motor speed and the commanded motor speed is not within tolerance. | 34 | ALARM_ID_HD_BLOOD_PUMP_ROTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed rotor speed check. Mismatch with rotor and motor speeds
x | PRS 494 | SA 250 SA 251 | SRSHD 542 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: DPi Current |  | Dialysis inlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | If the Motor Controller measured DPi pump current > 150.0 mA for more than 250.0 msec. | 35 | ALARM_ID_HD_DIAL_IN_PUMP_MC_CURRENT_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis inlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running
x | PRS 494 | SA 248 | SRSHD 1088 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: DPi Speed |  | Dialysis inlet pump failed motor off check. Measured speed while commanded off | If during command off state, the DPi pump motor speed is measured > 100.0 RMP. | 36 | ALARM_ID_HD_DIAL_IN_PUMP_OFF_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis inlet pump failed motor off check. Measured speed while commanded off
x | PRS 494 | SA 249 SA 253 | SRSHD 1050 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: DPi Direction |  | Dialysis inlet pump failed motor direction check. Measured vs. commanded | If the set direction versus hall sensors direction DO NOT agree for more than 250.0 msec. | 37 | ALARM_ID_HD_DIAL_IN_PUMP_MC_DIRECTION_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis inlet pump failed motor direction check. Measured vs. commanded
x | PRS 494 | SA 255 | SRSHD 1082 |  |  | HD | Dialysate Pump Error |  | A problem has been detected with the dialysate pump.\n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue. | HD Fault: DPi Rotor |  | Dialysis inlet pump failed rotor speed check. Mismatch with rotor and motor speeds | If the measured motor speed versus the measured motor speed while controlling the target > 5.0 for more than 12 seconds. | 38 | ALARM_ID_HD_DIAL_IN_PUMP_ROTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis inlet pump failed rotor speed check. Mismatch with rotor and motor speeds
x | PRS 494 | SA 258 SA 259 | SRSHD 554 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: DPo Current |  | Dialysis outlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running | If the DPo pump current > 150.0 mA while stopped or Dpo current > 2000.0 mA for more than 2 seconds while running. | 39 | ALARM_ID_HD_DIAL_OUT_PUMP_MC_CURRENT_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis outlet pump failed motor controller current check. Too high when pump should be off or out of range when pump should be running
x | PRS 494 | SA 256 | SRSHD 1089 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: DPo Speed |  | Dialysis outlet pump failed motor off check. Measured speed while commanded off | If the Dpo pump current > 100.0 RPM while commanded off for more than 5000 msec. | 40 | ALARM_ID_HD_DIAL_OUT_PUMP_OFF_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis outlet pump failed motor off check. Measured speed while commanded off
x | PRS 494 | SA 257 SA 261 | SRSHD 1051 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: DPo Direction |  | Dialysis outlet pump failed motor direction check. Measured vs. commanded | If the DPo pump set direction does not agree with hall sensors or set direction does not agree with the sign of the motor speed for more than 25 msec. | 41 | ALARM_ID_HD_DIAL_OUT_PUMP_MC_DIRECTION_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis outlet pump failed motor direction check. Measured vs. commanded
x | PRS 494 | SA 263 | SRSHD 1083 |  |  | HD | Dialysate Pump Error |  | A problem has been detected with the dialysate pump.\n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue. | HD Fault: DPo Rotor |  | Dialysis outlet pump failed rotor speed check. Mismatch with rotor and motor speeds | If DPo pump measured rotor speed does not agreed with the measured/commanded motor speed | 42 | ALARM_ID_HD_DIAL_OUT_PUMP_ROTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysis outlet pump failed rotor speed check. Mismatch with rotor and motor speeds
x | PRS 863 | SA 357 |  | SRSDG 954 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Iso ADC Comm |  | Main primary heater FPGA fault | If the FPGA does not report fresh data for a certain period of time | 43 | ALARM_ID_DG_MAIN_PRIMARY_HEATER_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG main primary heater FPGA fault
x | PRS 494 | SA 226 | SRSHD 1001 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: UI Comm |  | User interface communication timeout | If the sub-system communication timeout has occurred. UI is 5 seconds. DG is 2 seconds. | 44 | ALARM_ID_HD_UI_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI communication timeout
x | PRS 494 | SA 226 SA 477 | SRSHD 967 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: CAN CRC Com |  | Too many bad CRCs detected on received system messages | If too many bad messages CRCs are received within a set period of time | 45 | ALARM_ID_HD_COMM_TOO_MANY_BAD_CRCS |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD too many bad communications CRC
x | PRS 494 | SA 226 SA 477 | SRSHD 964 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: CAN ACK UI Com |  | System message that required acknowledgment was not acknowledged | If no retries left in the pending acknowlegement list. | 46 | ALARM_ID_HD_CAN_MESSAGE_NOT_ACKED_BY_UI |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD didn't get ACK on message to UI that required acknowledgment
x | [5 empty] | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: RO Pump High |  | DG RO pump out of upper range | If the RO pump > 95% PWM for 12 secs. | 47 | ALARM_ID_DG_MAX_RO_PUMP_PWM_EXCEEDED |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | DG maximum RO Pump PWM exceeded
x | PRS 494 | SA 452 | SRSHD 1029 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: UF Rate Tare Err |  | Ultrafiltration volume accuracy error during treatment | If the UF volume (refUFvolume - measureVolume) > 100.0 mL. | 48 | ALARM_ID_HD_UF_VOLUME_ACCURACY_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD ultrafiltration volume accuracy error during treatment
x | PRS 494 | SA 224 | SRSHD 1356 SRSHD 1490 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: FPGA Com |  | HD FPGA communication down for too long | If there are more than 3 comm errors in a 1 min period. | 49 | ALARM_ID_HD_FPGA_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD FPGA communication down for too long
x | PRS 863 | SA 481 |  | SRSDG 26 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Valve Control |  | DG FPGA not accepting commanded valve states | If the read back FPGA valve states does not match the commanded valve states for a certain amount of time. | 50 | ALARM_ID_DG_VALVE_CONTROL_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG FPGA not accepting commanded valve states
x | PRS 494 | SA 239 SA 243 | SRSHD 1098 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: BP Speed |  | Blood pump failed motor speed check. Measured vs. commanded | If measured motor speed versus commanded rotor speed exceeded limits. | 51 | ALARM_ID_HD_BLOOD_PUMP_MOTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood pump failed motor speed check. Measured vs. commanded
x | PRS 494 | SA 252 | SRSHD 1099 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: DPi Speed |  | Dialysate inlet pump failed motor speed check. Measured vs. commanded | If measured motor speed versus commanded motor speed exceeded limits for more than 5 secs. | 52 | ALARM_ID_HD_DIAL_IN_PUMP_MOTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysate inlet pump failed motor speed check. Measured vs. commanded
x | PRS 494 | SA 260 SA 256 | SRSHD 1198 SRSHD 1431 SRSHD 1670 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: DPo Speed |  | Dialysate outlet pump failed motor speed check. Measured vs. commanded | If measured motor speed versus commanded motor speed exceeded limits for more than 5 secs. | 53 | ALARM_ID_HD_DIAL_OUT_PUMP_MOTOR_SPEED_CHECK |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD dialysate outlet pump failed motor speed check. Measured vs. commanded
x | PRS 494 | SA 501 | SRSHD 1705 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Critical Data Err |  | HD critical data integrity check failed | If the critical data failed range check. | 54 | ALARM_ID_HD_CRITICAL_DATA_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD critical data integrity check failed
| PRS 863 | SA 500 |  | SRSDG 921 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Critical Data Err |  | DG critical data integrity check failed | If the critical data failed range check. | 55 | ALARM_ID_DG_CRITICAL_DATA_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG critical data integrity check failed
x | PRS 494 | SA 232 | SRSHD 1526 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Acc Com |  | HD accelerometer error (no readings or FPGA reports error) | If no fresh sample of accelerometer data has been detected for more than 100 msec. | 56 | ALARM_ID_HD_ACCELEROMETER_FAILURE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD accelerometer error (no readings or FPGA reports error)
x | PRS 863 | SA 346 |  | SRSDG 869 SRSDG 870 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Acc Com |  | DG accelerometer error (no readings or FPGA reports error) | If no fresh sample of accelerometer data has been detected for more than 100 msec. | 57 | ALARM_ID_DG_ACCELEROMETER_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG accelerometer error (no readings or FPGA reports error)
x | PRS 494 | SA 384 SA 385 SA 386 SA 387 | SRSHD 1340 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Valve Home |  | HD valve homing failed | If the number of failed homing > = 3 counts. | 58 | ALARM_ID_HD_VALVE_HOMING_FAILED |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD valve homing failed
x | PRS 494 | SA 264 SA 266 SA 268 SA 270 | SRSHD 1338 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Valve Transition Err |  | HD valve transition time out | if the valve's transition time has timed out (3000msec) | 59 | ALARM_ID_HD_VALVE_TRANSITION_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD valve transition time out
x | PRS 494 | SA 502 SA 505 SA 503 SA 504 | SRSHD 1515 SRSHD 1516 SRSHD 1632 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Valve Control |  | HD valve not functional | If the valve bit masks are not as expected in different valve control mode. | 60 | ALARM_ID_HD_VALVE_NOT_FUNCTIONAL |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD valve not functional
x | PRS 494 | SA 265 SA 267 SA 269 SA 271 | SRSHD 1339 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Valve Current |  | HD valve current out of range | If the valve current is over the threshold of 1.0 A for 1000 msec. | 61 | ALARM_ID_HD_VALVE_CURRENT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD valve current out of range
x | PRS 494 | SA 264 SA 266 SA 268 SA 270 | SRSHD 1336 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Valve Position |  | HD valve position out of target | If the current position has deviated from the expected position for more than 1 second. | 62 | ALARM_ID_HD_VALVE_POSITION_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD valve position out of target
x | PRS 494 | SA 453 | SRSHD 1586 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Pre- Tx Hep. Pump Timeout. |  | HD syringe pump timeout alarm | If the syringe pump priming > 5 seconds | 63 | ALARM_ID_HD_SYRINGE_PUMP_PRIME_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump prime timeout
x | PRS 863 | SA 350 |  | SRSDG 874 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Baro Range |  | Barometric pressure is out of range | If the barometric pressure is > 10.1 psi or < 15.4 psi for a certain period of time. | 64 | ALARM_ID_DG_BARO_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG barometric pressure sensor out of range
x | PRS 400 |  | SRSHD 1531 |  |  | HD | Communication Error |  | A communication problem has been detected.\n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | Loss of DG Comm (Reject) |  | HD requests DG command with invalid parameter fault | If the latest DG command response indicated invalid parameter | 65 | ALARM_ID_HD_DG_COMMAND_INVALID_PARAMETER_FAULT |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD DG requests DG command with invalid parameter fault
x | PRS 398 | SA 480 | SRSHD 1427 SRSHD 1429 SRSHD 1631 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: BLD Config Set Point |  | HD blood leak sensor set point set failure | If the set point of the blood leak sensor does not match the set point in the calibration record. | 66 | ALARM_ID_HD_BLOOD_LEAK_SENSOR_SET_POINT_SET_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood leak sensor set point set failure
No cartridge self-test (procedural) |  |  | SRSHD 1511 |  |  | HD | Self Test: Cartridge |  | A cartridge has been detected during self- test.\n- Open door\n- Remove cartridge\n- Close door\n- Press Resume/OK | Process: Pre- Tx Remove Cartridge |  | HD blood pump occlusion self-test failure alarm | If the blood pump pressure > 5000 | 67 | ALARM_ID_HD_BP_OCCLUSION_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 700 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD blood pump occlusion self-test failure alarm
agree to change rank and re-assess similar to temperature alarm Blaine |  |  | SRSHD 1426 |  |  | HD | Dialysate Generation Error |  | A problem has been detected with dialysate generation. \n- Treatment is paused and will be need to be resumed when \ndialysate is ready for use. | Dial. Gen: Dialysate not ready |  | HD active reservoir recirculation out of range | If the recirculation level has exceeded the max allowed | 68 | ALARM_ID_HD_ACTIVE_RESERVOIR_RECIRCULATION_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 421 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD active reservoir recirculation out of range
x | PRS 398 | SA 368 | SRSHD 532 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: BLD Cal |  | HD blood leak sensor invalid calibration record | If the BLD data that was received from NVDataMgmt is NOT legitimate. | 69 | ALARM_ID_HD_BLOOD_LEAK_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood leak sensor invalid calibration record
This should be in-treatment alarm? Not machine alarm? See alarm ID 275 for machine alarm (valid range check) agree - make this in-treatment alarms not faults Blaine to update - not a fault Sean to update the range to what is in the SRS |  |  | SRSHD 737 |  |  | HD | Self Test: Cartridge |  | A cartridge has been detected during self- test.\n- Open door\n- Remove cartridge\n- Close door\n- Press Resume/OK | Process: Pre- Tx Remove Cartridge |  | HD arterial pressure self-test failure alarm | If arterial pressure < -300 or > 100 mmHg | 70 | ALARM_ID_HD_ARTERIAL_PRESSURE_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 700 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD arterial pressure self-test failure alarm
This should be in-treatment alarm? Not machine alarm? See alarm ID 275 for machine alarm (valid range check) agree - make this in-treatment alarms not faults Blaine to update - not a fault Sean to update the range to what is in the SRS |  |  | SRSHD 738 |  |  | HD | Self Test: Cartridge |  | A cartridge was detected during self- test.\n- Open door\n- Remove cartridge\n- Close door\n- Press Resume/OK | Process: Pre- Tx Remove Cartridge |  | HD venous pressure self-test failure alarm | If the venous pressure < -100 or > 600 mmHg | 71 | ALARM_ID_HD_VENOUS_PRESSURE_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 700 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD venous pressure self-test failure alarm
x | PRS 400 |  | SRSHD 1126 |  |  | HD | Communication Error |  | A communication problem has been detected.\n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | Loss of DG Comm (LC) |  | HD No load cell data message receive at least once every 2 seconds | No new load cell data for 2 seconds. | 72 | ALARM_ID_HD_NEW_LOAD_CELL_DATA_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD No load cell data message receive at least once every 2 seconds
x | PRS 400 |  | SRSHD 1573 |  |  | HD | Communication Error |  | A communication problem has been detected.\n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | Loss of DG Comm (TDI/TRO) |  | HD No dialysate temperature data message received at least once every 2 seconds | No new dialysate temperature data for 2 seconds. | 73 | ALARM_ID_HD_NEW_DIALYSATE_TEMP_DATA_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD No dialysate temperature data message receive at least once every 2 seconds
x | PRS 863 | SA 364 |  | SRSDG 600 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: UVi On No Flow |  | DG inlet UV reactor is on while there is no flow | If the inlet UV reactor is on while VPi is closed for a certain period of time. | 74 | ALARM_ID_DG_INLET_UV_REACTOR_ON_WITH_NO_FLOW |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG inlet UV reactor is on with no flow
x | PRS 494 | SA 469 | SRSHD 1488 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Syringe DAC |  | HD syringe pump self-test failure alarm | If the calibration data that was received from NVDataMgmt and the DAC voltage with syringe retracted are not within tol. | 75 | ALARM_ID_HD_SYRINGE_PUMP_SELF_TEST_FAILURE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump self-test failure alarm
Sean to update code | PRS 494 | SA 217 SA 218 SA 219 | SRSHD 635 SRSHD 1678 SRSHD 1681 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Voltage Monitor |  | HD monitored voltage is out of range | If the HD monitored voltages are out of range for a certain period of time. | 76 | ALARM_ID_HD_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD monitored voltage is out of range
Sean to update code | PRS 863 | SA 295 SA 296 SA 297 |  | SRSDG 46 SRSDG 866 SRSDG 1088 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Voltage Monitor |  | DG monitored voltage is out of range | If the DG monitored voltages are out of range for a certain period of time. | 77 | ALARM_ID_DG_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG monitored voltage is out of range
x | PRS 494 | SA 275 | SRSHD 1324 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Syringe Pump Enc Dir |  | HD syringe pump direction (from encoder) error | If the measured encoder direction of the syringe pump does not agree with the expected direction | 78 | ALARM_ID_HD_SYRINGE_PUMP_ENCODER_DIRECTION_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump direction (from encoder) error
x | PRS 494 | SA 489 | SRSHD 1507 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Syringe Pump Drv Dir |  | HD syringe pump direction (from controller) error | If the measured direction of the syringe pump motor controller does not agree with the expected direction | 79 | ALARM_ID_HD_SYRINGE_PUMP_CONTROLLER_DIRECTION_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump direction (from controller) error
x | PRS 494 | SA 272 | SRSHD 1317 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Syringe Pump Driver |  | HD syringe pump fault reported by FPGA | If the FPGA reported the syringe pump fault | 80 | ALARM_ID_HD_SYRINGE_PUMP_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump fault reported by FPGA
x | PRS 494 | SA 488 | SRSHD 1330 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Syringe Pump OT |  | HD syringe pump over travel error | If the syringe pump position is deemed to have overtraveled | 81 | ALARM_ID_HD_SYRINGE_PUMP_OVER_TRAVEL_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump over travel error
x | PRS 494 | SA 277 | SRSHD 1506 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Syringe Pump DAC Err |  | HD syringe pump DAC write failure | If the syringe pump DAC write error is reported by the FPGA | 82 | ALARM_ID_HD_SYRINGE_PUMP_DAC_WRITE_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump DAC write failure
x | PRS 494 | SA 415 | SRSHD 1328 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Syringe Pump On Err |  | HD syringe pump is running while the BP is off | If the syringe pump is on while the blood pump is off | 83 | ALARM_ID_HD_SYRINGE_PUMP_RUNNING_WHILE_BP_OFF_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump is running while the BP is off
x | PRS 863 | SA 286 |  | SRSDG 815 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: DG RTC Time |  | DG RTC set year is invalid | If DG RTC set year is < 2021. | 84 | ALARM_ID_DG_SET_RTC_YEAR_INVALID |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG set RTC year is invalid
revisit - PRS has this as low priority - why did SRS make this high priority? Probably needs to be high priority as blood pump needs to stop - update PRS to make high priority | [5 empty] | HD | Cartridge Lever is Open |  | A problem has been detected with the cartridge.\n- Open the device door and push lever all the way down.\n- Close the device door.\n- Press Resume/OK button to continue. | Blood Stop: Cartridge Lever Open |  | HD pump track lever is not fully latched | If the pump track lever is detected as unlatched when it should be. | 85 | ALARM_ID_HD_PUMP_TRACK_LATCH_OPENED |  | ALARM_PRIORITY_HIGH | 209 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD pump track latch opened alarm
x | PRS 398 | SA 199 | SRSHD 1479 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: HD RTC Time |  | HD RTC set year is invalid | If HD RTC set year is < 2021. | 86 | ALARM_ID_HD_SET_RTC_YEAR_INVALID |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD set RTC year is invalid
x | PRS 863 | SA 369 |  | SRSDG 929 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Heat Cal |  | DG heating parameters invalid calibration record | If the heating paramters' non-volatile record's CRC does not match the calculated CRC. | 87 | ALARM_ID_DG_HEATING_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG heating invalid calibration record
[6 empty] | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: CP1 CP2 Speed Range |  | DG concentrate pump hall sensor out of range. | Hall sensor pulse to pulse timing is < 1ms (speed too fast). | 88 | ALARM_ID_DG_CONC_PUMP_HALL_SENSOR_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG concentrate pumps hall sensor out of range
x | PRS 863 | SA 365 |  | SRSDG 599 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: UVo On No Flow |  | DG outlet UV reactor is on while there is no flow | If the outlet UV reactor is on while the RO flow is less than 200 mL/min. | 89 | ALARM_ID_DG_OUTLET_UV_REACTOR_ON_WITH_NO_FLOW |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG outlet UV reactor on with no flow
x | PRS 863 | SA 307 SA 308 |  | SRSDG 604 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: LC A1/B1 Comm |  | DG load cells A1/B1 FPGA fault | If the FPGA does not report fresh data for a certain period of time. | 90 | ALARM_ID_DG_LOAD_CELL_A1_B1_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells A1/B1 FPGA fault
x | PRS 400 |  | SRSHD 1408 |  |  | HD | Communication Error |  | A communication problem has been detected.\n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | Loss of DG Comm (FMD) |  | HD No dialysate flow data received in the last 3 seconds | If not receiving new dialysate flow readings in expected interval (3 seconds) | 91 | ALARM_ID_HD_DIALYSATE_FLOW_DATA_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD No dialysate flow data receive in the last 3 seconds
Remove (redundant with SRSDG 913) | [13 empty] | 92
[14 empty] | 92 | ALARM_ID_AVAILABLE_1 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Available for use
x | PRS 863 | SA 380 |  | SRSDG 430 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: RO Permeate Cond High |  | DG outlet primary conductivity sensor out of range | If the outlet primary conductivity sensor, CPo is greater than 30 uS/cm for a certain period of time. | 93 | ALARM_ID_DG_OUTLET_PRIMARY_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG outlet primary conductivity out of range
x | PRS 863 | SA 303 SA 304 SA 305 SA 306 |  | SRSDG 629 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Pressure Sensor Range |  | DG pressure sensor out of range | If either of the pressure sensors is < 0 psia or > 200 psia for DG pressure sensors. | 94 | ALARM_ID_DG_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG pressure out of range
x | PRS 863 | SA 486 |  | SRSDG 907 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: WD Error |  | DG watchdog expired | If the watchdog signal is low while not in POST. | 95 | ALARM_ID_DG_WATCHDOG_EXPIRED |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG watchdog expired
x | PRS 406 |  |  | SRSDG 410 |  | DG | Incoming Water Temperature High |  | Incoming water temperature is high.\n- Treatment is currently paused.\n- Check the incoming water supply.\n- When the temperature is within target range\nResume/OK button will be available.\n- Press Resume/OK to continue treatment. | Dial. Gen: Inlet Water Temp High |  | DG inlet water temperature out of range | If the inlet water temperature is > 37 C. | 96 | ALARM_ID_DG_INLET_WATER_TEMPERATURE_IN_HIGH_RANGE |  | ALARM_PRIORITY_MEDIUM | 432 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water temperature in high range
Check that software behavior and PRS behavior are aligned | PRS 543 |  |  | SRSDG 397 |  | DG | Conductivity Out of Range |  | Treatment is currently paused.\n- Confirm that acid and bicarbonate concentrates\nare securely connected.\n- If empty, replace. \n- Treatment will need to be resumed when conductivity\nis within target range. | Dial. Gen: Dialysate Cond. |  | DG fill conductivity out of range | If the average acid or bicard conductivity is > 5% of the normal acid and bicarb conductivity values. | 97 | ALARM_ID_DG_FILL_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 420 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | DG fill conductivity out of range
x | PRS 872 | SA 227 SA 228 | SRSHD 1395 |  |  | HD | Service Required: Hemodialysis Device |  | A non- critical problem was detected with the hemodialysis\n device.\n- Call service to report the issue and schedule a repair. | HD Alarm (Low): Battery Com |  | HD battery or battery charger communication fault | If battery I2C communication fault persistence count > 5. | 98 | ALARM_ID_HD_BATTERY_COMM_FAULT |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | HD battery communication fault
x | PRS 494 | SA 274 | SRSHD 1683 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Syringe Pump Stall |  | HD syringe pump stall alarm | If syringe pump measured rate < stall speed threshold (0.05) and the ramp up has been restarted 3 times after stall has been detected. | 99 | ALARM_ID_HD_SYRINGE_PUMP_STALL |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump stall alarm
x | [13 empty] | 100 | ALARM_ID_AVAILABLE_12 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Available for use
[6 empty] | UI | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: UI- Proc Comm |  | UI POST FW Communication failure | If the HD is not communicating with the UI during POST. | 101 | ALARM_ID_UI_POST_HD_COMM |  | ALARM_PRIORITY_HIGH | 0 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | UI POST HD communication failure
cleaning | PRS 863 | SA 482 |  | SRSDG 856 SRSDG 930 |  | DG | Disinfection Interrupted |  | Disinfection terminated unexpectedly.\n- Restart disinfection.\n- If problem persists, call service to schedule a repair. | Cleaning: Tpo THD Temp Delta |  | DG heat disinfection temperature gradient out of range | If temperature gradient between Tpo and Thd exceeds maximum 15°C. | 102 | ALARM_ID_DG_HEAT_DISINFECT_TEMP_GRAD_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG heat disinfect temperature gradient out of range
x | PRS 398 | SA 196 | SRSHD 1059 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: Cal NVRAM CRC |  | HD non-volatile calibration group invalid record CRC | If the computed CRC of the calibration group record is not equal to the record CRC. | 103 | ALARM_ID_HD_NVDATAMGMT_CAL_GROUP_RECORD_CRC_INVALID |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD invalid calibration CRC
x | PRS 494 | SA 478 | SRSHD 1115 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Air Trap Sensor |  | Air trap level sensors reporting illegal combination of air/fluid | If the air trap level readings are not valid for more than 2 seconds. | 104 | ALARM_ID_HD_AIR_TRAP_ILLEGAL_LEVELS |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD air trap level sensors reporting illegal combination of air/fluid
x | PRS 863 | SA 283 |  | SRSDG 608 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Cal NVRAM CRC |  | DG non-volatile system group invalid record CRC | If the DG NVDATAMGMT CAL group calculated CRC does not matched the expected CRC. | 105 | ALARM_ID_DG_NVDATAMGMT_CAL_GROUP_RECORD_CRC_INVALID |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid calibration CRC
x | PRS 863 | SA 369 |  | SRSDG 820 SRSDG 821 SRSDG 1087 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Flow Sensor Cal |  | DG flow sensors calibration error | If the calculated CRC of the flow sensors calibration record does not match the calibration records' CRC. | 106 | ALARM_ID_DG_FLOW_SENSORS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG dialysate flow sensor invalid calibration record
No SRS ID - need to add then assess Move to DG machine alarm? (allow properties table to be aligned) Sean - to revisit code to assess | [5 empty] | HD | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: DG Restarted |  | HD reports DG restarted fault | If DG reports being in any mode other than idle, drain, fill or fault during pre-tx, tx, or post-tx modes - means DG restarted unexpectedly and is no longer in sync w/ HD. | 107 | ALARM_ID_HD_DG_RESTARTED_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD reports DG restarted fault
x | PRS 494 | SA 276 | SRSHD 1318 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Syringe Pump ADC Com |  | HD syringe pump ADC error | If the ADC read is not fresh after POST init and the ADC read counter has not changed. | 108 | ALARM_ID_HD_SYRINGE_PUMP_FPGA_ADC_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump ADC FPGA fault
x | PRS 494 | SA 490 | SRSHD 580 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Syringe Pump Vol |  | HD syringe pump volume check error | If the volume delivered is off by more than 5% or 0.1mL. | 109 | ALARM_ID_HD_SYRINGE_PUMP_VOLUME_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump volume check error
x | PRS 494 | SA 274 | SRSHD 579 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Syringe Pump Speed |  | HD syringe pump speed check error | If the syringe measured rate is off by more than 5% or 0.1mL. | 110 | ALARM_ID_HD_SYRINGE_PUMP_SPEED_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump speed check error
x | PRS 494 | SA 274 | SRSHD 1333 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Syringe Pump Speed |  | HD syringe pump not stopped in off state error | If the syringe pump position is not equal to its last position. | 111 | ALARM_ID_HD_SYRINGE_PUMP_NOT_STOPPED_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD syringe pump not stopped in off state error
x | PRS 863 | SA 312 SA 313 |  | SRSDG 953 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: LC A2/B2 Comm |  | DG load cells A2/B2 FPGA fault | If the FPGA does not report fresh data for a certain period of time. | 112 | ALARM_ID_DG_LOAD_CELL_A2_B2_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells A1/B1 FPGA fault
x | PRS 863 | SA 341 |  | SRSDG 890 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: DRP Current |  | DG drain pump current out of range | If the DG drain pump current > 2.2 A when it is in open loop or if the current > 0.1 A when it is off. | 113 | ALARM_ID_DG_DRAIN_PUMP_CURRENT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG Drain pump current out of range
move to wet self-test? | PRS 494 | SA 479 | SRSHD 1532 |  |  | HD | Self Test: Venous Air Detector |  | A problem has been detected with the venous air detector.\n- Check cartridge installation.\n- If problem persists, terminate treatment preparation and \n begin again using a new cartridge. | Process: Pre- Tx Wet Self- Test |  | HD venous air bubble detector self-test failure | If the venous air bubble self test did not detect bubble for more than 500 msec. | 114 | ALARM_ID_HD_VENOUS_BUBBLE_SELF_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD venous air bubble detector self-test failure
Sean - update code to remove temperatures from code that are no longer needed (see updated SRSDG 224) | PRS 863 | SA 321 SA 322 SA 323 SA 324 SA 328 SA 332 SA 336 SA 298 SA 460 |  | SRSDG 224 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Temp Sensor Range |  | DG temperature sensor out of range | If the measured temperature < 0 C or > 80 C for the non-fluid inline sensors and > 120 C for the fluid sensors. | 115 | ALARM_ID_DG_TEMPERATURE_SENSOR_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG temperature sensor out of range
[6 empty] | UI HD | Service Required: Dialysate Device |  | A non- critical problem was detected with the hemodialysis\n device.\n- Call service to report the issue and schedule a repair. | HD Alarm (Low): SD- Card Error |  | SD card is non-operational | SD card is non-operational | 116 | ALARM_ID_HD_UI_SDCARD_FAILURE |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | HD UI SD card failure
|  |  |  | SRSDG 909 |  | DG | Communication Error |  | A communication problem has been detected.\n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | Loss of HD Comm (HD ACK) |  | System message that required acknowledgment was not acknowledged | If no retries left in the pending acknowlegement list. | 117 | ALARM_ID_DG_CAN_MESSAGE_NOT_ACKED |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CAN message not acked
x | PRS 863 | SA 458 |  | SRSDG 43 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: RTC Config |  | DG RTC configuration error | If the RTC configuration registers report an error. | 118 | ALARM_ID_DG_RTC_CONFIG_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RTC config error
x | PRS 874 | SA 499 |  | SRSDG 908 |  | DG | Service Required: Dialysate Device |  | A non- critical problem was detected with the dialysate device.\n- Call service to report the issue and schedule a repair. | DG Alarm (Low): RTC Bat Low |  | DG RTC battery low | If the RTC battery bit report a low alarm. | 119 | ALARM_ID_DG_RTC_BATTERY_LOW |  | ALARM_PRIORITY_LOW | 851 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG RTC battery low error
|  |  | SRSHD 777 |  |  | HD | Self Test: Venous Air Detector |  | A problem has been detected with the venous air detector\n- Check cartridge installation.\n- If problem persists, terminate treatment preparation and \n start over using a new cartridge. | Process: Pre- Tx Wet Self- Test |  | HD pre-treatment wet prime test failure | If the HD pre-treatment wet self test failed. | 120 | ALARM_ID_HD_PRE_TREATMENT_WET_PRIME_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode wet self-test prime check failure
x | PRS 863 | SA 363 |  | SRSDG 611 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Prim Heater Lg Control |  | DG main primary heater voltage out of range | If the main primary heater's voltage is out of range. | 121 | ALARM_ID_DG_MAIN_PRIMARY_HEATER_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG main primary heater voltage out of range
x | PRS 863 | SA 362 |  | SRSDG 612 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Prim Heater Sm Control |  | DG small primary heater voltage out of range | If the small primary heater's voltage is out of range. | 122 | ALARM_ID_DG_SMALL_PRIMARY_HEATER_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG small primary heater voltage out of range
x | PRS 863 | SA 361 |  | SRSDG 613 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Trimmer Heater Control |  | DG trimmer heater voltage out of range | If the trimmer voltage is out of range. | 123 | ALARM_ID_DG_TRIMMER_HEATER_VOLTAGE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG trimmer heater voltage out of range
x | PRS 390 |  | SRSHD 954 |  |  | HD | Blood Pump Stop Time Exceeded |  | The blood pump has been off for too long.\n- Treatment must be terminated without rinseback. | Blood Stop: Blood Clot Risk |  | HD end of treatment alarm (high priority - user still not acting to end treatment) | In treatment end paused state for > 5 minutes. | 124 | ALARM_ID_HD_END_OF_TREATMENT_HIGH |  | ALARM_PRIORITY_HIGH | 11 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD end of treatment alarm (high priority - user still not acting to end treatment)
x | PRS 390 |  | SRSHD 948 |  |  | HD | Blood Pump Stop Time Exceeded |  | The blood pump has been off for too long.\n- Treatment must be terminated without rinseback. | Blood Stop: Blood Clot Risk |  | Blood sitting too long after treatment stopped by user alarm (> 5 min) | If in treatment stop sub mode when blood is in the tubing, the blood sitting counter exceeded 5 minutes. | 125 | ALARM_ID_HD_TREATMENT_STOPPED_NO_RINSEBACK |  | ALARM_PRIORITY_HIGH | 11 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood sitting too long after treatment stopped by user alarm (>5 min)
x | PRS 387 |  | SRSHD 925 |  |  | HD | Blood Leak |  | A blood leak has been detected. \n- Visually inspect dialysate return line for blood.\n- Press Resume/OK to proceed to the next step. Dialysate will\n remain in bypass. | Blood Stop: Dialyzer Blood Leak |  | HD blood leak detected alarm | If the blood leak is detected based on the blood leak status reported by the sensor. | 126 | ALARM_ID_HD_BLOOD_LEAK_DETECTED |  | ALARM_PRIORITY_HIGH | 150 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | TRUE | HD blood leak detected alarm
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 911 |  |  | HD | Low Venous Pressure |  | Venous pressure is low.\n- Ensure venous access connections are properly secured.\n- If there has been an access disconnection or needle\n dislodgement, follow facility policy to address and\nterminate treatment without rinseback.\n- If access connections are properly secured, \npress Resume/OK to continue treatment. | Blood Stop: Venous Pres. Low |  | Venous pressure too low during treatment | If the venous pressure < TREATMENT_PARAM_VEN_PRESSURE_LOW_LIMIT. Or expanded pressure limits based on mode of operation. | 127 | ALARM_ID_HD_VENOUS_PRESSURE_LOW |  | ALARM_PRIORITY_HIGH | 200 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | HD venous pressure too low during treatment
x | PRS 494 | SA 506 | SRSHD 919 |  |  | HD | Venous Air Detected |  | Air detected in the venous air trap.\n- Confirm arterial access placement and confirm\nconnections are secure.\n- Treatment must be terminated without rinseback. | Blood Stop: Venous Air Bubble |  | HD venous air bubble detected alarm | If bubbles are detected during the treatment. | 128 | ALARM_ID_HD_VENOUS_BUBBLE_DETECTED |  | ALARM_PRIORITY_HIGH | 12 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | HD venous air bubble detected alarm
x | [5 empty] | HD | Blood Leak: Bypass |  | The device is currently in bypass due to a potential blood leak.\n- If blood is present upon visual inspection, press\n END TREATMENT to terminate treatment without rinseback.\n- If blood is not visualized, check dialysate for blood using\n test strip.\n- Follow facility policy to Resume/OK or END TREATMENT. | Blood Stop: Dialyzer Blood Leak Bypass |  | HD blood leak detected and it is in the process of attempting to remove the bubbles/foam | If the blood leak is detected in treatment, the HD device transitions to treatment stop and runs in bypass until there is no blood detected. The user can hit end treatment anytime. | 129 | ALARM_ID_HD_BLOOD_LEAK_RECOVERING_PLEASE_WAIT |  | ALARM_PRIORITY_MEDIUM | 301 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | HD blood leak recovering please wait
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 913 |  |  | HD | High Venous Pressure |  | Venous pressure is high. \n- Check venous bloodlines for kinks or clamps.\n- Check needle placement and/or access position.\n- Resolve access issues, if applicable.\n- If clots are present, end treatment without rinseback. | Blood Stop: Venous Pres. High |  | Venous pressure too high during treatment | If the venous pressure > TREATMENT_PARAM_VEN_PRESSURE_HIGH_LIMIT. Or expanded pressure limits based on mode of operation. | 130 | ALARM_ID_HD_VENOUS_PRESSURE_HIGH |  | ALARM_PRIORITY_HIGH | 201 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | HD venous pressure too high during treatment
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 915 |  |  | HD | Low Arterial Pressure |  | Arterial pressure is low.\n- Inspect arterial bloodlines for kinks or clamps.\n- Check needle placement or access position.\n- Resolve access issues, if applicable.\n- If clots are present, end treatment without rinseback. | Blood Stop: Arterial Pres. Low |  | Arterial pressure too low during treatment | If the arterial pressure < TREATMENT_PARAM_ART_PRESSURE_LOW_LIMIT. Or expanded pressure limits based on mode of operation. | 131 | ALARM_ID_HD_ARTERIAL_PRESSURE_LOW |  | ALARM_PRIORITY_HIGH | 202 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | HD arterial pressure too low during treatment
meeting to discuss - see eliza's email on changes to venous/arterial alarms |  |  | SRSHD 917 |  |  | HD | High Arterial Pressure |  | Arterial pressure is high.\n- Ensure arterial access connections are properly secured.\n- Check needle placement and/or access position.\n- Resolve access issues if applicable.\n- If clots are present, end treatment without rinseback. | Blood Stop: Arterial Pres. High |  | Arterial pressure too high during treatment | If the arterial pressure > TREATMENT_PARAM_ART_PRESSURE_HIGH_LIMIT. Or expanded pressure limits based on mode of operation. | 132 | ALARM_ID_HD_ARTERIAL_PRESSURE_HIGH |  | ALARM_PRIORITY_HIGH | 203 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | HD arterial pressure too high during treatment
x | PRS 863 | SA 419 |  | SRSDG 328 |  | DG | Dialysate Device Fluid Leak |  | Fluid leak detected in the dialysate device.\n- Treatment must be terminated.\n- Turn off water source. \n- Disconnect water and electrical connections from the device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Fluid Leak in DG |  | DG fluid leak detected alarm | If the fluid leak is detected as reported by the fluid leak detector. | 133 | ALARM_ID_DG_FLUID_LEAK_DETECTED |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | DG fluid leak detected alarm
x | PRS 409 |  | SRSHD 927 |  |  | HD | Fluid Leak |  | Fluid leak detected in hemodialysis device.\n- If in treatment, blood pump is off.\n- Identify and correct the cause of the fluid leak.\n- Clean up fluid and dry leak sensor.\n- Once fluid has been cleaned up, Resume/OK treatment. | Blood Stop: Fluid Leak in HD |  | HD fluid leak detected alarm | If the fluid leak is detected as reported by the fluid leak detector. | 134 | ALARM_ID_HD_FLUID_LEAK_DETECTED |  | ALARM_PRIORITY_HIGH | 208 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | TRUE | TRUE | FALSE | TRUE | HD fluid leak detected alarm
x | PRS 411 |  | SRSHD 938 |  |  | HD | Blood Pump Stopped: Excessive Movement |  | Excessive movement of the hemodialysis device has been\ndetected. \n- Treatment is currently paused. \n- When system movement has stablized, \ntreatment will need to be resumed.\n- To resume treatment, press Resume/OK. | Blood Stop: HD Shock Detected |  | HD experienced large acceleration | If the system experienced excessive shock > 2.5 g. | 135 | ALARM_ID_HD_SHOCK |  | ALARM_PRIORITY_HIGH | 220 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | HD experienced large acceleration
x | PRS 686 |  |  | SRSDG 327 |  | DG | Dialysate Device: Excessive Movement |  | Excessive movement of the dialysate device has been\ndetected. \n- Treatment is currently paused. \n- When system movement has stablized, \ntreatment will need to be resumed.\n- To resume treatment, press Resume/OK. | Dial. Gen: Shock- Vibration |  | DG experienced large acceleration | If the system experienced excessive shock > 2.5 g. | 136 | ALARM_ID_DG_SHOCK |  | ALARM_PRIORITY_MEDIUM | 440 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | DG experienced large acceleration
x | PRS 412 |  | SRSHD 936 |  |  | HD | Blood Pump Stopped: Excessive Tilt |  | Excessive tilt of the hemodialysis device has been detected. \n- Treatment is currently paused. \n- When the device is stationary and level, \ntreatment will need to be resumed.\n- To resume treatment, press Resume/OK. | Blood Stop: HD Tilted |  | HD tilt exceeds maximum | If the system experienced excessive tilt x and y axis angle > 3 deg for more than 1 sec. | 137 | ALARM_ID_HD_EXCESSIVE_TILT |  | ALARM_PRIORITY_HIGH | 221 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | HD tilt exceeds maximum
x | PRS 687 |  |  | SRSDG 326 |  | DG | Dialysate Device: Excessive Tilt |  | Excessive tilt of the dialysate device has been detected.\n- Treatment is paused.\n- Confirm the device is stationary and level before continuing. | Dial. Gen: Tilt |  | DG tilt exceeds maximum | If the system experienced excessive tilt x and y axis angle > 3 deg for more than 1 sec. | 138 | ALARM_ID_DG_EXCESSIVE_TILT |  | ALARM_PRIORITY_MEDIUM | 441 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | DG tilt exceeds maximum
x | PRS 823 |  | SRSHD 1593 |  |  | HD | Power Loss |  | AC Power has been lost.\n- Check all electrical connections.\n- Press Resume/OK to return to welcome screen. | Loss of AC Power |  | HD AC power lost alarm | The HD Software shall trigger an alarm if an AC power loss is detected while the HD system is in standby, service, or treatment parameters mode. | 139 | ALARM_ID_HD_AC_POWER_LOST |  | ALARM_PRIORITY_LOW | 852 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD AC power lost alarm
x | PRS 400 |  | SRSHD 1124 |  |  | HD | Communication Error |  | A communication problem has been detected.\n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | Loss of DG Comm (Timeout) |  | Dialysate generator communication timeout | If DG communication timeout is detected for more than 2 seconds. | 140 | ALARM_ID_HD_DG_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | HD DG communication timeout
|  |  | SRSHD 1169 |  |  | HD | Blood Pump Stopped: Venous Air Detector |  | A problem was detected with the venous air detector.\n- Confirm arterial access placement and ensure\nconnections are secure.\n- Treatment must be terminated without rinseback. | Blood Stop: Air Trap Fill |  | Air trap fill timeout during treatment | If the air trap fill timeout during treatment. | 141 | ALARM_ID_HD_AIR_TRAP_FILL_DURING_TREATMENT |  | ALARM_PRIORITY_HIGH | 13 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | HD air trap fill timeout during treatment
x | PRS 496 |  | SRSHD 1361 SRSHD 1362 |  |  | HD | Blood Pump Stopped: Occlusion |  | Occlusion detected after the blood pump.\n- Inspect bloodlines for kinks or clamps.\n- Check needle placement or access position.\n- Resolve access issues, if applicable.\n- If clots are present, end treatment without rinseback. | Blood Stop: Blood Pump Occlusion |  | Blood pump occlusion detected | If blood occlusion is detected > OCCLUSION_THRESHOLD_OFFSET + bloodPumpOcclusionAfterCartridgeInstall | 142 | ALARM_ID_HD_OCCLUSION_BLOOD_PUMP |  | ALARM_PRIORITY_HIGH | 206 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | HD blood pump occlusion detected
x | PRS 863 | SA 338 |  | SRSDG 912 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: TRo TDi Delta |  | DG dialysate temperature sensors drift timeout. | If TDi and TRo have drifted for more 1 C for a certain period of time. | 143 | ALARM_ID_DG_DIALYSATE_TEMPERATURE_SENSORS_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG dialysate temperature sensors out of range
x | PRS 853 |  |  | SRSDG 1066 SRSDG 1054 SRSDG 970 |  | DG | Disinfection: Water Conductivity High |  | Incoming water conductivity is too high.\n- Disinfection has been terminated.\n- Check incoming water before initiating disinfection again. | Cleaning: Inlet Water Cond High |  | DG cleaning mode inlet water conductivity out of high range | Inlet water conductivity > 2000 uS/cm. Inlet water conductivity > 100 uS/cm (PURE WATER MODE). In cleaning mode | 144 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_COND_TOO_HIGH |  | ALARM_PRIORITY_MEDIUM | 603 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water conductivity out of high range.
x | PRS 394 |  |  | SRSDG 846 |  | DG | Conductivity Error During Prime |  | Conductivity out of range during prime.\n- Confirm that acid and bicarbonate concentrates\nare securely connected.\n- If empty, replace. \n- Treatment will need to be resumed when conductivity\nis within target range. | Dial. Gen: Acid Conc. Cond. |  | DG Concentrate conductivity after adding acid out of range alarm | If the conductivity value for the acid concentrate is out of range for more than 45 secs. | 145 | ALARM_ID_DG_ACID_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 410 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | DG concentrate conductivity after adding acid out of range alarm
x | PRS 863 | SA 285 |  | SRSDG 312 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: DG RTC Timer |  | DG RTC vs. Timer accuracy error | If RTC and timer are not aligned on 1 second elapsed time. | 146 | ALARM_ID_DG_RTC_OR_TIMER_ACCURACY_FAILURE |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RTC or timer accuracy failure.
x | PRS 543 |  |  | SRSDG 398 |  | DG | Dialysate Generation Error |  | A problem has been detected with dialysate generation. \n- Treatment is paused and will need to be resumed when \nthe dialysate is ready for use. | Dial. Gen: Dialysate not ready |  | DG dialysate generation conductivity fault alarm | If dialysate conductivity is not with 5% of the TCD | 147 | ALARM_ID_DG_CREATING_DIALYSATE_PLEASE_WAIT |  | ALARM_PRIORITY_MEDIUM | 421 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | DG wait for the DG to produce dialysate
x | PRS 557 |  |  | SRSDG 1065 SRSDG 1053 SRSDG 969 |  | DG | Disinfection: Water Temperature High |  | Incoming water temperature is too high.\n- Disinfection has been terminated.\n- Check incoming water before initiating disinfection again. | Cleaning: Inlet Water Temp High |  | Cleaning Mode Inlet water temperature too high | Inlet temperature > 37 C. | 148 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_TEMP_TOO_HIGH |  | ALARM_PRIORITY_MEDIUM | 601 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG Cleaning mode inlet water temperature too high
x | PRS 494 | SA 475 | SRSHD 646 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: WD Error |  | HD watchdog expired | If the watchdog signal is low while not in POST. | 149 | ALARM_ID_HD_WATCHDOG_EXPIRED |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD watchdog expired
x | PRS 404 |  |  | SRSDG 391 |  | DG | Incoming Water Conductivity High |  | Conductivity of incoming water is high.\n- Treatment is currently paused.\n- Check the incoming water supply.\n- When the conductivity is within target range\nResume/OK button will be available.\n- Press Resume/OK to continue treatment. | Dial. Gen: Inlet Water Cond High |  | Inlet water conductivity too high | Inlet water conductivity > 2000 uS/cm. Inlet water conductivity > 100 uS/cm (PURE WATER MODE). | 150 | ALARM_ID_DG_INLET_WATER_CONDUCTIVITY_IN_HIGH_RANGE |  | ALARM_PRIORITY_MEDIUM | 434 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water conductivity in high range
x | PRS 403 |  |  | SRSDG 826 |  | DG | Incoming Water Conductivity Low |  | Conductivity of incoming water is low.\n- Treatment is currently paused.\n- Check the incoming water supply.\n- When the conductivity is within target range\nResume/OK button will be available.\n- Press Resume/OK to continue treatment. | Dial. Gen: Inlet Water Cond Low |  | Inlet water conductivity too low | Inlet water conductivity < 200 uS/cm. Inlet water conductivity low disabled (PURE WATER MODE). | 151 | ALARM_ID_DG_INLET_WATER_CONDUCTIVITY_IN_LOW_RANGE |  | ALARM_PRIORITY_MEDIUM | 435 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water conductivity in low range
x | PRS 401 |  |  | SRSDG 417 |  | DG | Incoming Water Pressure Low |  | Water pressure of incoming water is low.\n- Treatment is currently paused.\n- Check the incoming water supply.\n- When the pressure is within target range\nResume/OK button will be available.\n- Press Resume/OK to continue treatment. | Dial. Gen: Inlet Water Pres Low |  | Inlet water pressure too low | Inlet water pressure < 25 PSI. | 152 | ALARM_ID_DG_INLET_WATER_PRESSURE_IN_LOW_RANGE |  | ALARM_PRIORITY_MEDIUM | 431 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water pressure in low range
x | PRS 529 |  | SRSHD 1122 |  |  | HD | Pre-Treatment Process Terminated |  | The pre- treatment process has been terminated because \ntreatment did not begin within two hours of priming. \n- Press Resume/OK for continue with next steps to remove \ncartridge from the device. | Process: Pre- Tx Terminated |  | HD prime completed high priority alarm | In Pre-Treatment Recirculation for 120 minutes (escalation of alarm 213) | 153 | ALARM_ID_HD_PRIME_COMPLETED_HIGH |  | ALARM_PRIORITY_HIGH | 292 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD prime completed high priority alarm
x | PRS 854 |  |  | SRSDG 1067 SRSDG 1055 SRSDG 971 |  | DG | Disinfection: Water Conductivity Low |  | Incoming water conductivity is too low.\n- Disinfection has been terminated.\n- Check incoming water before initiating disinfection again. | Cleaning: Inlet Water Cond Low |  | Cleaning mode inlet water conductivity too low | Inlet water conductivity < 200 uS/cm. Inlet water conductivity low disabled (PURE WATER MODE). In cleaning mode | 154 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_COND_TOO_LOW |  | ALARM_PRIORITY_MEDIUM | 604 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water conductivity out of low range.
x | PRS 400 |  | SRSHD 1574 |  |  | HD | Communication Error |  | A communication problem has been detected.\n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | Loss of DG Comm (Reservoir) |  | HD No reservoir data message receive at least once every 2 seconds | No new reservoir data for 2 seconds. | 155 | ALARM_ID_HD_NEW_RESERVOIRS_DATA_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD No reservoirs data message receive at least once every 2 seconds
x | PRS 400 |  | SRSHD 1575 |  |  | HD | Communication Error |  | A communication problem has been detected.\n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | Loss of DG Comm (Op Mode) |  | HD No DG op mode message receive at least once every 2 seconds | No DG op mode for 2 seconds. | 156 | ALARM_ID_HD_DG_NEW_OPERATION_MODE_MESSAGE_NOT_RECEIVE |  | ALARM_PRIORITY_HIGH | 111 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD No DG operation mode message receive at least once every 2 seconds
cleaning |  |  |  | SRSDG 988 |  | DG | Chemical Disinfectant Not Detected |  | Unable to sense proper chemical disinfectant.\n- Disinfection process has terminated.\n- Confirm chemical disinfectant is connected to the\nconcentrate inlet. \n- Confirm chemical disinfectant container is not empty.\n- Restart disinfection.\n- If problem persists, call service to schedule a repair. | Cleaning: Chem Disinf PAA Prime |  | Chemical disinfect priming the acid line timeout | If the disinfectant acid line is not primed within a certain period of time. | 157 | ALARM_ID_DG_CHEM_DISINFECT_PRIME_ACID_LINE_TIME_OUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG chemical disinfect prime acid line timeout.
x | PRS 405 |  |  | SRSDG 408 |  | DG | Incoming Water Temperature Low |  | Incoming water temperature is low.\n- Treatment is currently paused.\n- Check the incoming water supply.\n- When the temperature is within target range\nResume/OK button will be available.\n- Press Resume/OK to continue treatment. | Dial. Gen: Inlet Water Temp Low |  | Inlet water temperature low | Inlet temperature < 24 °C for 5 seconds. | 158 | ALARM_ID_DG_INLET_WATER_TEMPERATURE_IN_LOW_RANGE |  | ALARM_PRIORITY_MEDIUM | 433 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water temperature is in the low range
cleaning |  |  |  | SRSDG 1044 |  | DG | Remove Chemical Disinfectant |  | The chemical disinfectant is no longer needed. \n- Disconnect chemical disinfectant from the device. \n- Place the concentrate connection cap on the device.\n- Keep dialysate cap on the device. | Cleaning: Chem Disinf Connect Cap |  | Close concentrate line be able to flush it | Start of chemical disinfect flush. | 159 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_REMOVE_ACID |  | ALARM_PRIORITY_LOW | 751 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG chemical disinfect flush remove acid and close the concentrate cap.
x | PRS 494 | SA 239 | SRSHD 972 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: BP Rotor Max Speed |  | Blood pump rotor speed too high | If blood pump motor speed exceeded 100.0 RPM. | 160 | ALARM_ID_HD_BLOOD_PUMP_ROTOR_SPEED_TOO_HIGH |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD blood pump rotor speed too high
x | PRS 863 | SA 497 |  | SRSDG 122 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: HD- DG Comm |  | Too many bad CRCs detected on received system messages | If >= 5 bad CRCs detected on received CAN messages within 10 minutes. | 161 | ALARM_ID_DG_COMM_TOO_MANY_BAD_CRCS |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG comm too many bad CRCs
x | PRS 863 | SA 300 |  | SRSDG 867 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Proc- FPGA Clock |  | FPGA clock speed failure | If the FPGA clock speed is relatively close to the processor clock for a certain period of time. | 162 | ALARM_ID_DG_FPGA_CLOCK_SPEED_CHECK_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG FPGA clock speed check failure.
revisit - determine if this is blood stopped or dialysate may need to add to PRS |  |  | SRSHD 1260 |  |  | HD | Ultrafiltration Error |  | A problem was detected with the dialysate device. \n- Confirm the device is level with minimal movement. \n- Resume/OK treatment. \n- If problem persists, call service to report the issue and\nschedule a repair. | Dial. Gen: UF Rate Check |  | HD sees primary load cell for reservoir 1 change too much too fast | If the delta volume of the active reservoir exceeded 50.0 g (MAX_ACTIVE_LOAD_CELL_CHANGE_G). | 163 | ALARM_ID_HD_LOAD_CELL_ACCELERATION_RES_1_ALARM |  | ALARM_PRIORITY_MEDIUM | 442 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD sees primary load cell for reservoir 1 change too much too fast
revisit - determine if this is blood stopped or dialysate may need to add to PRS |  |  | SRSHD 1601 |  |  | HD | Ultrafiltration Error |  | A problem was detected with the dialysate device. \n- Confirm the device is level with minimal movement. \n- Resume/OK treatment. \n- If problem persists, call service to report the issue and\nschedule a repair. | Dial. Gen: UF Rate Check |  | HD sees primary load cell for reservoir 2 change too much too fast | If the delta volume of the active reservoir exceeded 50.0 g (MAX_ACTIVE_LOAD_CELL_CHANGE_G). | 164 | ALARM_ID_HD_LOAD_CELL_ACCELERATION_RES_2_ALARM |  | ALARM_PRIORITY_MEDIUM | 442 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD sees primary load cell for reservoir 2 change too much too fast
sean remove from code 15min limint (SRS 1556 has been updated) |  |  | SRSHD 1228 SRSHD 1399 SRSHD 1556 SRSHD 1709 |  |  | HD | Blood Pump Stop Time Exceeded |  | The blood pump has been off for too long.\n- Treatment must be terminated without rinseback. | Blood Stop: Blood Clot Risk |  | HD in treatment rinseback operation timeout | If the maximum time inrinseback operation has been exceeded. | 165 | ALARM_ID_HD_TREATMENT_RINSEBACK_TIMEOUT_ALARM |  | ALARM_PRIORITY_HIGH | 11 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD in treatment rinseback operation timeout
[7 empty] | Test Alarm - Alarm 166 | [6 empty] | 166 | ALARM_ID_TEST_ALARM_NO_BLOOD_RECIRC_CLR_TOP_ONLY |  | ALARM_PRIORITY_HIGH | 200 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | Test alarm - No blood recirculation, clear top only
[6 empty] | HD | Blood Pump Stopped: Door Open |  | The device door is open.\n Close the device door. \n- Press Resume/OK to continue. | Blood Stop: Door Open |  | HD cartridge door opened alarm | If the cartridge door is opened as reported by the FPGA. | 167 | ALARM_ID_HD_CARTRIDGE_DOOR_OPENED |  | ALARM_PRIORITY_HIGH | 210 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD cartridge door opened alarm
[7 empty] | Test Alarm - Alarm 168 | [6 empty] | 168 | ALARM_ID_TEST_ALARM_NO_DIAL_RECIRC_CLR_TOP_ONLY |  | ALARM_PRIORITY_HIGH | 110 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | Test alarm - No dialysate recirculation, clear top only
| PRS 863 | SA 493 |  | SRSDG 643 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Dialysate Flow Sens Rg |  | Dialysate flow rate out of range | If the dialysate flow rate > 2.00 L/min. | 169 | ALARM_ID_DG_DIALYSATE_FLOW_RATE_OUT_OF_MAX_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG Dialysate flow rate out of maximum range
[6 empty] | HD | Heparin Syringe Empty |  | Heparin syringe is empty.\n- Treatment will continue without heparin.\n- Refer to facility policy if additional heparin is needed to\ncomplete treatment. | Heparin Syringe Empty |  | HD syringe empty alarm | If the syringe pump position is near empty position. | 170 | ALARM_ID_HD_SYRINGE_PUMP_SYRINGE_EMPTY |  | ALARM_PRIORITY_LOW | 730 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | HD syringe empty alarm
|  |  | SRSHD 1525 |  |  | HD | Heparin Pump Occlusion |  | A problem has been detected with the heparin pump.\n- Ensure heparin line is not kinked or clamped. | Process: Pre- Tx Hep. Pump Occ. |  | HD syringe pump occlusion alarm | If the pressure sensor voltage reading > SYRINGE_FORCE_OCCLUSION_THRESHOLD_V. | 171 | ALARM_ID_HD_SYRINGE_PUMP_OCCLUSION |  | ALARM_PRIORITY_LOW | 716 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD syringe pump occlusion alarm
|  |  | SRSHD 1383 |  |  | HD | Not Enough Heparin |  | There is not enough heparin available for treatment.\n- Ensure the syringe pump is touching the syringe plunger.\n- If the syringe pump is touching the syringe plunger,\nend treatment.\n- If syringe pump is not touching the syringe plunger, \nensure the correct amount of heparin has been drawn and\nreload the syringe. | Process: Pre- Tx Not Enough Hep. |  | HD syringe pump not enough Heparin alarm | If the estimated syringe volume needed for treatment is not sufficient | 172 | ALARM_ID_HD_SYRINGE_PUMP_NOT_ENOUGH_HEPARIN_ALARM |  | ALARM_PRIORITY_LOW | 715 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD syringe pump not enough Heparin alarm
| PRS 494 | SA 474 | SRSHD 659 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: RTC Config |  | HD RTC configuration error | If the RTC configuration registers report an error. | 173 | ALARM_ID_HD_RTC_CONFIG_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD RTC configuration error
| PRS 398 | SA 198 | SRSHD 630 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: HD RTC Timer |  | HD RTC vs. Timer accuracy error | If RTC and timer are not aligned on 1 second elapsed time. | 174 | ALARM_ID_HD_RTC_OR_TIMER_ACCURACY_FAILURE |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD RTC or timer accuracy failure
| PRS 494 | SA 245 SA 254 SA 262 | SRSHD 1502 SRSHD 1503 SRSHD 1568 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Pump Commutation Err |  | HD pump commutation error | If the set pump direction does not agree with the hall sensor direction. | 175 | ALARM_ID_HD_PUMP_DIRECTION_STATUS_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD pump direction status error
[7 empty] | Test Alarm - Alarm 176 | [6 empty] | 176 | ALARM_ID_TEST_ALARM_NO_RECIRC_CLR_TOP_ONLY |  | ALARM_PRIORITY_HIGH | 200 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | Test alarm - No blood or dialysate recirculation, clear top only
| PRS 863 | SA 417 |  | SRSDG 919 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Software Invalid State |  | DG software fault. Software found itself in an unexpected state | If the specific software faults listed in alarm management SW faults has been detected | 177 | ALARM_ID_DG_SOFTWARE_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG software fault. Software found itself in an unexpected state
| PRS 494 | SA 225 |  |  | SRSUI 1066 | UI | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: UI- Proc Comm |  | HD communication timeout | If the HD stops communicating with the UI. | 178 | ALARM_ID_HD_COMM_TIMEOUT | OK | ALARM_PRIORITY_HIGH | 0 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD communication timeout
|  |  |  | SRSDG 605 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: FPGA Com |  | DG FPGA communication down for too long | If retries for commands exceeds limit or FPGA reports comm error. | 179 | ALARM_ID_DG_FPGA_COMM_TIMEOUT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG FPGA communication down for too long
| PRS 863 | SA 494 |  | SRSDG 883 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: RO Flow Sensor Range |  | RO flow rate out of range | If the RO flow rate > 2.00 L/min. | 180 | ALARM_ID_DG_RO_FLOW_RATE_OUT_OF_MAX_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO flow out of maximum range
|  |  |  | SRSDG 849 SRSDG 838 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Tare Out of Range |  | DG load cells weight out of range for tare | If the load cell weight is out of range for tare. | 181 | ALARM_ID_DG_LOAD_CELLS_TARE_WEIGHT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells weight out of range for tare
| PRS 863 | SA 369 |  | SRSDG 162 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: LC Cal |  | DG load cells invalid calibration | If the CRC of load cells calibration data does not match the calculated CRC. | 182 | ALARM_ID_DG_LOAD_CELLS_INVALID_CAL_RECORD |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells invalid calibration
|  |  |  | SRSDG 851 |  | DG | Disinfection Interrupted |  | Disinfection terminated unexpectedly.\n- Restart disinfection.\n- If problem persists, call service to schedule a repair. | Cleaning: Invalid Resv Transfer |  | DG invalid load cell value | If the half full reservoir is filled prior to filling the reservoir that is supposed to overflow to the half full reservoir. | 183 | ALARM_ID_DG_INVALID_LOAD_CELL_VALUE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG invalid load cell value
| PRS 863 | SA 355 SA 356 |  | SRSDG 559 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: UVi Error |  | Inlet reactor not healthy | If the inlet reactor are not healthy. | 184 | ALARM_ID_DG_INLET_UV_REACTOR_NOT_HEALTHY |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG inlet UV reactor not healthy
| PRS 874 | SA 299 |  | SRSDG 628 |  | DG | Service Required: Dialysate Device |  | A non- critical problem was detected with the dialysate device.\n- Call service to report the issue and schedule a repair. | DG Alarm (Low): Fan Speed |  | DG fan RPM out of range | If the fan RPM is out of Min/Max range. | 185 | ALARM_ID_DG_FAN_RPM_OUT_OF_RANGE |  | ALARM_PRIORITY_LOW | 851 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG fan RPM out of range
| PRS 863 | SA 351 SA 353 |  | SRSDG 891 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: CP1/CP2 Fault |  | DG concentrate pump fault | If the concentrate pumps fault is reported by the FPGA. | 186 | ALARM_ID_DG_CONCENTRATE_PUMP_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG concentrate pump fault
| PRS 863 | SA 352 |  | SRSDG 905 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: CP1 Speed Fault |  | DG concentrate pump CP1 speed control error | If the concentrate pump CP1 speed > 0.02 tolerance for a certain period of time. | 187 | ALARM_ID_DG_CP1_SPEED_CONTROL_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG concentrate pump CP1 speed control error
| PRS 863 | SA 354 |  | SRSDG 906 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: CP2 Speed Fault |  | DG concentrate pump CP2 speed control error | If the concentrate pump CP1 speed > 0.02 tolerance for a certain period of time. | 188 | ALARM_ID_DG_CP2_SPEED_CONTROL_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG concentrate pump CP2 speed control error
| PRS 863 | SA 339 SA 343 |  | SRSDG 448 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: DRP Speed Fault |  | DG drain pump RPM out of range | If the drain pump RPM > 100.0 of the target RPM. | 189 | ALARM_ID_DG_DRAIN_PUMP_RPM_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain pump RPM out of range
[6 empty] | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: DRP Speed Not Off |  | DG drain pump off fault | If the pump is in off state and the RPM is greater than the minimum RPM of 300.0 RPM. | 190 | ALARM_ID_DG_DRAIN_PUMP_OFF_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain pump off fault
| PRS 863 | SA 498 |  | SRSDG 885 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: RO Flow High |  | DG flow rate out of upper range | If the RO flow rate > 10% from the target flow rate. | 191 | ALARM_ID_DG_FLOW_RATE_OUT_OF_UPPER_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG flow rate out of upper range
| PRS 863 | SA 498 |  | SRSDG 945 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: RO Flow Low |  | DG flow rate out of lower range | If the RO flow rate < 10% from the target flow rate. | 192 | ALARM_ID_DG_FLOW_RATE_OUT_OF_LOWER_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG flow rate out of lower range
[14 empty] | 193 | ALARM_ID_AVAILABLE_5 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD dialysate temperature out of low safety range.
| PRS 863 | SA 360 |  | SRSDG 881 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: RO Driver Feedback |  | DG RO pump duty cycle out of range | If the DG RO pump duty cycle is out of range by 5% of the target duty cycle for a certain period of time. | 194 | ALARM_ID_DG_RO_PUMP_DUTY_CYCLE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO pump duty cycle out of range
| PRS 863 | SA 487 |  | SRSDG 925 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: RO Pump High Pressure |  | DG RO pump pressure out of range | If the DG RO pump pressure > 135 psi for a certain period of time. | 195 | ALARM_ID_DG_RO_PUMP_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO pump pressure out of range
| PRS 863 | SA 348 |  | SRSDG 955 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: CPi/CPo Comm |  | DG CPi/CPo conductivity sensors FPGA fault | If the FPGA reports error for a certain period of time. | 196 | ALARM_ID_DG_CPI_CPO_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CPi/CPo sensors FPGA fault
| PRS 863 | SA 347 |  | SRSDG 956 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: CD1/CD2 Comm |  | DG CD1/CD2 conductivity sensors FPGA fault | If the FPGA reports error for a certain period of time. | 197 | ALARM_ID_DG_CD1_CD2_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CD1/CD2 sensors FPGA fault
| PRS 863 | SA 367 |  | SRSDG 872 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: HP On With No Flow |  | DG primary heater while RO flow is low | If the RO flow is < 0.2 L/min for a certain period of time. | 198 | ALARM_ID_DG_RO_FLOW_TOO_LOW_WHILE_PRIMARY_HEATER_IS_ON |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO flow too low while primary heater is on
| PRS 863 | SA 366 |  | SRSDG 871 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: HR On With No Flow |  | DG trimmer heater while dialysate flow is low | If the FMD flow is ≤ 0 mL/min (raw) for a certain period of time. | 199 | ALARM_ID_DG_DIALYSATE_FLOW_TOO_LOW_WHILE_TRIMMER_HEATER_IS_ON |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG Dialysate flow too low while trimmer heater is on
| PRS 863 | SA 293 SA 358 SA 359 |  | SRSDG 178 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Thermistor Range |  | DG thermistor or temperature sensor out of range | If thermistor > 80.0 C or < 0.0 C. | 200 | ALARM_ID_DG_THERMISTORS_TEMPERATURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG thermistors/sensors temperature out of range
|  |  | SRSHD 746 SRSHD 747 |  |  | HD | Self Test: Dialysate Flow Rate |  | A problem has been detected when measuring dialysate flow.\n- Confirm cartridge is properly installed and primed.\n- If problem persists, terminate treatment preparation and \n start over using a new cartridge. | Process: Pre- Tx Wet Self- Test |  | HD pre-treatment wet flow test failure | If the pre-treatment wet flow test has failed. | 201 | ALARM_ID_HD_PRE_TREATMENT_WET_FLOW_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode wet self-test lc vs FMD failure
[6 empty] | HD | Self Test: Pressure |  | A problem has been detected during self- test. \n- Confirm the cartridge is properly installed. \n- Ensure arterial and venous pressure sensor lines are\nnot kinked.\n- Press Resume/OK to resume. \n- If problem persists, install a new cartridge. | Process: Pre- Tx Dry Self- Test |  | HD pre-treatment dry pressure normal test failure | If the pre-treatment dry pressure normal test failed. | 202 | ALARM_ID_HD_PRE_TREATMENT_DRY_PRESSURE_NORMAL_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode dry pressure normal self-test failure
[14 empty] | 203 | ALARM_ID_AVAILABLE_6 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Available for use
| PRS 863 | SA 373 |  | SRSDG 429 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: RO Rejection |  | RO rejection ratio out of range | If RO rejection ratio > 0.1 | 204 | ALARM_ID_DG_RO_REJECTION_RATIO_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG RO rejection ratio out of range
|  |  |  | SRSDG 913 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Cond Sensors Error |  | Conductivity sensor fault | If conductivity sensor reports 0x11 for more than 2 consecutive reads. | 205 | ALARM_ID_DG_CONDUCTIVITY_SENSOR_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG conductivity sensor fault
|  |  |  | SRSDG 437 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Reservoir Fill Timeout |  | DG dialysate fill runs out of time | If DG dialysate fill > 5 minutes. | 206 | ALARM_ID_DG_DIALYSATE_FILL_OUT_OF_TIME |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG dialysate fill runs out of time
|  |  |  | SRSDG 240 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: FMP Flow Rate Error |  | DG flow meter check failure alarm | If integrated Volume To Load Cell Reading Percent > 0.1. | 207 | ALARM_ID_DG_FLOW_METER_CHECK_FAILURE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG flow meter check failure alarm
|  |  | SRSHD 1569 |  |  | HD | High Venous Pressure |  | Venous pressure is high. \n- Check venous bloodlines for kinks or clamps.\n- Check needle placement and/or access position.\n- Resolve access issues, if applicable.\n- If clots are present, end treatment without rinseback. | Blood Stop: Venous Pres. High |  | HD venous line occlusion fault | If the venous pressure is > set limit + 50 mmHg if HD is in treatment or pre-treatment mode. | 208 | ALARM_ID_HD_VENOUS_LINE_OCCLUSION |  | ALARM_PRIORITY_HIGH | 201 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | HD venous pressure occlusion alarm
|  |  |  | SRSDG 980 |  | DG | Disinfection Interrupted |  | Disinfection terminated unexpectedly.\n- Restart disinfection.\n- If problem persists, call service to schedule a repair. | Cleaning: Flush Circ Error |  | DG drain circulation line timeout | If the integrated volume < 0.1 for more than 60 secs. | 209 | ALARM_ID_DG_DRAIN_CIRCULATION_LINE_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG drain circulation line timeout
| PRS 872 | SA 229 | SRSHD 1397 |  |  | HD | Service Required: Hemodialysis Device |  | A non- critical problem was detected with the hemodialysis\n device.\n- Call service to report the issue and schedule a repair. | HD Alarm (Low): Battery Fault |  | HD battery pack detected an error | If the battery pack status error is detected from the slave device I2C address 0x0B. | 210 | ALARM_ID_HD_BATTERY_PACK_ERROR_DETECTED |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | HD battery pack detected an error
| PRS 389 |  | SRSHD 946 |  |  | HD | Approaching MAX Blood Pump Stop Time |  | The blood has been stopped for nearly 5 minutes.\n- You must resume treatment before the blood pump stop time\ntimer expires.\n- If treatment is not resumed before the timer expires,\ntreatment will be terminated without rinseback. | Blood Stopped: Warning |  | Blood sitting too long warning (> 4 min) | During treatment stop sub-mode, if the blood is in the line for more than 4 minutes. | 211 | ALARM_ID_HD_BLOOD_SITTING_WARNING |  | ALARM_PRIORITY_HIGH | 290 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD blood sitting too long warning (>4 min)
| PRS 389 |  | SRSHD 952 |  |  | HD | Approaching MAX Blood Pump Stop Time |  | The blood has been stopped for nearly 5 minutes.\n- You must Resume/OK treatment before the blood pump stop \ntime timer expires.\n- If treatment is not Resume/OK before the timer expires,\ntreatment will be terminated without rinseback. | Blood Stopped: Warning |  | Blood sitting too long warning (> 4 min) | During treatment end paused state, if the blood is in the line for more than 4 minutes. | 212 | ALARM_ID_HD_END_OF_TREATMENT_ALARM |  | ALARM_PRIORITY_HIGH | 290 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD end of treatment alarm (user not acting to end treatment)
| PRS 530 |  | SRSHD 1120 |  |  | HD | Prime Complete - 30 MIN WARNING |  | The priming process completed 90 minutes ago. \n- Treatment must begin within 30 minutes or the\npre- treatment process will be terminated.\n- Press Resume/OK to continue. | Process: Pre- Tx Prime Complete |  | HD prime completed medium priority alarm | In Pre-Treatment Recirculation for 90 minutes (escalation of alarm 226) | 213 | ALARM_ID_HD_PRIME_COMPLETED_MEDIUM |  | ALARM_PRIORITY_MEDIUM | 590 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | HD prime completed medium priority alarm
[14 empty] | 214 | ALARM_ID_AVAILABLE_30 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | Available for use
[14 empty] | 215 | ALARM_ID_AVAILABLE_31 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | Available for use
| PRS 393 |  | SRSHD 1405 |  |  | HD | Treatment Complete |  | Dialysis treatment is complete.\n- Press RINSEBACK to end treatment and rinseback. | Process: End of Dialysis Tx |  | HD end treatment sub-mode timeout alarm | If the treatment ended successfully. | 216 | ALARM_ID_HD_END_TREATMENT_TIMEOUT_ALARM |  | ALARM_PRIORITY_HIGH | 291 | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD end treatment sub-mode timeout alarm
[14 empty] | 217 | ALARM_ID_AVAILABLE_32 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | Available for use
|  |  | SRSHD 1294 |  |  | HD | Syringe Detected |  | Heparin syringe was not removed from device.\n- Remove heparin syringe.\n- Press Resume/OK to continue. | Process: Remove Syringe |  | HD syringe detected alarm | If the syringe detection switch > 2.0 volts. | 218 | ALARM_ID_HD_SYRINGE_DETECTED |  | ALARM_PRIORITY_LOW | 745 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD syringe detected alarm
[6 empty] | HD | Blood Stopped: Syringe Not Detected |  | There is a problem detecting a syringe in the heparin pump.\n- Confirm the heparin syringe is properly installed. \n- Press Resume/OK to resume treatment. | Blood Stop: Syringe Not Detected |  | HD syringe pump syringe removed alarm | If the syringe detection switch < 2.0 volts. | 219 | ALARM_ID_HD_SYRINGE_PUMP_SYRINGE_REMOVED |  | ALARM_PRIORITY_HIGH | 207 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | HD syringe pump syringe removed alarm
| PRS 863 | SA 334 SA 335 |  | SRSDG 947 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: THD RTD Comm |  | DG THd sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 220 | ALARM_ID_DG_THD_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG THd sensors FPGA fault
|  |  | SRSHD 1306 SRSHD 1307 SRSHD 1308 |  |  | HD | Blood Pump Stopped: Saline Bag Empty |  | The saline bag appears empty.\n- Check the saline bag and replace if empty. | Blood Stop: Saline Bag Empty |  | HD Saline bag is empty | If the measured arterial pressure < -300.0 mmHg for more than 250 msec. | 221 | ALARM_ID_HD_EMPTY_SALINE_BAG |  | ALARM_PRIORITY_HIGH | 204 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | HD Saline bag is empty
| PRS 494 | SA 230 | SRSHD 1585 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: BP Occ Com |  | HD occlusion sensor FPGA fault | If the FPGA does not report fresh data or indicates an error for a certain period of time. | 222 | ALARM_ID_HD_OCCLUSION_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD occlusion sensor FPGA fault
| PRS 494 | SA 236 | SRSHD 1583 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Art. Pres. Com |  | HD arterial sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 223 | ALARM_ID_HD_ARTERIAL_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD arterial sensor FPGA fault
| PRS 388 |  | SRSHD 944 |  |  | HD | Blood Pump Stopped: Treatment Paused |  | Treatment has been stopped by user.\n- Press Resume/OK to resume treatment. \n- If treatment is not resumed within five minutes,\ntreatment will be terminated without rinseback. | Blood Stopped: Treatment Pause |  | Treatment stopped by user action - pressed stop button | If the user has pressed the stop button. | 224 | ALARM_ID_HD_TREATMENT_STOPPED_BY_USER |  | ALARM_PRIORITY_MEDIUM | 300 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD treatment stopped by user action - pressed stop button
| PRS 391 |  | SRSHD 810 |  |  | HD | Treatment Complete |  | Dialysis treatment is complete.\n- If blood sampling is required, draw blood at\n the arterial sample port\n- Press RINSEBACK to end treatment and rinseback. | Process: End of Treatment |  | HD end of treatment warning | If the treatment duration has exceeded the prescribed treatment time entered by the user. | 225 | ALARM_ID_HD_END_OF_TREATMENT_WARNING |  | ALARM_PRIORITY_LOW | 739 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD end of treatment warning
| PRS 531 |  | SRSHD 1118 |  |  | HD | Priming Complete |  | Priming is complete. \n- Treatment must begin within 120 minutes \nor the pre- treatment process will be terminated.\n- Press Resume/OK to continue. | Process: Pre- Tx Priming Complete |  | HD prime completed low priority alarm | Start of Pre-Treatment Recirculation - notifiy user to proceed to patient connection | 226 | ALARM_ID_HD_PRIME_COMPLETED_LOW_PRIORITY |  | ALARM_PRIORITY_LOW | 729 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | HD prime completed low priority alarm
[14 empty] | 227 | ALARM_ID_AVAILABLE_15 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | Available for use
|  |  | SRSHD 1571 |  |  | HD | Priming Error |  | A problem has been detected during prime. \n- Prime was unable to complete. \n - Check blood and dialysate lines for clamps, kinks or leaks. | Process: Pre- Tx Prime Timeout |  | HD time out on prime sub-mode purge air state. | During the pre-treatment prime state, if the prime saline purge air state exceeded 60 secs. | 228 | ALARM_ID_HD_PRIME_SALINE_PURGE_AIR_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD time out on prime saline purge air state.
|  |  | SRSHD 1189 |  |  | HD | Priming Error |  | A problem has been detected during prime. \n- Prime was unable to complete. \n - Check blood and dialysate lines for clamps, kinks or leaks. | Process: Pre- Tx Prime Timeout |  | HD prime dialysate dialyzer time out alarm | If the priming of the dialyzer dialysate flow path has exceeded the time limit. | 229 | ALARM_ID_HD_PRIME_DIALYSATE_DIALYZER_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD prime dialysate dialyzer time out alarm
|  |  | SRSHD 1190 |  |  | HD | Priming Error |  | A problem has been detected during prime. \n- Prime was unable to complete. \n - Check blood and dialysate lines for clamps, kinks or leaks. | Process: Pre- Tx Prime Timeout |  | HD prime dialysate bypass time out alarm | If the priming of the dialyzer bypass flow path has exceeded the time limit. | 230 | ALARM_ID_HD_PRIME_DIALYSATE_BYPASS_TIME_OUT |  | ALARM_PRIORITY_LOW | 720 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD prime dialysate bypass time out alarm
|  |  | SRSHD 771 SRSHD 1599 SRSHD 1579 SRSHD 1580 SRSHD 772 |  |  | HD | Self Test: Cartridge Check |  | A problem has been detected during self- test. \n- Confirm cartridge is properly installed without kinks or clamps.\n- Ensure all connections are secure. \n- Press Resume/OK to resume. \n- If problem persists, install a new cartridge. | Process: Pre- Tx Dry Self- Test |  | HD pre-treatment mode dry self-tests failure | If the following dry self tests have failed: 1. handleDrySelfTestPressureSensorsState 2. handleDrySelfTestPressureSensorsNormalState | 231 | ALARM_ID_HD_PRE_TREATMENT_DRY_PRESSURE_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode dry pressure self-test failure
|  |  | SRSHD 1534 SRSHD 1535 |  |  | HD | Self Test: Dialysate Flow Rate |  | A problem has been detected when measuring dialysate flow.\n- Confirm cartridge is properly installed and primed.\n- If problem persists, terminate treatment preparation and \n start over using a new cartridge. | Process: Pre- Tx Wet Self- Test |  | HD pre-treatment mode wet self-tests failure | If the following wet self tests have failed: 1. handleWetSelfTestPrimeCheckState 2. handleWetSelfTestFirstDisplacementVerifyState 3. handleWetSelfTestSecondDisplacementVerifyState | 232 | ALARM_ID_HD_PRE_TREATMENT_WET_LC_TEST_FAILURE |  | ALARM_PRIORITY_LOW | 725 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD pre-treatment mode wet self-test LC vs. LC failure
| PRS 863 | SA 461 |  | SRSDG 964 SRSDG 949 SRSDG 1089 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Baro Cal CRC |  | DG barometric pressure sensor bad coefficients CRC | If the following happens: 1. The stored CRC in the sensor does not match the calculated CRC. 2. The CRC and coefficients have not been received from the sensor after 20 seconds from the power up. | 233 | ALARM_ID_DG_BAROMETRIC_SENSOR_COEFFS_BAD_CRC |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG barometric sensor coefficients bad CRC
| PRS 841 |  |  | SRSDG 1037 |  | DG | Incoming Water Pressure High |  | Incoming water pressure is high.\n- Treatment is currently paused.\n- Check the incoming water supply.\n- When the pressure is within target range\nResume/OK button will be available.\n- Press Resume/OK to continue treatment. | Dial. Gen: Inlet Water Pres High |  | DG inlet water pressure is too high | If the inlet water pressure is > 80 psig. | 234 | ALARM_ID_DG_INLET_WATER_PRESSURE_IN_HIGH_RANGE |  | ALARM_PRIORITY_MEDIUM | 430 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | DG inlet water pressure in high range.
[6 empty] | HD | Rinseback Complete |  | Rinseback is complete. \n- Follow facililty policy prior to disconnecting patient. \n- Press Resume/OK to continue. | Process: Rinseback Complete |  | HD in treatment stopped sub-mode after rinseback completed (no escalation) | If there is no active alarm and rinseback has completed. | 235 | ALARM_ID_HD_TREATMENT_STOPPED_AFTER_RINSEBACK |  | ALARM_PRIORITY_LOW | 738 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | HD in treatment stopped sub-mode after rinseback completed (no escalation)
|  |  | SRSHD 1205 |  |  | HD | Self Test: Fluid Present |  | A problem has been detected during self- test. \n- Cartridges are single- use only and fluid is present.\n- Install new cartridge to resume. | Process: Pre- Tx Dry Self- Test |  | HD needs new cartridge to be installed | If the air trap sensor reported AIR and bubble sensor detected bubbles. | 236 | ALARM_ID_HD_INSTALL_NEW_CARTRIDGE |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD needs new cartridge to be installed
[14 empty] | 237 | ALARM_ID_AVAILABLE_13 |  | ALARM_PRIORITY_LOW | 999 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | Available for use
|  |  | SRSHD 1359 |  |  | HD | Cartridge Not Detected |  | A problem has been detected during self- test. \n- Cartridge is not detected.\n- Confirm cartridge is properly installed. | Process: Pre- Tx Dry Self- Test |  | HD no cartridge loaded or installed improperly alarm | If the measured blood pump occlucsion pressure reading is < 5000 or > 24500. | 238 | ALARM_ID_HD_NO_CARTRIDGE_LOADED |  | ALARM_PRIORITY_LOW | 710 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD no cartridge loaded or installed improperly alarm
|  |  | SRSHD 1542 |  |  | HD | Cartridge Still Detected |  | A problem has been detected during self- test. \n- Cartridge is still present or was unloaded incorrectly.\n- Confirm the cartridge has been fully removed. | Process: Post- Tx Remove Cartridge |  | HD fail to remove cartridge alarm | In the post treatment state, if the cartridge has not been removed. | 239 | ALARM_ID_HD_CARTRIDGE_REMOVAL_FAILURE |  | ALARM_PRIORITY_LOW | 746 | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD fail to remove cartridge alarm
| PRS 542 |  |  | SRSDG 482 |  | DG | Conductivity Error During Prime |  | Conductivity out of range during prime.\n- Confirm that acid and bicarbonate concentrates\nare securely connected.\n- If empty, replace. \n- Treatment will need to be resumed when conductivity\nis within target range. | Dial. Gen: Bicarb Conc. Cond. |  | Bicarb conductivity out of range during bicarb pump check alarm | If the conductivity value for bicarbonate concentrate is out of range for more than 45 secs. | 240 | ALARM_ID_DG_BICARB_CONDUCTIVITY_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 411 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | DG bicarb conductivity out of range during bicarb pump check alarm
| PRS 873 |  |  | SRSDG 848 SRSDG 90 SRSDG 561 SRSDG 577 SRSDG 590 SRSDG 591 SRSDG 336 SRSDG 879 SRSDG 340 SRSDG 496 SRSDG 501 SRSDG 505 SRSDG 532 SRSDG 535 SRSDG 556 SRSDG 557 SRSDG 1048 SRSDG 1049 SRSDG 651 SRSDG 655 SRSDG 744 SRSDG 745 SRSDG 1004 SRSDG 1008 SRSDG 1019 SRSDG 1023 SRSDG 1033 SRSDG 1034 |  | DG | Disinfection Interrupted |  | Disinfection terminated unexpectedly.\n- Restart disinfection.\n- If problem persists, call service to schedule a repair. | Cleaning: Resv Drain Timeout |  | DG reservoir drain time out | If the DG reservoir drain is not complete in 120 secs. | 241 | ALARM_ID_DG_RESERVOIR_DRAIN_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG reservoir drain time out
| PRS 873 |  |  | SRSDG 850 SRSDG 573 SRSDG 574 SRSDG 579 SRSDG 580 SRSDG 492 SRSDG 493 SRSDG 498 SRSDG 499 SRSDG 508 SRSDG 509 SRSDG 519 SRSDG 690 SRSDG 691 SRSDG 703 SRSDG 711 SRSDG 1020 SRSDG 1024 |  | DG | Disinfection Interrupted |  | Disinfection terminated unexpectedly.\n- Restart disinfection.\n- If problem persists, call service to schedule a repair. | Cleaning: Resv Fill Timeout |  | DG reservoir fill time out | If the DG reservoir fill is not complete in 120 secs. | 242 | ALARM_ID_DG_RESERVOIR_FILL_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG reservoir fill time out
|  |  |  | SRSDG 853 |  | DG | Disinfection Interrupted |  | Disinfection terminated unexpectedly.\n- Restart disinfection.\n- If problem persists, call service to schedule a repair. | Cleaning: Fluid Loss Detected |  | DG reservoir leak time out | If the reservoir level is drifted by 100 mL for more than 30 seconds. | 243 | ALARM_ID_DG_RESERVOIR_LEAK_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG reservoir leak time out
|  |  |  | SRSDG 852 |  | DG | Disinfection Interrupted |  | Disinfection terminated unexpectedly.\n- Restart disinfection.\n- If problem persists, call service to schedule a repair. | Cleaning: Temp Sensor Check |  | DG temperature sensors difference out of range | If any of the temperature sensors are out of the average of the sensors by more than 3 C. | 244 | ALARM_ID_DG_CLEANING_MODE_TEMP_SENSORS_DIFF_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode temperature sensors difference out of range
|  |  |  | SRSDG 854 |  | DG | Disinfection Interrupted |  | Disinfection terminated unexpectedly.\n- Restart disinfection.\n- If problem persists, call service to schedule a repair. | Cleaning: Heat Disinfect Temp TO |  | DG heat disinfect target temperature time out (could not reach to temperature) | If heating up to minimum temperature for heat disinfect failed after 4 hours. | 245 | ALARM_ID_DG_HEAT_DISINFECT_TARGET_TEMP_TIMEOUT |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG heat disinfect target temperature time out (could not reach to temperature)
[6 empty] | DG | Disinfection Interrupted |  | Disinfection terminated unexpectedly.\n- Restart disinfection.\n- If problem persists, call service to schedule a repair. | Cleaning: RO Water Cond Error |  | DG chemical disinfect inlet conducitivty out of range | If CPo and CD2 conductivity values are above 100 uS/cm. | 246 | ALARM_ID_DG_CLEANING_MODE_COND_SENSORS_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode conductivity sensors out of range
[14 empty] | 247 | ALARM_ID_AVAILABLE_7 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Alarm Id available
[6 empty] | DG | Disinfection Interrupted |  | Disinfection terminated unexpectedly.\n- Restart disinfection.\n- If problem persists, call service to schedule a repair. | Cleaning: Chem Disinfect Proc |  | DG chemical disinfect out of target temperature | If TPo > 50 C after 1 second. | 248 | ALARM_ID_DG_CHEM_DISINFECT_TARGET_TEMP_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG chemical disinfect could not reach to target temperature out of range
|  |  |  | SRSDG 992 |  | DG | Chemical Disinfectant Not Detected |  | Unable to sense proper chemical disinfectant.\n- Disinfection process has terminated.\n- Confirm chemical disinfectant is connected to the\nconcentrate inlet.\n- Confirm chemical disinfectant container is not empty.\n- Restart disinfection.\n- If problem persists, call service to schedule a repair. | Cleaning: Chem Disinf PAA Cond |  | DG chemical disinfect out of target conductivity | If CD2 < 150 uS/cm or CD2 > 650 uS/cm after 10 seconds. | 249 | ALARM_ID_DG_CHEM_DISINFECT_TARGET_COND_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG chemical disinfect target conductivity out of range
|  |  |  | SRSDG 1046 |  | DG | Connect Chemical Disinfectant |  | Connect chemical disinfectant to the device. \n- Remove concentrate cap.\n- Keep dialysate cap on the device.\n- Place disinfectant straw in the disinfectant bottle. \n- Attach straw to concentrate connection and ensure\nconnection is secure.\n- Press Resume/OK to continue. | Cleaning: Chem Disinf Connect PAA |  | Insert the acid to start chemical disinfect | Start of chemical disinfect | 250 | ALARM_ID_DG_CHEM_DISINFECT_INSERT_ACID |  | ALARM_PRIORITY_LOW | 750 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG chemical disinfect insert acid and remove the concentrate cap
| PRS 398 | SA 196 | SRSHD 1058 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: Sys Rec NVRAM CRC |  | HD invalid system record | If the calculated CRC of the HD system record image does not match the expected CRC. | 251 | ALARM_ID_HD_INVALID_SYSTEM_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD invalid system record
| PRS 398 | SA 196 | SRSHD 1062 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: Serv Rec NVRAM CRC |  | HD invalid service record | If the calculated CRC of the HD service record image does not match the expected CRC. | 252 | ALARM_ID_HD_INVALID_SERVICE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD invalid service record
| PRS 863 | SA 283 |  | SRSDG 609 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Sys Rec NVRAM CRC |  | DG invalid system record | If the calculated CRC of the DG system record image does not match the expected CRC. | 253 | ALARM_ID_DG_INVALID_SYSTEM_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid system record
| PRS 863 | SA 283 |  | SRSDG 610 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Serv Rec NVRAM CRC |  | DG invalid service record | If the calculated CRC of the DG service record image does not match the expected CRC. | 254 | ALARM_ID_DG_INVALID_SERVICE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid service record
| PRS 398 | SA 206 | SRSHD 1358 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: UI Compatibility |  | HD and UI software builds are not compatible | If the compatibility revision field of the UI is not equal to SW_COMPATIBILITY_REV. | 255 | ALARM_ID_HD_UI_COMPATIBILITY_ERROR |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD and UI software builds are not compatible
| PRS 398 | SA 205 |  |  | SRSUI 497 | HD | Chemical Flush in Progress |  | Chemical flush in progress.\n- Please wait for the process to complete. \n- You will not be able to create a treatment until the process\nis finished.\n- To end this process, press STOP on the front of the device. | Cleaning: Chem Flush In Progress |  | DG flush started | If DG flush successfully started from the UI. | 256 | ALARM_ID_HD_DISINFECT_CHEM_FLUSH | OK | ALARM_PRIORITY_LOW | 843 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD processor is in Disinfect Chemical Flush mode
[14 empty] | 257 | ALARM_ID_AVAILABLE_8 |  | ALARM_PRIORITY_LOW | 999 | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | Alarm ID available for use
[6 empty] | UI HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: OS Version |  | UI POST OS Version compatibility failure | If the UI POST reports the OS Version compatibility failure. | 258 | ALARM_ID_HD_UI_POST_FAILURE_OS_VERSION |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST OS version compatibility failure
| PRS 494 | SA 220 SA 215 SA 216 SA 235 SA 238 | SRSHD 1390 SRSHD 1388 SRSHD 1387 SRSHD 1497 SRSHD 1499 SRSHD 1682 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Temp Sensors Error |  | HD temperatures out of range | If thermistor reports temperature > 80.0 C or < 0.0 C for a certain period of time. | 259 | ALARM_ID_HD_TEMPERATURES_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD temperatures out of range
| PRS 398 | SA 200 | SRSHD 977 |  | SRSUI 1067 | UI HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: UI Image |  | UI POST Application Integrity (Sha256Sum) failure | If the UI POST reports a shasum failure. | 260 | ALARM_ID_HD_UI_POST_FAILURE_SHASUM | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST Application Integrity (Sha256Sum) failure
| PRS 398 | SA 201 | SRSHD 1655 |  | SRSUI 1068 | UI HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: UI CAN Driver |  | UI POST CANBus failure | If the UI POST reports a can bus failure. | 261 | ALARM_ID_HD_UI_POST_FAILURE_CANBUS | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST CANBus failure
[12 empty] | UI POST Display failure | If the UI POST reports a display failure. | 262 | ALARM_ID_HD_UI_POST_FAILURE_DISPLAY | OK | ALARM_PRIORITY_LOW | 999 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | Alarm Id available
| PRS 398 | SA 201 | SRSHD 1657 |  | SRSUI 1070 | UI HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: UI Touch Driver |  | UI POST Touch failure | If the UI POST reports a touch failure. | 263 | ALARM_ID_HD_UI_POST_FAILURE_TOUCH | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST Touch failure
| PRS 398 | SA 201 | SRSHD 1658 |  | SRSUI 1071 | UI HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: UI SD- Card |  | UI POST SD-Card failure | If the UI POST reports SD card failure. | 264 | ALARM_ID_HD_UI_POST_FAILURE_SDCARD | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST SD-Card failure
| PRS 398 | SA 201 SA 202 SA 203 | SRSHD 1659 |  | SRSUI 1072 | UI HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: UI RTC |  | UI POST RTC failure | If the UI POST reports RTC failure. | 265 | ALARM_ID_HD_UI_POST_FAILURE_RTC | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST RTC failure
| PRS 398 | SA 201 | SRSHD 1660 |  | SRSUI 1073 | UI HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: UI WiFi |  | UI POST WiFi failure | If the UI POST reports WiFi failure. | 266 | ALARM_ID_HD_UI_POST_FAILURE_WIFI | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST WiFi failure
| PRS 398 | SA 201 | SRSHD 1661 |  | SRSUI 1074 | UI HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: UI Bluetooth |  | UI POST Bluetooth failure | If the UI POST reports bluetooth failure. | 267 | ALARM_ID_HD_UI_POST_FAILURE_BLUETOOTH | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST Bluetooth failure
| PRS 398 | SA 201 | SRSHD 1662 |  | SRSUI 1075 | UI HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: UI Ethernet |  | UI POST Ethernet failure | If the UI POST reports ethernet failure. | 268 | ALARM_ID_HD_UI_POST_FAILURE_ETHERNET | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST Ethernet failure
| PRS 398 | SA 201 | SRSHD 1663 |  | SRSUI 1076 | UI HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: UI Sound |  | UI POST Sound failure | If the UI POST reports sound failure. | 269 | ALARM_ID_HD_UI_POST_FAILURE_SOUND | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST Sound failure
| PRS 398 | SA 459 | SRSHD 1481 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: Proc Safestate |  | HD POST Safety Shutdown failure | If the 24 volts power source is greater than 5.0 volts during safety shutdown self test state. If the 24 volts power source is less than 22.6 volts during safety shutdown self test recover state. | 270 | ALARM_ID_HD_SAFETY_SHUTDOWN_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD POST Safety Shutdown failure
| PRS 863 | SA 292 |  | SRSDG 822 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Proc Safestate |  | DG POST Safety Shutdown failure | If the 24 volts power source is greater than 5.0 volts during safety shutdown self test state. If the 24 volts power source is less than 22.6 volts during safety shutdown self test recover state. | 271 | ALARM_ID_DG_SAFETY_SHUTDOWN_POST_TEST_FAILED |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG POST Safety Shutdown failure
| PRS 872 | SA 221 | SRSHD 1394 |  |  | HD | Service Required: Hemodialysis Device |  | A non- critical problem was detected with the hemodialysis\n device.\n- Call service to report the issue and schedule a repair. | HD Alarm (Low): Fan Speed |  | HD Fan RPM out of range | If the fan RPM is out of Min/Max range. | 272 | ALARM_ID_HD_FAN_RPM_OUT_OF_RANGE |  | ALARM_PRIORITY_LOW | 850 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | HD Fan RPM out of range
[14 empty] | 273 | ALARM_ID_AVAILABLE_10 |  | ALARM_PRIORITY_LOW | 999 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | Available for use
[6 empty] | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Reservoir Fault |  | DG idle reservoir weight out of range | If the reservoir weight > 200.0 of the target mL. | 274 | ALARM_ID_DG_INACTIVE_RESERVOIR_WEIGHT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG inactive reservoir weight out of range
| PRS 494 | SA 237 | SRSHD 1706 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Arterial Press Range |  | HD arterial pressure sensor is reading out of range | If the arterial pressure > 2000.0 mmHg or < -1500.0 mmHg. | 275 | ALARM_ID_HD_ARTERIAL_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD arterial pressure sensor is reading out of range
| PRS 494 | SA 234 | SRSHD 1707 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Venous Press Range |  | HD venous pressure sensor is reading out of range | If the venous pressure > 2000.0 mmHg or < -1500.0 mmHg. | 276 | ALARM_ID_HD_VENOUS_PRESSURE_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD venous pressure sensor is reading out of range
| PRS 494 | SA 231 | SRSHD 1708 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Occlusion Sens Range |  | HD BP occlusion sensor is reading out of range | If the blood pump occlusion count is < 2000. | 277 | ALARM_ID_HD_BP_OCCLUSION_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD BP occlusion sensor is reading out of range
|  |  | SRSHD 793 |  |  | HD | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Reservoir Overfill |  | HD active reservoir weight out of range. | If the active reservoir weight is > 1900.0 g. | 278 | ALARM_ID_HD_ACTIVE_RESERVOIR_WEIGHT_OUT_OF_RANGE |  | ALARM_PRIORITY_MEDIUM | 421 | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD active reservoir weight out of range.
|  |  |  | SRSDG 752 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Reservoir Drain |  | DG dialysate drain time out | If the DG dialysate drain exceeded 120 seconds. | 279 | ALARM_ID_DG_DIALYSATE_DRAIN_TIME_OUT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG dialysate drain time out
| PRS 863 | SA 340 SA 344 |  | SRSDG 888 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: DRP Direction Fault |  | DG drain pump invalid direction | If the drain pump's direction is not 1 when draining. | 280 | ALARM_ID_DG_DRAIN_PUMP_DIRECTION_INVALID |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain pump direction invalid
| PRS 827 |  |  | SRSDG 836 |  | DG | Acid Concentrate Low |  | The acid concentrate is low.\n- Replace acid concentrate container. | Dial. Gen: Acid Conc. Low |  | DG Acid concentration bottle low volume alarm | If the acid concentration bottle volume > 3240 (3340 - 100). | 281 | ALARM_ID_DG_ACID_BOTTLE_LOW_VOLUME |  | ALARM_PRIORITY_MEDIUM | 412 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | Acid concentration bottle low volume alarm
| PRS 828 |  |  | SRSDG 837 |  | DG | Bicarb Concentrate Low |  | Bicarbonate concentrate is low.\n- Replace the bicarbonate concentrate container. | Dial. Gen: Bicarb Conc. Low |  | DG Bicarbonate concentration bottle low volume alarm | If the bicarbonate concentration bottle volume > 3680 (3780 - 100). | 282 | ALARM_ID_DG_BICARB_BOTTLE_LOW_VOLUME |  | ALARM_PRIORITY_MEDIUM | 413 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | Bicarbonate concentration bottle low volume alarm
| PRS 863 | SA 309 SA 310 SA 314 SA 315 |  | SRSDG 615 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: LC Range |  | DG load cells weight out of range | If the weight of the DG load cells > 4500.0 g or < 0.0 g. | 283 | ALARM_ID_DG_LOAD_CELL_WEIGHT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells weight out of range
| PRS 863 | SA 317 SA 318 |  | SRSDG 383 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Redundant LC Check |  | DG load cells primary/back up drift out of range | If redundant load cells deviate by more than 60g after tare for 5 seconds. | 284 | ALARM_ID_DG_LOAD_CELL_PRIMARY_BACKUP_DRIFT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG load cells primary/back up drift out of range
[6 empty] | HD | RO Permeate Sample in Progress |  | RO Permeate Sample is in progress.\n- Please wait for the process to complete. \n- You will not be able to create a treatment until the process\nis finished.\n- To end this process, press STOP on the front of the device. | Cleaning: RO Permeate Sample In Progress |  | DG RO permeate sample started | If DG RO permeate sample successfully started from UI. | 285 | ALARM_ID_HD_RO_PERMEATE_SAMPLE |  | ALARM_PRIORITY_LOW | 845 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD processor is in RO permeate sample mode
[6 empty] | DG | Concentrate Connection Cap Removed |  | The concentrate connection cap was removed from the \ndialysate device during disinfection.\n- Disinfection terminated.\n- Replace cap before attempting disinfection again. | Cleaning: Conc Cap Removed |  | DG cap in wrong state | If the concentrate cap is not in the right position in the cleaning modes. | 286 | ALARM_ID_DG_CONCENTRATE_CAP_NOT_IN_PROPER_POSITION |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG concentrate cap not in proper position
| PRS 494 | SA 222 | SRSHD 1403 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Proc- FPGA Clock |  | HD processor clock speed checks against FPGA clock failure | If during initialization mode, the difference between the processor clock speed and the FPGA clock speed > 1. | 287 | ALARM_ID_HD_FPGA_CLOCK_SPEED_CHECK_FAILURE |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD processor clock speed checks against FPGA clock failure
|  | SA 491 SA 492 | SRSHD 562 |  |  | HD | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: LC Drift Check |  | HD load cell inconsistency | If redundant load cells deviate from start to end of a reservoir use during treatment by more than 10g. | 288 | ALARM_ID_HD_LOAD_CELL_PRIMARY_BACKUP_DRIFT_OUT_OF_RANGE |  | ALARM_PRIORITY_HIGH | 110 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD load cells primary/back up drift out of range
|  |  |  | SRSDG 928 |  | DG | Dialysate Connection Cap Removed |  | The dialysate connection cap was removed from the\ndialysate device during disinfection.\n- Disinfection terminated.\n- Replace cap before attempting disinfection again. | Cleaning: Dialysate Cap Removed |  | DG cap in wrong state | If the dialysate cap is not in the right position in the cleaning modes. | 289 | ALARM_ID_DG_DIALYSATE_CAP_NOT_IN_PROPER_POSITION |  | ALARM_PRIORITY_MEDIUM | 600 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG dialysate cap not closed
| PRS 494 | SA 233 | SRSHD 1584 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Ven. Pres. Com |  | HD venous sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 290 | ALARM_ID_HD_VENOUS_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD venous pressure sensor FPGA fault
[6 empty] | HD | Flush In Progress |  | System flush is in progress\n- Please wait for the process to complete. \n- You will not be able to create a treatment until the process\nis finished.\n- To end this process, press STOP on the front of the device. | Cleaning: Flush In Progress |  | DG flush started | If DG flush successfully started from the UI. | 291 | ALARM_ID_HD_DISINFECT_FLUSH |  | ALARM_PRIORITY_LOW | 840 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD processor is in Disinfect Flush mode
[6 empty] | HD | Heat Disinfection in Progress |  | Heat disinfection is in progress.\n- Please wait for the process to complete. \n- You will not be able to create a treatment until the process\nis finished.\n- To end this process, press STOP on the front of the device. | Cleaning: Heat Disinf In Progress |  | DG heat disinfect started | If DG heat disinfect successfully started from the UI. | 292 | ALARM_ID_HD_DISINFECT_HEAT |  | ALARM_PRIORITY_LOW | 841 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD processor is in Disinfect Heat mode
[6 empty] | HD | Chemical Disinfection in Progress |  | Chemical disinfection in progress.\n- Please wait for the process to complete. \n- You will not be able to create a treatment until the process\nis finished.\n- To end this process, press STOP on the front of the device. | Cleaning: Chem Disinfect In Progress |  | DG chem disinfect started | If DG flush successfully started from the UI. | 293 | ALARM_ID_HD_DISINFECT_CHEM |  | ALARM_PRIORITY_LOW | 842 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD processor is in Disinfect Chemical mode
[6 empty] | HD | Active Cool in Progress |  | Active Cool is in progress.\n- Please wait for the process to complete. \n- You will not be able to create a treatment until the process\nis finished.\n- To end this process, press STOP on the front of the device. | Cleaning: Active Cool In Progress |  | DG active cool started | If DG active cool successfully started. | 294 | ALARM_ID_HD_DISINFECT_HEAT_COOL |  | ALARM_PRIORITY_LOW | 844 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | HD processor is in Disinfect active cool mode
| PRS 398 | SA 470 | SRSHD 1664 |  | SRSUI 1238 | UI HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: UI CloudSync |  | UI cloud sync POST failure | If the UI POST reports cloud sync failure. | 295 | ALARM_ID_HD_UI_POST_FAILURE_CLOUDSYNC | OK | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST CloudSync failure
| PRS 863 | SA 330 SA 331 |  | SRSDG 951 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: TDi RTD Comm |  | DG TDi sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 296 | ALARM_ID_DG_TDI_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG TDi sensors FPGA fault
| PRS 863 | SA 326 SA 327 |  | SRSDG 950 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: TRo RTD Comm |  | DG TRo sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 297 | ALARM_ID_DG_TRO_SENSORS_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG TRo sensors FPGA fault
| PRS 863 | SA 349 |  | SRSDG 948 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Baro Pres Comm |  | DG barometric sensor FPGA fault | If the FPGA does not report fresh data for a certain period of time | 298 | ALARM_ID_DG_BARO_SENSOR_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG baro sensor FPGA fault
| PRS 863 | SA 483 |  | SRSDG 952 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG POST: Serial Number Check |  | DG invalid serial number | If the DG serial number is empty and does not have any characters. | 299 | ALARM_ID_DG_INVALID_SERIAL_NUMBER |  | ALARM_PRIORITY_HIGH | 2 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG invalid serial number
| PRS 398 | SA 473 | SRSHD 1582 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: Serial Number Check |  | HD invalid serial number | If the HD serial number is empty and does not have any characters. | 300 | ALARM_ID_HD_INVALID_SERIAL_NUMBER |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD invalid serial number
[6 empty] | DG | Remove Dialysate Cap |  | Remove the dialysate cap\n - Place the concentrate connection cap on the device.\n- Keep dialysate cap on the device. | Cleaning: RO Permeate Sample Connect Cap |  | Open dialysate cap to be able to collecte the dispensing fluid | RO permeate sample collection state | 301 | ALARM_ID_DG_RO_PERMEATE_SAMPLE_REMOVE_DIA_CAP |  | ALARM_PRIORITY_LOW | 753 | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG RO permeate sample remove the dialysate cap
| PRS 863 | SA 485 |  | SRSDG 126 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Processor RAM error |  | DG_CPU RAM error | If the RAM status indicates any of the following: Double bit error, Address decode failure, Address decode logic element failure, Read address parity failure, Write address parity failure. | 302 | ALARM_ID_DG_CPU_RAM_ERROR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG processor RAM error
| PRS 863 | SA 345 |  | SRSDG 958 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: DRP Hall Commutation |  | DG drain pump direction FPGA fault | If the FPGA indicates an error for a certain period of time. | 303 | ALARM_ID_DG_DRAIN_PUMP_DIRECTION_FPGA_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG drain pump direction FPGA fault
| PRS 398 | SA 197 | SRSHD 1587 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST: Usage Rec NVRAM CRC |  | HD invalid usage reocord | If the CRC of the HD usage record does not match the calculated CRC. | 304 | ALARM_ID_HD_INVALID_USAGE_RECORD_CRC |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD invalid usage record
| PRS 494 | SA 484 | SRSHD 1577 |  |  | HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD Fault: Processor RAM error |  | HD_CPU RAM error | If the RAM status indicates any of the following: Double bit error, Address decode failure, Address decode logic element failure, Read address parity failure, Write address parity failure. | 305 | ALARM_ID_HD_CPU_RAM_ERROR |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD processor RAM error
[6 empty] | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: CD1 CD2 Check |  |  | CD1 and CD2 conductivity measurement during acid check is not within 5% of each other | 306 | ALARM_ID_DG_CD1_CD2_CHECK_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CD1 CD2 check fault
| PRS 396 |  | SRSHD 1594 |  |  | HD | Power Loss |  | Treatment must be terminated due to a power loss.\n- Follow clinic instructions on when to initiate manual rinseback.\n- Check all electrical connections.\n- If power returns within 5 minutes of outage, \n- rinseback will be allowed. | Blood Stop: Loss of AC Power |  | HD AC power lost alarm | The HD Software shall trigger an alarm if an AC power loss is detected while the HD system is in pre-treatment, treatment, or post-treatment mode. | 307 | ALARM_ID_HD_AC_POWER_LOST_IN_TREATMENT |  | ALARM_PRIORITY_HIGH | 100 | FALSE | FALSE | TRUE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | HD AC power lost second alarm
| PRS 863 | SA 348 |  | SRSDG 959 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: CPi Comm |  | DG CPi conductivity sensor transmits invalid characters | If the CPi conductivity sensor transmits invalid characters for a certain period of time. | 308 | ALARM_ID_DG_CPI_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CPi conductivity sensor invalid character received
| PRS 863 | SA 348 |  | SRSDG 960 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: CPo Comm |  | DG CPo conductivity sensor transmits invalid characters | If the CPo conductivity sensor transmits invalid characters for a certain period of time. | 309 | ALARM_ID_DG_CPO_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CPo conductivity sensor invalid character received
| PRS 863 | SA 347 |  | SRSDG 962 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: CD1 Comm |  | DG CD1 conductivity sensor transmits invalid characters | If the CD1 conductivity sensor transmits invalid characters for a certain period of time. | 310 | ALARM_ID_DG_CD1_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CD1 conductivity sensor invalid character received
| PRS 863 | SA 347 |  | SRSDG 961 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: CD2 Comm |  | DG CD2 conductivity sensor transmits invalid characters | If the CD2 conductivity sensor transmits invalid characters for a certain period of time. | 311 | ALARM_ID_DG_CD2_COND_SENSOR_INVALID_CHAR |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG CD2 conductivity sensor invalid character received
|  |  | SRSHD 1600 |  |  | HD | Low Dialysate Flow |  | Dialysate flow rate is low. \n- Confirm dialysate concentrate connections are secure. \n- When the flow rates are within target range \nResume/OK button will be available.\n- Press Resume/OK to continue. | Dial. Gen: Dialysate In Flow |  | The dialysate inlet flow is too low. | If the dialysate inlet flow < 10 mL/min (raw) while in treatment mode or pre-treatment mode after prime and DPi is running. | 312 | ALARM_ID_HD_DIAL_IN_FLOW_CHECK_FAILURE |  | ALARM_PRIORITY_MEDIUM | 400 | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | HD Dialysate In flow rate to too low.
| PRS 852 |  |  | SRSDG 1064 SRSDG 1052 SRSDG 968 |  | DG | Disinfection: Water Temperature Low |  | Incoming water temperature is too low.\n- Disinfection has been terminated.\n- Check incoming water before initiating disinfection again. | Cleaning: Inlet Water Temp Low |  | DG cleaning mode inlet water temperature too low | If inlet water temperature is < 24 C. | 313 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_TEMP_TOO_LOW |  | ALARM_PRIORITY_MEDIUM | 602 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water temperature too low.
| PRS 855 |  |  | SRSDG 1063 SRSDG 1051 SRSDG 967 |  | DG | Disinfection: Water Pressure High |  | Incoming water pressure is too high. \n- Disinfection has been terminated.\n- Check incoming water before initiating disinfection again. | Cleaning: Inlet Water Press High |  | DG cleaning mode inlet water pressure too high | If inlet water temperature is > 80 psig. | 314 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_PRESSURE_TOO_HIGH |  | ALARM_PRIORITY_MEDIUM | 605 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water pressure too high.
| PRS 856 |  |  | SRSDG 1068 SRSDG 1056 SRSDG 966 |  | DG | Disinfection: Water Pressure Low |  | Incoming water pressure is too low. \n- Disinfection has been terminated.\n- Check incoming water pressure\n before initiating disinfection again. | Cleaning: Inlet Water Press Low |  | DG cleaning mode inlet water pressure too low | If inlet water pressure is < 14 psig. | 315 | ALARM_ID_DG_CLEANING_MODE_INLET_WATER_PRESSURE_TOO_LOW |  | ALARM_PRIORITY_MEDIUM | 606 | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | DG cleaning mode inlet water pressure too low.
| PRS 863 | SA 495 |  | SRSDG 983 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Acid CP Pump Park |  | Acid concentrate pump did not park successfully | If the park took more than 1 second. | 316 | ALARM_ID_DG_ACID_CONCENTRATE_PUMP_PARK_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG acid concentrate pump park fault.
| PRS 863 | SA 496 |  | SRSDG 984 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: Bicarb CP Pump Park |  | Bicarb concentrate pump did not park successfully | If the park took more than 1 second. | 317 | ALARM_ID_DG_BICARB_CONCENTRATE_PUMP_PARK_FAULT |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG bicarb concentrate pump part fault.
|  |  | SRSHD 1609 |  |  | HD | Air Trap Control Error |  | A problem has been detected with the air trap level\ncontrol system.\n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue. | HD Fault: Air Pump Timeout |  | HD Air Pump operation timeout | If air pump has been turned on for greater than 10s. | 318 | ALARM_ID_HD_AIR_PUMP_TIMEOUT |  | ALARM_PRIORITY_HIGH | 10 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | TRUE | HD Air Pump timeout alarm
|  |  |  | SRSDG 1040 |  | DG | Residual Chemical Sampling |  | A water sample is needed to confirm the \nabsence of residual chemical following disinfection. \n- Test sample according to facility policy. \n- Confirm absence of residual chemical disinfectant. | Cleaning: Chem Residual Sampling |  | Collect sample to check the chemical disinfectant acid | Chemical disinfect flush sample alarm is prompted. | 319 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_FLUSH_SAMPLE |  | ALARM_PRIORITY_LOW | 752 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG chemical disinfect flush flush sample.
|  |  |  | SRSDG 1043 |  | DG | Chemical Residual Timeout |  | A chemical residual sample was not collected \n- within 30min of chemical flush.\n- To collect a sample for chemical residual,\nyou must repeat the chemical flush. | Cleaning: Chem Residual Timeout |  | The user did not proceed with collecting chemical disinfect flush sample for more than 30 minutes | Chemical disinfect flush sample was not collected form more than 30 minutes. | 320 | ALARM_ID_DG_CHEM_DISINFECT_FLUSH_SAMPLE_TIMEOUT |  | ALARM_PRIORITY_LOW | 752 | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | DG chemical disinfect flush sample timeout.
| PRS 863 | SA 356 |  | SRSDG 1069 |  | DG | Service Required: Dialysate Device |  | A problem was detected with the dialysate device. \n- Treatment must be terminated.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | DG Fault: UVo Error |  | Outlet reactor not healthy | If the outlet reactor are not healthy. | 321 | ALARM_ID_DG_OUTLET_UV_REACTOR_NOT_HEALTHY |  | ALARM_PRIORITY_HIGH | 110 | FALSE | TRUE | TRUE | TRUE | FALSE | TRUE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | FALSE | TRUE | FALSE | FALSE | DG outlet UV reactor not healthy
[6 empty] | UI HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST Alarm: UI Invalid Year |  | UI POST Invalid Year | If the UI POST reports a invalid year. | 322 | ALARM_ID_HD_UI_POST_FAILURE_INVALID_YEAR |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST Year check failure
[6 empty] | UI HD | Service Required: Hemodialysis Device |  | A problem was detected with the hemodialysis device.\n- Locate the ID code found in the bottom left corner\nof the alarm screen.\n- Call service to report the issue and schedule a repair. | HD POST Alarm: UI Bad Settings |  | UI POST Bad Settings | If the UI POST reports a bad settings. (e.g. no settings read, empty settings file, invalid/empty parms, CRC failure) | 323 | ALARM_ID_HD_UI_POST_FAILURE_SETTINGS_BAD |  | ALARM_PRIORITY_HIGH | 1 | TRUE | FALSE | TRUE | TRUE | TRUE | TRUE | TRUE | TRUE | FALSE | FALSE | FALSE | FALSE | TRUE | TRUE | FALSE | FALSE | HD UI POST Configuration check failure

## Property Definitions
Property Name | Definition |  | Column | Description
AlarmID | Unique alarm identification number |  | PRS ID | PRS requirement ID associated with the alarm.
Alarm Name | Name of the alarm |  | HD SRS ID | HD SRS requirement ID (if applicable) associated with the alarm.
Priority | Priority of the alarm (High, Medium, Low) |  | DG SRS ID | DG SRS requirement ID (if applicable) associated with the alarm.
Rank | Within the same priority, higher rank alarms have higher priority |  | Device | Sub-system that detects the alarm.
Fault | The alarm activation causes an HD fault |  | Description | Detailed description of the alarm.
DGFault | The alarm activation causes an DG fault |  | Displayed Title | Alarm title displayed prominantly at top of alarm dialog when the alarm is top priority.
Stop | The alarm activation causes a controlled stop |  | Displayed Instructions | Instructions displayed in alarm dialog when the alarm is top priority.
Clear Immediate | Alarm condition is cleared immediately after alarm activation |  | Trigger Conditions | Condition(s) required to trigger the alarm.
Non Recoverable | Alarm cannot be cleared (unrecoverable) |  | AlarmID | Unique alarm ID #.
No Resume | Alarm activation prevents treatment resume |  | Alarm Name | Enumeration name (code reference) for the alarm.
No Rinseback | Alarm activation prevents rinseback |  | Priority | Priority of the alarm (high, medium or low per 60601-1-8.
No End Treatment | Alarm activation prevents ending treatment |  | Rank | Rank of the alarm (used to determine which alarm is top priority) (low rank=high priority).
Block in Rinseback | If user opts to rinseback on this alarm, alarm will not re-trigger during rinseback |  | Fault | Does HD f/w consider the alarm to be a system fault.
Block in End Treatment | If user opts to end tx on this alarm, alarm will not re-trigger during post-tx |  | DGFault | Does DG f/w consider the alarm to be a system fault (will not prevent HD from some activities).
No Blood Recirc | Alarm activation prevents blood re-circulation |  | Stop | Does the HD f/w need to pause treatment (or perhaps other activities) when alarm is active.
No Dialysate Recirc | Alarm activation prevents dialysate re-circulation |  | Clear Immediate | Should the f/w immediately clear the alarm condition after trigger (not clear the alarm itself).
Clear Top Only | User acknowledgment clears only displayed alarm |  | Non Recoverable | Is the alarm recoverable (can user clear the alarm).
Treatment Log | Alarm will send a msg to UI for inclusion in Tx log |  | No Resume | Should treatment resume be prevented when the alarm is active.
|  |  | No Rinseback | Should rinseback be prevented when the alarm is active.
|  |  | No End Treatment | Should ending treatment be prevented when the alarm is active.
|  |  | Block in Rinseback | If user opts to rinseback on this alarm, alarm will not re-trigger during rinseback
|  |  | Block in End Treatment | If user opts to end tx on this alarm, alarm will not re-trigger during post-tx
|  |  | No Blood Recirc | Should blood recirculation be prevented when the alarm is active (treatment is paused). Dialyzer is bypassed regardless.
Category | Priority |  | No Dialysate Recirc | Should dialysate recirculation be prevented when the alarm is active (treatment is paused). Dialyzer is bypassed regardless.
HD POST | High |  | Clear Top Only | If user acknowledges alarm, should all active alarms be cleared or just the alarm top (displayed) alarm.
HD Machine Alarm | High |  | TreatmentLog | Should alarm be included in treatment log.
Blood Stopped - Rinseback Allowed | High |  | Doxygen Title | Used for design documentation purposes.
End Pre-Tx | Medium
Concentrate Warning | Low
Concentrate Alarm | Medium
Saline Warning | Low
Saline Alarm | Medium
Dialysate Alarms (Temp, conductivity, recirc?) | Medium
Other Procedural (setup, pre-tx, etc) | Medium/Low
Disinfection Alarms | Medium
HD Machine Alert | Low
DG Machine Alert | Low
Blood Pressure Alarms | N/A - not supported
Mode
Priority | Type